## applications
Applications and Application groups for CIB water quality instruments
version | 2017-04-25
Application groups
| Application group |  | Real-time multi-parameter |  |  |  |  | Autonomous multi-parameter |  | Autonomous and readable |  |  | Anti-fouling real-time multi-parameter
Applications
Application reference |  |  | A | B | C1 | C2 | C3 | D | E | F | G | H | I
Application title |  |  | marine tow fish (rugged and fast) | marine profiler | inland waterways / coastal tow fish and profiler: Zeeland | inland waterways / coastal tow fish and profiler: Wadden | inland waterways / coastal tow fish and profiler: IJM and rivers | CTD measurement project | Multi-parameter measurement project | manual measurements general multi-parameter | Manual temperature measurement | manual PAR light intensity measurement | Measuring pontoon in flow-through system
Measuring platform |  |  | tow fish | Profiler | Tow fish | Tow fish | Tow fish | Quayside, shore | Quayside, shore | From bridge, boat, shore | From bridge, boat, shore | Boat | Measuring pontoon
Indicative number in use | Expected number used individually |  | 1 | 2 | 2 | 2 | 2 | 40 | 3 | 5 | 7 | 6 | 4
Test-specific requirements
Parameters | temperature (T) | test range | -4 « 40 °C | -4 « 40 °C | -4 « 40 °C | -4 « 40 °C | -4 « 40 °C | -4 « 40 °C | -4 « 40 °C | -4 « 40 °C | -4 « 40 °C | Not requested | -4 « 40 °C
|  | accuracy | +/- 0.1 °C | +/- 0.1 °C | +/- 0.1 °C | +/- 0.1 °C | +/- 0.1 °C | +/- 0.1 °C | +/- 0.1 °C | +/- 0.1 °C | +/- 0.1 °C |  | +/- 0.1 °C
| conductivity (Ec) | measurement range | 0 « 6000 mS/m | 0 « 6000 mS/m | 0 « 6000 mS/m | 0 « 6000 mS/m | 0 « 6000 mS/m | 0 « 6000 mS/m | 0 « 6000 mS/m | 0 « 6000 mS/m | Not requested | Not requested | 0 « 200 mS/m
|  | accuracy | +/- 1% | +/- 1% | +/- 0.5% | +/- 0.5% | +/- 0.5% | +/- 1% | +/- 0.5% | +/- 0.5% |  |  | +/- 0.5%
| dissolved oxygen (DO) | measurement range | 0 « 20 mg/l | 0 « 20 mg/l | 0 « 20 mg/l | 0 « 20 mg/l | 0 « 20 mg/l | Not requested | 0 « 20 mg/l | 0 « 20 mg/l | Not requested | Not requested | 0 « 20 mg/l
|  | accuracy | +/- 0.2 mg/l | +/- 0.2 mg/l | +/- 0.2 mg/l | +/- 0.2 mg/l | +/- 0.2 mg/l |  | +/- 0.2 mg/l | +/- 0.2 mg/l |  |  | +/- 0.2 mg/l
| turbidity (D)- nephelometric | measurement range | Not requested | Not requested | 0 « 4000 NTU | 0 « 4000 NTU | 0 « 4000 NTU | Not requested | 0 « 4000 NTU | 0 « 4000 NTU | Not requested | Not requested | 0 « 4000 NTU
` |  | accuracy |  |  | +/- 3% | +/- 3% | +/- 3% |  | +/- 3% | +/- 3% |  |  | +/- 3%
| turbidity (D) - backscatter | measurement range | 0 « 4000 FTU | 0 « 4000 FTU | Not requested | Not requested | Not requested | Not requested | Not requested | Not requested | Not requested | Not requested | Not requested
|  | accuracy | +/- 3% | +/- 3%
| acidity (pH) | measurement range | 4 « 10 pH | 4 « 10 pH | 4 « 10 pH | 4 « 10 pH | 6 « 10 pH | Not requested | 6 « 10 pH | 6 « 10 pH | Not requested | Not requested | 6 « 10 pH
|  | accuracy | +/- 0.1 pH | +/- 0.1 pH | +/- 0.1 pH | +/- 0.1 pH | +/- 0.1 pH |  | +/- 0.1 pH | +/- 0.1 pH |  |  | +/- 0.1 pH
| fluorescence - Chlorophyl (fChl) | measurement range | 0 « 133 RFU | 0 « 133 RFU | 0 « 133 RFU | 0 « 133 RFU | 0 « 133 RFU | Not requested | 0 « 133 RFU | 0 « 133 RFU | Not requested | Not requested | 0 « 133 RFU
|  | accuracy | +/- 5% | +/- 5% | +/- 5% | +/- 5% | +/- 5% |  | +/- 5% | +/- 5% |  |  | +/- 5%
| fluorescence - cyanobacteria (fBGA) | measurement range | Not requested | Not requested | 0 « 133 RFU | 0 « 133 RFU | 0 « 133 RFU | Not requested | 0 « 133 RFU | 0 « 133 RFU | Not requested | Not requested | Not requested
|  | accuracy |  |  | +/- 2% | +/- 2% | +/- 2% |  | +/- 2% | +/- 2%
| PAR light intensity (Eo) for extinction | measurement range | Not requested | 1«5000 μmol/(m2s) | 1«5000 μmol/(m2s) | Not requested | Not requested | Not requested | Not requested | Not requested | Not requested | 1«5000 μmol/(m2s) | Not requested
|  | accuracy | - | +/- 5% | +/- 5% | [6 empty] | +/- 5%
| depth (d) | measurement range | 0 « 60 m | 0 « 60 m | 0 « 60 m | 0 « 60 m | 0 « 60 m | Not requested | 0 « 60 m | 0 « 60 m | Not requested | Not requested | 0 « 60 m
|  | accuracy | +/- 0.1 m | +/- 0.1 m | +/- 0.1 m | +/- 0.1 m | +/- 0.1 m |  | +/- 0.1 m | +/- 0.1 m |  |  | +/- 0.1 m
Additional parameters | CDOM | measurement range | 0 « 3000 ppb | 0 « 3000 ppb | Not requested | Not requested | Not requested | Not requested | Not requested | Not requested | Not requested | Not requested | Not requested
|  | accuracy | +/- 5% | +/- 5%
| pCO2 | measurement range | 0 « 4000 ppb | 0 « 4000 ppb | Not requested | Not requested | Not requested | Not requested | Not requested | Not requested | Not requested | Not requested | Not requested
|  | accuracy | +/- 5% | +/- 5%
| Accurate acidity (pH) | measurement range | 4 « 10 pH | 4 « 10 pH | Not requested | Not requested | Not requested | Not requested | Not requested | Not requested | Not requested | Not requested | Not requested
|  | accuracy | +/- 0.01 pH | +/- 0.01 pH
Measurement-specific preconditions
| Minimum measurement frequency (Hz) |  | ≥12 | ≥24 | ≥1 | n/a | n/a | ≥ 2 | n/a | ≥1 | n/a | ≥1 | ≥0.1
| Response time (per parameter) * The response time includes any pumping time, if a sensor pump is used for water supply. |  | Oxygen, acidity, fluorescence: around 1 minute. The rest 1 sec | Oxygen, acidity, fluorescence: around 1 minute. The rest 1 sec | Oxygen, acidity, fluorescence: around 1 minute. The rest 1 sec | Oxygen, acidity, fluorescence: around 1 minute. The rest 1 sec | Oxygen, acidity, fluorescence: around 1 minute. The rest 1 sec | About 1 second | Oxygen, acidity, fluorescence: around 1 minute. The rest 1 sec | Oxygen, acidity, fluorescence: around 1 minute. The rest 1 sec | 1 sec | 1 sec | Oxygen, acidity, fluorescence: around 1 minute. The rest 1 sec
Usage aspects
Data | Internal data logger |  | - | - | - | - | - | ≥ 100,000 readings | ≥ 100,000 readings | ≥ 100,000 readings | - | ≥ 1000 readings | ≥ 1000 readings
| Real-time data output |  | Yes, RS232 ASCII | Yes, RS232 ASCII | Yes, RS232 ASCII | Yes, RS232 ASCII | Yes, RS232 ASCII | - | - | - | - | - | Yes, RS232 ASCII
| Readable on instrument |  | - | - | - | - | - | - | - | Yes | Yes | Yes | -
Power supply | Internal power supply (battery) |  | - | - | - | - | - | Yes | Yes | Yes | Yes | - | -
| permanent (sufficient capacity available) |  | Yes | Yes | Yes | Yes | Yes | - | - | - | - | - | Yes
Version | standard |  | - | Yes | Yes | Yes | Yes | Yes | - | Yes | Yes | Yes | Yes
| compact/mobile |  | - | - | - | - | - | - | Yes | Yes | Yes | Yes | -
| rugged |  | Yes | - | - | - | - | - | - | - | - | - | -
Requirements from data-acquisition process | Frequency of use in water |  | Weekly | Zirfaea weekly: Arca only in emergency cases | Weekly | Weekly | Weekly | Project-based for a few days | Half-yearly, for one month | Weekly | Weekly | Weekly | Continually
| Anti-fouling |  | No | No | No | No | No | No | No | No | No | No | Yes
Aspects of environmental conditions | Max. current or navigation speed |  | 15 knots | n/a | 10 knots | 10 knots | 10 knots | n/a | n/a | n/a | n/a | n/a | n/a
| Temperature of body of water |  | 0 to 40 °C | 0 to 40 °C | 0 to 40 °C | 0 to 40 °C | 0 to 40 °C | 0 to 40 °C | 0 to 40 °C | 0 to 40 °C | 0 to 40 °C | 0 to 40 °C | 0 to 40 °C
| Maximum water depth (watertight) |  | To 60 meters | To 60 meters | To 60 meters | To 60 meters | To 60 meters | To 60 meters | To 60 meters | To 60 meters | To 20 meters | To 60 meters | To 20 meters
Integration in measurement configuration | Connection of interfaces beyond scope |  | PC, umbilical (2-core) | Echo sounder, PC, umbilical (2-core) | Echo sounder, PC, Umbilical (multiple cores available) | PC | PC | None | None | None | None | Wadden: PC and pressure sensor | PC
| Additional functionality within scope |  | Bundling measurement signals underwater | Bundling measurement signals underwater, incl. echo sounder | Bundling measurement signals underwater, incl. echo sounder | Bundling measurement signals underwater | Bundling measurement signals underwater | Stand-alone with own power supply and data storage | Stand-alone with own power supply and data storage | Convenient in use, stand-alone | Convenient in use, stand-alone | Bundling measurement signals, incl. pressure sensor | Bundling measurement signals
Software functions | Programming and calibrating |  | Yes | Yes | Yes | Yes | Yes | Yes | Yes | Yes | Yes | Yes | Yes
| Data-gathering interfaces beyond scope |  | No | Yes, echo sounder | Yes, echo sounder | No | No | No | No | No | No | Yes, pressure sensor (Wadden Sea) | No
| Presenting and exporting |  | Yes | Yes | Yes | Yes | Yes | Yes | Yes | Yes (after reading logger) | Yes (after reading logger) | No | Yes

## Blad1
|  | parameters | [10 empty] | data issue/storage | energy supply | used YSI equipment |  |  | alternative CIB measurement instrument
information required | measurement platform used | temperature | conductivity | dissolved oxygen | turbidity | acidity | fluorescence - Chlorophyl | fluorescence - cyanobacteria | nitrate | phosphate | ammonium | PAR |  |  | 6600v2 | 600XL | 600 OMS | TRDI Citadel
salt intrusion | permanent measurement arrays. Checks using manual measurements, buoy (project-related) | X | X | - | - | - | - | - | - | - | - | - | direct or data logger | internal and external | - | X | - | X
eutrophication | permanent measurement arrays, measurements while under way | X | X | X | X | X | X | X | X | X | X | X | direct or data logger | external | X (only "chemical") | - | - | Licor PAR
swimming-water quality checks | manual measurements | X | X | X | X | X | X | X | - | - | - | - | data logger | internal | X | X | -
enforcement | manual measurements | X | X | X | X | ? | ? | ? | ? | ? | ? | ? | direct or data logger | internal | X | X | -
various project measurements | permanent measurement arrays, measurements while under way and manual measurements | X | X | X | X | X | X | X | X | X | X | X |  | internal and external | X (only "chemical") | X (only "chemical") | X (only "chemical")
KRW | permanent measurement arrays, measurements while under way and manual measurements | X | X | X | X | X | X | X | X | X | X | X | direct | external | X (only "chemical") | - | -
KRM | Measurements while under way and manual measurements | X | X | X | X | X | X | X | X | X | X | X | direct | external | X (only "chemical") | - | -
power supply
application
| permanent |  |  |  | under way |  |  | manual
| stand-alone | connected to data-acquisition system |  |  | stand-alone | connected to data-acquisition system
project monitoring
parameter
temperature
conductivity
dissolved oxygen
turbidity
acidity
fluorescence - Chlorophyl
fluorescence - cyanobacteria
| permanent | under way | manual (incl. 650 MDS)
YSI 6600 v2 | x | x | X
YSI 600XL /650MDS |  |  | x
YSI 600 OMS | x |  | x
combination of parameters
parameter | temperature | conductivity | dissolved oxygen | turbidity | acidity | fluorescence - Chlorophyl | fluorescence - cyanobacteria
temperature | x
conductivity | x | x
dissolved oxygen | x | x | x
turbidity |  |  |  | x
acidity | x | x |  |  | x
fluorescence - Chlorophyl | x | x | x | x | x | x
fluorescence - cyanobacteria | x | x | x | x | x | x | x
conditions
salt/fresh
deep/shallow
ice-resistant