BIJLAGE 3a:	PROGRAMMA VAN EISEN, WENSEN EN OPEN VRAGEN

Behorende bij de Europese aanbesteding inzake de levering van een MR-Linac systeem, kenmerk 2048/SP

Inleiding

In dit document worden de doelstellingen van het Radboudumc met de MR-Linac gepresenteerd. Deze zijn vertaald in een programma van eisen, wensen en open vragen.

In hoofdstuk 2 van dit document zijn eisen geformuleerd waaraan Ondernemer moet voldoen. Ondernemer wordt gevraagd om per eis te bevestigen dat zij voldoet aan de gestelde eis. Dit doet zij door in de laatste kolom de eisen aan te vinken voor akkoord. Voldoet Ondernemer niet aan één of meerdere eisen, dan geldt dit als knock-out en komt zij niet in aanmerking voor gunning.

Hoofdstuk 3 bevat een drietal open vragen met betrekking tot samenwerking met Ondernemer, behorend bij gunningcriterium A. In dit hoofdstuk zijn ook de wensen geformuleerd, behorend bij de gunningcriteria B, C en D. Ondernemer wordt gevraagd om de open vragen te beantwoorden en voor de wensen per wens aan te geven of zij kan voldoen aan deze wensen met een ‘Ja’ of een ‘Nee’.

Voor een aantal eisen en wensen wordt Ondernemer gevraagd om aanvullende toelichting te geven. Ondernemer gebruikt hiervoor ‘Bijlage 3_Bijlage 3b_Programma van eisen en wensen EA MR-Linac_toelichting’. Desgewenst mag Ondernemer ook haar antwoord op andere eisen en wensen toelichten in dit document.

Indien er eisen en wensen zijn waar Ondernemer ten tijde van deze aanbestedingsprocedure niet aan kan voldoen, maar zij deze eis/wens wel in ontwikkeling heeft, dan dient Ondernemer dit separaat toe te lichten (zie pagina 12 van dit document). Ondernemer geeft daarbij een tijdspad aan wanneer Radboudumc mag verwachten dat wordt voldaan aan de eis/wens, waarbij een maximum geldt van 3 jaar na ontvangst Inschrijvingen. Ondernemer calculeert deze eisen/wensen al wel in haar inschrijfprijs. Indien blijkt dat Ondernemer niet binnen het voorgestelde tijdspad aan de eis/wens voldoet, dan geldt er een boeteclausule (zie artikel 19 Koopovereenkomst).

- Introduction

In this document, the Radboudumc presents its clinical and research objectives with respect to the MR-Linac and the requirements that should be fulfilled by the offered solution. Vendor is invited to submit an offer that suits these objectives and requirements. Vendor should clearly indicate to what extent the offered solution fulfills the requirements and is indeed equipped with the functionalities mentioned in this document. If for the system not all options are released (yet), but will be before the foreseen installation date, these options should be offered as a prototype and Vendor should support these until they are released as a product. Vendor takes into account in the offer that this should not lead to extra costs for the Radboudumc.

The Radboudumc department of Radiation Oncology currently possesses eight linac vaults, namely four at the Radboudumc main location and two at each of the two satellite departments (CWZ Nijmegen and Maasziekenhuis Pantein Boxmeer). Furthermore, the department features a CT simulator and an HDR brachytherapy vault. The Radboudumc wishes to replace the oldest linac with an MR-linac. The department plays a leading role in the treatment of cancers of head & neck, prostate, breast, lung, and gastro-intestinal origin. State-of-the-art techniques are used for both irradiation and imaging, the latter in cooperation with the Radboudumc department of Radiology & Nuclear Medicine. Besides clinical research and studies regarding patient involvement in medical decision making, the department also features the Radiotherapy & OncoImmunology Laboratory (ROI), which studies the biology of tumors at the macroscopic and microscopic level, comparing different functional imaging modalities for the same tumor. Two of the focus areas are molecular immunology and the tumor microenvironment. The department of Radiology & Nuclear Medicine encompasses renowned research groups on biomedical MR, MRI-guided interventions, holmium therapy and diagnostic image analysis. The diagnostic image analysis group (DIAG) is a world leader in deep learning-based image analysis of chest, breast and prostate. In addition, in the Radboudumc the cost-effectiveness of innovative technologies is studied using health technology assessments.

Objectives

MRI-guidance is the logical next step in image-guided radiotherapy, striving for optimal locoregional cancer treatment, which should be effective, safe, patient-friendly, and cost-efficient. An MR-Linac offers superior image quality as compared to CBCT, both anatomical and functional imaging, online and inline guidance in 3D or 4D, reduction of geometric uncertainties and maximum potential for (daily) plan adaptation. This generates opportunities for not only decreasing margins and toxicity, but also for dose painting/escalation based on dynamic functional imaging, analysis of online images for response monitoring/prediction, dosimetry, radiomics, and immunogenicity, and extreme hypofractionation with “one-stop-shop” treatment as the ultimate goal. The properties of MRI-guided radiotherapy are essential for the clinical application of and research in the field of two key objectives defined by the department of Radiation Oncology in close collaboration with the department of Radiology & Nuclear Medicine, namely 1) SBRT and 2) radioimmunotherapy (RIT).

Ad 1): For SBRT of early stage tumors and (oligo)metastases in (central) lung, abdomen and pelvis, we see opportunities in single-modality treatment but also in combination with systemic and other focal ablative therapies. Advanced imaging techniques such as MR lymphography, nano-MRI, and PSMA and PD-L1 PET can be combined with online MRI-guidance for dose painting and redistribution and response monitoring and prediction.

Ad 2): With respect to RIT, we would like to optimize the dose, fractionation and timing of RIT in a tumor-specific way. In order to do this, we need to analyze radiation-induced changes during and after the irradiation of tumor, lymph node and peritumoral area. We expect to be able to monitor and predict the immunogenicity of tumors, and test the abscopal effect of radioimmunotherapy. For this key objective, machine learning is important, as we strive to use “immunoradiomics” and deep learning-based approaches to predict the biological behavior of tumors.

Both objectives will exploit the advantage of treatment adaptation based on repeated (daily) MR-imaging. This requires a different diagnostic-therapeutic workflow allowing fast and frequent image registration, processing, delineation, treatment replanning, dose delivery and verification. The departments of Radiology & Nuclear Medicine and Radiation Oncology have a strong interest and complementary expertise in machine learning based autocontouring/autosegmentation, which is a critical component for adaptive MR-guided radiotherapy, response monitoring/prediction and single fraction based SBRT. Therefore, developing a workflow that allows rapid treatment adaptation will be another important collaborative research program of both departments within the MR-linac project.

In order to be able to study dose painting and (immune system) response monitoring, to relate to the advanced diagnostic imaging methods, and to collect features for (immuno)radiomics, it is essential to have an MR system with diagnostic qualities. The signal-to-noise ratio should be high and spatial resolution should be excellent, and there should be ample opportunities for functional imaging (i.e. DWI, DCE).

We have the ambition to occupy a so far empty spot in the spectrum of MR-Linac research, and the offered infrastructure should support this ambition. We would like to have a prominent position in the user research community and gain access to its protocols and databases. Practically, protocol adaptation should be possible and access to new MR sequences should be easy and fast. We hope to be supported by both Vendor and Vendor’s diagnostic partner to this end. In addition, besides the practical support, we are interested to join forces with respect to resources that can be used for research.

- EISEN / REQUIREMENTS

| Nr. | General | Conform? |
| --- | --- | --- |
|  | The system has a CE mark and complies with the applicable laws and regulations. |  |
|  | The system and available software (e.g. planning software) can be delivered as determined in the sales agreement ("Koopovereenkomst"). |  |
|  | The product is "state of the art" during delivery. "State of the art" means: At the moment of installation the system has to be equipped with all hardware and software developments, which came available commercially between registration to the European tender and installation of the system. If commercially available developments are excluded from the offer, please specify these in a list of options attached to your offer. In appendix 2 Vendor can describe the options. |  |
|  | Future enhancements of the product will be made available on the offered platforms, as such that the systems can stay “state of the art” in the next 10 years after installation. |  |
|  | The vendor provides radiation safety calculations for the current bunker on basis of the building drawings as attached in Appendix 1. These calculations can be limited to primary beam shielding, as the labyrinth and thus the scattered beam path might be subject to changes. All assumptions and values, which are necessary to reproduce the calculations, have to be supplied. After award, and when the building construction is clear, vendor will provide definitive radiation safety calculations. |  |
|  | The Faraday cage is included in the offer. |  |
|  | The vendor provides all accessories, phantoms, positioning utilities and software necessary for quality control and quality assurance. Please provide a list of all provided tools. The following equipment has to be included in the offer: - Waterphantom: Beamscan MR – L981561 (or similar) - Array dosimetry: Octavius 4D with 1500 detector (or similar) - Geometric distortion analysis system: QUASAR MRID3D (or similar) - 4D motion phantom: QUASAR MRI 4D Motion Phantom (or similar) All details about technical support are described in the service agreement ("Onderhoudsovereenkomst"). |  |
|  | Users will be trained to a level that they can operate the MR Linac in a qualified manner. All details about user training are described in the "Onderhoudsovereenkomst". |  |

| Nr. | Linac | Conform? |
| --- | --- | --- |
|  | The MR scanner and the linear accelerator are physically integrated in one system to allow for imaging and treatment simultaneously in one patient position. |  |
|  | The system facilitates patient treatment simulation by MR imaging on the treatment facility, subsequent planning using the acquired images and subsequent delivery of the treatment. |  |
|  | The following system specifications are fulfilled: - Photon energy ≥ 6MV - MLCs: leaf speed ≥4cm/s, full overtravel of MLC leafs possible - Leaf positioning accuracy at isocenter < 1 mm - Large beam field size ≥ 20cm × 25cm at isocenter |  |

|  | The system has a very robust design of beam forming system, resulting in < 2 MU stability of the beam after switching on the beam. Please specify time after which the beam profile is stable. | 3b: R12 |
| --- | --- | --- |
|  | The system is suited for 3D conformal and Intensity Modulated Radiation Therapy (IMRT) treatment. |  |
|  | The treatment couch contains very limited beam disturbing material, e.g. carbon fibre couch top. Please specify the max. water equivalent density of the couch components. | 3b: R14 |
|  | The system features a low collimator leakage within IEC requirements. Please specify collimator leakage. | 3b: R15 |
|  | The system features a low out of field leakage within IEC requirements. Please specify the out of field leakage. | 3b: R16 |

| Nr. | MRI | Conform? |
| --- | --- | --- |
|  | The system provides MR imaging of the following treatment regions: head&neck, thorax, breast, abdomen and pelvis. |  |
|  | The system is capable of acquiring dynamic MR imaging for motion monitoring during treatment with a minimum frame rate of 4/s. |  |
|  | The system is capable of performing functional imaging in order to analyse radiation induced tissue changes. Evaluation/post processing of functional imaging can be performed on the delivered workstations. |  |
|  | The system provides a high geometrical precision that is sufficient for precise treatment delivery. If necessary, methods to correct for geometrical deformation (e.g. 3D gradient warp) are provided. |  |
|  | The system allows for adaptation of the MR protocol parameters in case the images are not used for treatment planning, but treatment response measurement. |  |
|  | All coils, which are needed to perform imaging of the brain, head&neck, thorax, abdomen and pelvis region, are provided. All delivered coils are suited to be used during radiotherapy treatment. |  |

| Nr. | Process & Workflow | Conform? |
| --- | --- | --- |
|  | The system provides an online console (supporting image acquisition, online plan adaptation, motion monitoring and treatment delivery) and a separate imaging console in the control room. Furthermore two offline treatment planning stations are provided. |  |
|  | The planning system accounts for densities of the MR coils and the couch. |  |
|  | The dose calculation and optimization algorithm of the planning system accounts for effects related to the magnetic field. |  |
|  | The vendor provides a laser system for patient positioning, with a laser line from the ceiling and two crosshair lasers from the sides. |  |
|  | The system facilitates the use of treatment masks for patient positioning. A thorax board is included in the offer. Please specify if other positioning utilities are provided. All patient positioning utilities are MR compatible. | 3b: R27 |
|  | The coordinate systems of the MR scanner and of the linear accelerator have to be identical with a precision ≤ 1 mm. |  |
|  | The patient can be moved to the isocenter of the bore automatically. |  |
|  | The vendor provides solutions to facilitate patient comfort during image acquisition and treatment (e.g. music, light). Please specify your solutions. | 3b: R30 |

| Nr. | ICT | Conform? |
| --- | --- | --- |
|  | The system fulfils the requirements given in ‘Bijlage 7 Connection Requirements’ regarding network communication. |  |
|  | The system complies with DICOM(-RT) compatibility. Please include DICOM conformance statement. |  |
|  | A record and verify (R&V) system is necessary to store the technical treatment planning and delivery data for radiation therapy patients. This can be realized by Elekta Mosaiq or a vendor specific R&V solution. Alternatively HL7-ADT standardized export of treatment delivery data has to be provided. Please specify your R&V solution. | 3b: R33 |
|  | Acquisition and treatment computers provide a solution to prevent loss of data during (power) failure, e.g. redundant execution of critical ICT data. Please specify your solution. | 3b: R34 |
|  | The vendor provides a backup system to backup all treatment data. This backup system is part of the offer. |  |
|  | After an update and/or upgrade of the system the connectivity and interoperability with further equipment or systems will be warranted by the vendor. (Verification of connectivity and interoperability is a joint action as part of acceptation. This has to be checked after every system update/upgrade). |  |
|  | All subsystems use the same patient-ID unambiguously. |  |

| Nr. | Safety | Conform? |
| --- | --- | --- |
|  | The system is suited to be placed in a group 1 clinically used room (NEN 1010:2015). |  |
|  | The patient couch can be pulled out of the bore manually to evacuate the patient in case of (power) failure. |  |
|  | The system is equipped with a CCTV camera system with monitor to watch the patient from the control room. |  |
|  | The treatment room has an oxygen sensor which automatically gives an alarm in case of too low oxygen level. |  |
|  | The system has a device to communicate bi-directionally with the patient during measurement and/or treatment, like e.g. earphones. If a head coil is used, communication still has to be guaranteed. Please describe your solution. |  |
|  | Acquisition and treatment computers have a provision to not lose data in case of power failure. |  |
|  | The system has a door interlock which stops treatment in case the door is open. |  |
|  | The console offers various access levels. Technicians, sufficiently certified by the vendor, have full access. |  |
|  | The accelerator has an emergency stop to stop the treatment. The already given part of the fraction is recorded and treatment can be continued on a later moment. |  |

|  | The delivery and machine parameters of rotational and non-rotational treatment are continuously registered. In case of machine break down the already delivered part of the interrupted treatment is stored and can be resumed later. |  |
| --- | --- | --- |
|  | Quench buttons to shut down the magnetic field are mounted in the scanning room and control room. |  |
|  | The 5 Gauss line is marked on the floor. |  |

- WENSEN / PREFERENCES

| Nr. | Research cooperation | Answer |
| --- | --- | --- |
|  | Radboudumc has great research ambitions with main focus on radio-immunotherapy, (immuno-)radiomics, auto-contouring, auto-segmentation and auto-tracking. We wish to be part of a research consortium and have access to its joint database and protocols, to be able to contribute to these research fields on a high international level. Please specify the possibilities and timeline of research collaborations and integration in a research consortium, to facilitate our ambitions on fundamental research. What are the conditions for a membership in a research consortium? Please provide this information according to the SMART criteria. | separate |
|  | We would like to jointly develop three research collaborations on 1) radio-immunotherapy, 2) (immuno-)radiomics and 3) auto-contouring, auto-segmentation and auto-tracking. These collaborative projects will describe topics, duration, and manufacturer resources and support available for a successful 10 year partnership in the respective research lines. Please provide your research proposal according to the SMART criteria. | separate |
|  | To facilitate our research lines, the newest MR sequences and innovations on functional imaging should be implemented soon after development. Implementation and verification of newly developed functional imaging sequences can be part of a “work in progress” research collaboration. Please specify the possibilities of a research collaboration on MR sequence innovations. Please also specify the conditions under which innovations that were part of a research collaboration can continue to be used by Radboudumc after they received product status. | separate |

| Nr. | Linac | Answer |
| --- | --- | --- |
|  | The system provides a robust method and failsafe delivery of an IMRT plan. Please specify open quality issues and improvement strategies for plan delivery. | Yes No 3b: P4 |
|  | The system provides the capability to deliver a VMAT plan. | Yes No |
|  | The system provides the capability to present a 3D in-vivo delivered dose distribution after treatment for dose accumulation. | Yes No |
|  | The system provides the capability to perform treatment gating based on MR motion monitoring. | Yes No |

| Nr. | MRI | Answer |
| --- | --- | --- |
|  | The MR scanner has a magnetic field strength of ≥1,5T to provide a high intrinsic signal to noise ratio and excellent spatial resolution comparable to that of diagnostic MR scanners, also allowing for functional imaging. | Yes No |

|  | The MR scanner has to deliver excellent imaging quality with a high soft tissue contrast and high geometrical precision. Please provide pseudonymisized imaging data sets (T1w, T2w, real time imaging, DWI (EPI and other technique, if available), and a treatment planning imaging data set) of the following treatment regions: -head & neck • Thorax • Abdomen • Pelvis Scoring: The images will be evaluated on image quality, image resolution and artefacts. | Yes No 3b: P9 |
| --- | --- | --- |
|  | Radboudumc is especially interested in functional imaging in order to analyse radiation induced tissue changes and dynamic imaging techniques such as parallel imaging, golden angle radial or other non-Cartesian sampling, in combination with compressed sensing or deep learning-based reconstruction. Furthermore MR sequences which provide a MR-only workflow are of interest (e.g. DIXON). Please provide a list of all available MR sequences and a timeline of sequences which will become available till 1/1/2025. Please also specify for which body parts sequences are approved. | P10 |
|  | The system is capable to acquire dynamic MR imaging during treatment in all three imaging planes with a minimum frame rate of 4/s. | Yes No |
|  | The system is capable of performing 1H single voxel spectroscopy (SVS) and magnetic resonance spectroscopic imaging (MRSI), to analyse radio induced changes in metabolism. SVS and MRSI evaluation/post processing can be performed on the delivered workstations. | Yes No |
|  | The vendor provides a head&neck coil for detailed imaging of the head and neck region. The coil can be used in combination with a treatment mask. | Yes No |
|  | The vendor provides surface coils for detailed imaging of the thorax and abdomen region. | Yes No |

| Nr. | Process & Workflow | Answer |
| --- | --- | --- |
|  | The system allows for treatment planning based on MR images (pseudo CT), without the need of CT images. | Yes No |
|  | The system is suited for adaptive treatment strategies. Please provide a description of the current workflow of adaptive treatment, including MR imaging, image registration, contouring, plan adaptation and treatment. Please specify the workflow steps, mouse clicks and key strokes and the time necessary for the following treatments: - Prostate - Head&neck - liver stereotactic treatment Scoring: The workflow will be evaluated according to user friendliness and time efficiency. | Yes No 3b: P16 |
|  | Two delineation and plan review licences are included in the offer. | Yes No |
|  | The planning software provides auto-contouring. | Yes No |
|  | The system complies with IHE-RO conformity. | Yes No |
|  | The system supports the DICOM Modality Worklist (DMWL). | Yes No |
|  | The system supports communication with the Hospital RIS / PACS systems. | Yes No |
|  | The system provides the possibility to review results of imaging, positioning, dose tracking and planning remotely. | Yes No |
|  | The system provides access to its internal databases (read only) for reasons of logistics and / or clinical research. | Yes No |

Toelichting eisen en wensen

Ondernemer geeft in onderstaande tabel een toelichting bij de eisen en wensen waar zij ten tijde van deze aanbestedingsprocedure niet aan kan voldoen, maar waarvan zij garandeert dat zij binnen 3 jaar na ontvangst van de Inschrijvingen wel kan voldoen.

| Nr. requirement /wish | Toelichting door Ondernemer |
| --- | --- |
|  | Tijdspad: |
|  | Tijdspad: |

Signing of this document

The undersigned declares that all questions and attachments have been completely and truthfully answered and that the information provided in this questionnaire corresponds to reality and is accurate and complete.

The undersigned is aware and agrees that Radboudumc may audit this document and the submitted proofs and attachments or will have them audited. The undersigned will give every assistance if Radboudumc wishes to audit the data. The undersigned will likewise cooperate in an investigation into the ‘source of funds’ performed by or on behalf of Radboudumc

Drawn up and completed truthfully and validly (by authorised agent) and signed by;

| Name supplier |  |
| --- | --- |
| Name signatory |  |
| Date |  |
| Signature |  |

Appendix 1

Appendix 2:

| Requirement 5: | The product is "state of the art" during delivery. "State of the art" means: At the moment of installation the system has to be equipped with all hardware and software developments, which came available commercially between registration to the European tender and installation of the system. If commercially available developments are excluded from the offer, please specify these in a list of options attached to your offer. (Ondernemer gebruikt hiervoor bijlage 8 Prijsopgaveformulier, tabblad opties) |
| --- | --- |
| Description options: | …. |