## Appendix A05
Appendix A05, Conformity list 'Programme of Requirements' Guidelines for the National Open Tender for a Light Source, ref. TNO 2020 FPL-INK 182 dated 16-10-2020
By the following Programme of Requirements TNO will assess if Tenderer meets the relevant requirements as laid down in Section 6 and 8. The Tenderer must indicate for each item, the conformity for the requirements in the table below.
If to any of the requirements, as stated in section 8, are not met by Tenderer, he will be excluded for any further assessment. By answering a question with “compliant” will mean that in the event of awarding the contract, non-conformity to the requirement shall count as failure to meet an obligation from the Agreement to be concluded.
Requirement | Decription | Tenderer is compliant to the requirement | Reference in Tenderer's documentation where compliances with the requirement is described | Remark
Requirement 8.1.1. | Tenderer must have a service department. - Confirmation of receipt of maintenance / or service request or complaint, within 2 days*. - Service request – Light Source down – within warranty, within 5 days* at site; - Service Request – Light Source down – after warranty, within 10 days* at site. * (office) hours and regular working days | Click the drop-down box to select
Requirement 8.1.2 | The service organization must be able to carry out all relevant (field) repairs too keep the light source fully operational during 50 weeks a year. | Click the drop-down box to select
Requirement 8.1.3 | To reduce down time, tenderer must keep all common replacement parts, critical replacement parts and consumables on stock. The tenderer must provide such a list as well as the prices, subdivided into the mentioned subgroups, and where the prices are fixed during the guarantee period as stated in the AIV-TNO, clause 11.1. | Click the drop-down box to select
Requirement 8.1.4 | All common replacement parts, critical replacement parts and consumables shall be available for at least 10 years after delivery of the SAT | Click the drop-down box to select
Requirement 8.1.5 | Tenderer must be able to supply telephone assistance through a service hotline during regular working hours. Tenderer must have a 5 working day warranty response time. | Click the drop-down box to select
Requirement 8.1.6 | All members of the Tenderer’s staff who are responsible for the performance of the Contract must have an adequate command of written and spoken English. All correspondence will be conducted in English or Dutch. | Click the drop-down box to select
Requirement 8.2.1 | The light-source has a spectrum resembling the solar spectrum. Clarification: • It does not need to be an exact resemblance of the solar spectrum, a xenon-arc light source ,as is, would suffice, • However: we need to know the exact emission spectrum of the light source. At least an emission spectrum between 350-2500 nm, emitting light in all these wavelengths. | Click the drop-down box to select
Requirement 8.2.2 | The intensity of the light source should be variable between 3-12 sun intensities | Click the drop-down box to select
Requirement 8.2.3 | Target size is 200mm*200 mm | Click the drop-down box to select
Requirement 8.2.4 | Uniformity +/-10% non-uniformity of irradiance The point to be measured the uniformity the direct at the light exit point. | Click the drop-down box to select
Requirement 8.2.5 | The length of the simulator must not exceed the following dimensions 275 x 65 x 65 cm (Length x Depth x Height). It should be able to be placed on the lab-table in picture, where the point of light outcoupling should be via an opening in the side of the fumehood. Hereby the light should be traveling in the direction parallel to the length. | Click the drop-down box to select
Requirement 8.2.6 | It should be possible to align the Light source emission point with the reactor, for more information see §2.2.24 | Click the drop-down box to select
Requirement 8.2.7 | In the design it should be foreseen that optical filters / long pass filters can be fitted (for instance for filtering out all UV below 400nm). | Click the drop-down box to select
Requirement 8.3.1 | Documentation of third party hardware The Tenderer shall deliver certificates, CE declarations of conformity, and manuals for all the third party hardware that is part of the Light Source Assembly. Remark: Delivery of documentation is expected during SAT. Rationale: TNO quality requirement. | Click the drop-down box to select
Requirement 8.3.2 | Safety document: A document should be included how to safely work with the high intensity light source, including a description of it hazerds | Click the drop-down box to select
Requirement 8.3.2 | The light source assembly system shall be delivered with a manual. This manual shall include the necessary drawings and electric layouts to perform simple troubleshooting and shall include basic operation instructions. Rationale: Necessary for simple trouble shooting or for operating the light source assembly. | Click the drop-down box to select
Requirement 8.3.3 | The light source assembly dimensions The Tenderer shall state all the dimensions of the Light source assembly in the tender document. | Click the drop-down box to select
Requirement 8.3.4 | Documentation: • Instruction manual • Maintenance manual including all required technical documentation for Preventive maintenance • and Corrective Maintenance • Quick- start guide • List of spare parts including price for TNO • CE certification • Safety measures and their implementation using the setup | Click the drop-down box to select
Requirement 8.3.5 | All documentation must be available in English, both in PDF format and on paper. | Click the drop-down box to select
Requirement 8.4.1 | The offered Light source assembly and associated (necessary) accessories comply at all times with the applicable Dutch and European legal requirements, such as (but not limited to): NEN, EN and CE standards and their use including the working conditions - legislation. (93/42 / EEC); | Click the drop-down box to select
Requirement 8.4.2 | The Tenderer declares that goods are always provided with a valid CE marking upon delivery. The CE marking complies with all current and future European and Dutch legislation and regulations, even if the scope of the legislation and regulations changes in the period between the Contract awarding and the final delivery date; | Click the drop-down box to select
Requirement 8.4.3 | The Tenderer guarantees the continuity of delivery in the event of changes in European or national legislation, even if the scope changes. | Click the drop-down box to select
Requirement 8.4.4 | Tenderer complies with international sanctions and export control legislation. | Click the drop-down box to select
Requirement 8.4.5 | TNO does not accept additional costs resulting from the departure of the United Kingdom from the European Union. These (extra) costs include, but are not limited to, additional import duties, extra costs as a longer delivery time, exchange rate fluctuations, increased costs in the supply chain of the Contractor et cetera. | Click the drop-down box to select

## validatie (hidden)
Click the drop-down box to select
Yes, compliant to requirement
No, not compliant to requirement
note: dit blad 'verbergen' indien je dit gaat delen intern/extern (behoudens aan inkoop collega's)