## P.v.E.
ANNEX 1a - PROGRAM OF REQUIREMENTS European tender " Radiation device "
Rating: Failure to comply with 1 or more Knock-Outs (KO) means that your Tender will no longer be eligigle to compete for the award of this contract. Compliance with all KO is considered as an condition for the European Tender. Completing all KOs does not earn extra points.
Explanation: Failure to (incorrectly) fill in will result in the Tender being declared null and void, so please follow the instructions carefully. The PoR must be completed truthfully. - Tenderer only to fill in the yellow cells. - In column C (KO), the Tenderer indicates with "yes" if the Tenderer meets a requirement and with "no", if the Tenderer does not meet a requirement. - If 'Yes' is noted in column D (Evidence required for registration), then the Tenderer must provide documentation with the tender (see column E) so that Erasmus MC can confirm that the stated requirement has been met. For all requirements where explicit documentation is required for Tender (column D), failure to provide the requested information means that your Tender will no longer be eligible to compete for the award of this contract.
Criteria No. | Requirements | KO | Evidence required for registration (Yes/No) | Type of Evidence required if applicable
Section 1 | Regulatory Approval
1.1 | The Tenderer guarantees that the irradiator offered complies with the highest legal requirements and other government regulations, according to European laws and regulations and regulations by local authorities, as well as the highest requirements of safety, quality and environmental standards used within industry. | O Yes O No | No
1.2 | The irradiator must be CE marked (medical device directive 93/42/EEC and ISO13485). The Tenderer must provide quotation evidence of CE certification in their tender. | O Yes O No | Yes | CE Certification
1.3 | The Tenderer is able to provide, upon request by Erasmus MC, evidence of the unique identificationof the irradiator. | O Yes O No | No
1.4 | The irradiator must be designed to ensure Health and Safety statutory compliance as amended, and that hazards from ionising radiation are ALARP (As Low As Reasonably Practicable) | O Yes O No | Yes | Through report. The dose rate, measured at a distance of 0.1 m from the X-ray machine, must not exceed 1 microsievert per hour at any point around the X-ray machine.
1.5 | Blood irradiation dosimetry procedures must be compliant with ISO/ASTM 51939:2016/2017+B16 | O Yes O No | Yes | The supplier must provide a health and safety risk assessment, including design features to ensure safety for operators (according to the ALARA principle)
1.6 | Tenderer shall comply with all other applicable laws, especially Dutch and EU laws. Key Dutch laws are: "Kernenergiewet", "ANVS-Verordening" and "Besluit basisveiligheidsnormen stralingsbescherming". Key EU law is Euratom Treaty with as latest directive 2013/59. | O Yes O No | No | On request Erasmus MC
Section 2 | Operational requirements
2.1 | Upon installation, The irradiator must come supplied with a comprehensive user manual (including all user operational use and safety instructions) in English. | O Yes O No | Yes | User manual
2.2 | The irradiator must use 230V power. No other inputs or outlets will be accommodated. Erasmus MC will not be responsible for incorrect power supplies, cables or connectors provided with the irradiator. | O Yes O No | Yes | Statement - reference to page in user manual
2.3 | The irradiator must have visible / audible displays and / or warning lights. These must indicate status before, during and post operation to clearly indicate irradiator status. | O Yes O No | Yes | Statement - reference to page in user manual
2.4 | The irradiator must clearly confirm completion of irradiation cycle (visible, audible display or panel illumination). | O Yes O No | Yes | Statement - reference to page in user manual
2.5 | The irradiator must have a calibrated timer or method for control of irradiation dose. The timer must be capable of being pre-set to a given time. The adjustment of the time must only be available to authorised personnel and not accessible during routine operation e.g. individual keys or password protection. | O Yes O No | Yes | Statement - reference to page in user manual
2.6 | The irradiator must have an independent, calibrated back up timer or alternative method available to verify the timer setting. | O Yes O No | Yes | Statement - reference to page in user manual
2.7 | The correct dose must be applied irrespective of the volume or variety of components inserted into the irradiator canisters. | O Yes O No | No
2.8 | Irradiated red cell quality markers i.e. potassium leakage and haemolysis as a minimum, must be equivalent to published data for components irradiated at the same absorbed dose rate. Transfusion volume 58, april 2018 (BEST) 918-926, Timing of gamma irradiation and blood donor sex influences in vitro characteristics of red blood cells. | O Yes O No | No
2.9 | In routine procedure the core temperature of any product will not increase more than 4°C compared with the starting point of the products as a result of the irredation cycle. | O Yes O No | No
2.10 | The irradiator must be capable of delivering the recommended minimum dose of 25Gy to components within irradiation field, with no part receiving >50 Gy (Erasmus MC Guidelines Bloodproducts PT009.RL.BB current edition, EDQM Guide to the preparation, use and quality assurance of bloodcomponents / Council of Europe current version). | O Yes O No | No
2.11 | The irradiator warm-up cycle should not exceed 5 minutes. | O Yes O No | No
2.12 | The time to complete irradiation to the required dose (25-50Gy) must not exceed 7 minutes over the lifetime of the equipment. | O Yes O No | No
2.13 | The irradiator should be capable of performing at least 30 radiation cycles within a 10 hour period and capable of irradiating up to 25,000 components per year. | O Yes O No | No
2.14 | The dimensions of the x-ray will be such that the x-ray machine will be easily placed in excisting sites of Erasmus MC (minium modification in excisting buildings). | O Yes O No | No
2.15 | The irradiator cooling system must be built into the main body as a closed system or casing of the irradiator i.e. within the same footprint. Irradiators requiring external inputs or equipment such as external power supplies (not including cables), coolers, heat exchangers etc. will not be accepted. The X-ray device must have a self-contained water reservoir for cooling the X-ray tube. The energy consumption in kilowatts per hour must be as low as reasonably possible. | O Yes O No | No
2.16 | The irradiator coolant must be changeable and designed in such a way as to enable management/control of microbial growth. | O Yes O No | No
2.17 | There must be a mechanism available to monitor coolant as necessary. | O Yes O No | No
Section 3 | Training Requirements
3.1 | The Tenderer must provide the outline of the proposed training programme to be provided to new users. | O Yes O No | No
3.2 | The Tenderer must provide on-site training for the use of the irradiator for a minimum of four individuals free of charge per site. | O Yes O No | No
3.3 | The Tenderer must include the use of the installed irradiator's software, if applicable, as part of the training. | O Yes O No | No
3.4 | The Tenderer must include in-house maintenance requirements as part of the training. | O Yes O No | No
3.5 | The Tenderer must include in-house cleaning requirements as part of the training including compatible cleaning materials and instructions. | O Yes O No | No
3.6 | The Tenderer must provide installation and training for any hardware or software upgrades during the life of the irradiator free of charge. | O Yes O No | No
3.7 | The Tenderer must ensure all training provided is logged and certificates of competence provided on completion. This must include names, dates, times, places and details of the training undertaken. | O Yes O No | No
Section 4 | Health and Safety
4.1 | The irradiator must be adequately and securely shielded to ensure that dose rate is as low as reasonably achievable at all accessible points in all service modes of operation compliant with must comply with the Decree on basic safety standards for radiation protection (Bbs). | O Yes O No | Yes | Statement - reference to relevant page(s) in user manual.
4.2 | No ionising radiation must be produced when the chamber door is open. | O Yes O No | Yes | Statement - reference to relevant page(s) in user manual.
4.3 | The irradiator must be designed such that it is not possible for the operator to come into contact with moving parts of the equipment whilst the machine is in operation. | O Yes O No | Yes | Statement - reference to relevant page(s) in user manual.
4.4 | There must be no danger of the irradiator tipping or becoming unstable during normal operation. | O Yes O No | Yes | Statement - reference to relevant page(s) in user manual.
4.5 | The Tenderer must disclose to Erasmus MC any health and safety risk associated with the use of the machine. | O Yes O No | Yes | Statement - reference to relevant page(s) in user manual.
4.6 | The design of the irradiator must allow for relocation or access for support, ensuring that operator and technician health and safety is not compromised. | O Yes O No | Yes | Statement - reference to relevant page(s) in user manual.
4.7 | The Tenderer must provide reference to the page in the user manual that explains how the irradiator can be moved or accessed. | O Yes O No | Yes | Statement - reference to relevant page(s) in user manual.
4.8 | The irradiator must allow unobstructed access to the operator panel free of any health and safety implications for all Erasmus MC operators. | O Yes O No | Yes | Statement - reference to relevant page(s) in user manual.
4.9 | The position and value of the maximum sound pressure must be specified. This must be conform directive 2006/42/EG from the European Parlement and the council of 17 May 2006 concerning machines and amending directive 95/16/EG). Placing the X-ray device in the workplace must not lead to an A-weighted noise emission pressure of 70 dB (A) being exceeded. | O Yes O No | Yes | Statement - reference to relevant page(s) in user manual.
Section 5 | Installation
5.1 | Following installation the Tenderer must undertake a critical examination of the irradiator to confirm that any safety features and warning devices operate correctly and that there is sufficient protection for persons from exposure to ionising radiation, with must be in compliance with the Decree on basic safety standards for radiation protection (Bbs). | O Yes O No | No
5.2 | The Tenderer must map the absorbed dose rate (with components) to confirm dose distribution is compliant with Sanquin guidlines Sanquin Guidelines for Blood products PT009.RL.SQ - current edition (25-50Gy - see definitions) and Guide to the preparation, use and quality assurance of blood components (25-50Gy - see definitions) under both minimum and maximum loading configurations. Dosimeters compatible with ISO/ASTM 51939:2016/2017 must be used and the procedure to assure dose distribution must be to the satisfaction of the Responsable Person Sanquin Bloodbank, The Tenderer must conduct a dose map of products. The Tenderer must provide a written report which will form part of the installation qualification and operational qualification procedure. | O Yes O No | No
Section 6 | Data
Note | The following refers to all and any data collected as part of normal operation of servicing of the supplied irradiator.
6.1 | Data Integrity, that is the extent to which data is complete, consistent and accurate (and retained) throughout its lifecycle, must be determined for equipment using computerised systems to perform critical operations for a desired product quality. | O Yes O No | Yes | Define what data is collected and how it is stored throughout its lifecycle on irradiator and / or software
6.2 | The Computerised systems should comply with the requirements of EU GMP Annex 11. Eudralex volume 4 part 1 chapter 3. The Tenderer should provide a statement. | O Yes O No | Yes | Please provide a statement
6.3 | The irradiator must have access controls to ensure that users have access only to functionality that is appropriate for their job role, and that actions are attributable to a specific individual. | O Yes O No | No
6.4 | The system design must always provide for the retention of full audit trails to show all changes to the data while retaining previous and original data. | O Yes O No | No
6.5 | All changes to critical data must detail the users making those changes, and changes must be time stamped and a reason given. | O Yes O No | No
6.6 | The Tenderer must design the computer systems to ensure that the execution of critical operations are recorded contemporaneously | O Yes O No | No
6.7 | The Tenderer must identify critical settings which may impact outcome, if changed inadvertently, and ensure there is an ability to confirm settings have not been changed by the end user. | O Yes O No | No
6.8 | Data and document retention arrangements must ensure the protection of records from deliberate or inadvertent alteration or loss | O Yes O No | No
Section 7 | Maintenance, repair and documentation
7.1 | All irradiators supplied must use the same software version. Any changes to the software version must first be agreed with Erasmus MC. Where agreed changes are made, the irradiator Tenderer must notify Erasmus MC no less than 1 month before the software version is changed. | O Yes O No | No
7.2 | The Tenderer actively communicates (during the entire lifespan of the system) about software updates to the Department of Clinical Chemistry of the Erasmus MC and makes these available. | O Yes O No | No
7.3 | Costs of software updates are included in the offer for the entire duration of the system. | O Yes O No | No
7.4 | Changes to the documentation will be communicated to the medical technology coordinator within 30 days after publication | O Yes O No | No
7.5 | The service manual is to be supplied digitally both in Dutch and / or in English. | O Yes O No | No
7.6 | The irradiator and all peripherals must be covered by a minimum 2-year warranty period which includes (preventive) maintenance, calibrations to regulatory requirements, all breakdown and emergency callouts, all labour costs, replacement parts, travel and subsistence, telephone help desk provision and IT support (where relevant). The maintenance service visits must be completed within periods during the warranty period. | O Yes O No | No
7.7 | The Tenderer must provide ad-hoc repairs as part of the 2 year warranty period and for the duration of the framework agreement . | O Yes O No | No
7.8 | The Tenderer must provide a point of contact (e-mail, Whatsapp, telephone) with a response time within 24 hours, with the exception of weekends and holidays. | O Yes O No | No
7.9 | The Tenderer is to provide Dutch or English speaking representation who can also be contacted in the event of malfunctions and other problems with the equipment offered. | O Yes O No | No
7.10 | The Tenderer must diagnose faults within 1 working days after reporting from a failure/malfunction. | O Yes O No | No
7.11 | The Tenderer must complete repairs and ensure the machine is fully functional within 3 working days after reporting the failure. | O Yes O No | No
7.12 | The Tenderer gives a statement of maintenance and prices thereof, (based on appendix 4c Wibaz SSO version 3) of full maintenance by the company, during 10 years after purchase. | O Yes O No | No
7.13 | The Tenderer provides a brief description of the maintenance that must be carried out for each part of the system. | O Yes O No | No
7.14 | The TCO price quoted by the Tenderer in the Quotation Form is an fixed quotation which is to include full maintenance and costs of all necessary equipment, parts, accessories, licenses, support and installation. | O Yes O No | No
7.15 | The Tenderer must provide a plan to Erasmus MC that details how repairs and planned preventative maintenance will be effected. This must include: - Escalation plan mapped out with respective telephone numbers and email addresses The Tenderer must accommodate Erasmus MC's required planned preventative maintenance (PPM) within 1 month of the intended maintenance expiry date. The Tenderer must carry out Planned Preventive Maintanance (PPM) | O Yes O No | Yes | Statement from manufacturer indicating PPM frequency. Statement from Tenderer that determines methodology for justifying increase/decrease in PPM frequency. Planned Preventive Maintenance plan
7.16 | The Tenderer must provide Erasmus MC with an list of crtitical spare parts | O Yes O No | No
7.17 | The Tenderer must provide Erasmus MC with a report detailing expected replacement cycles for parts including those specified to last 5 years in 2.16, 2.17 and 2.24 including costs and quantity. | O Yes O No | Yes | Tenderer report
7.18 | Each X-ray tube used in the irradiator must have a minimum operating life of 5 years. | O Yes O No | No
7.19 | The irradiator power supply must have a minimum operating life of 5 years. | O Yes O No | No
7.20 | Canisters must be composed of a material that does not degrade at an excessive rate as a result of exposure during normal operation of the irradiator. | O Yes O No | No
7.21 | The Tenderer must state the minimum lifetime of canisters against the expected Erasmus MC usage under 2.14. | O Yes O No | No
7.22 | All mechanical components of the irradiator and canisters within the irradiation field must either be composed of materials that resist degradation under normal operation of the irradiator or have adequate shielding so that no parts become degraded requiring replacement more often than every 5 years. | O Yes O No | No
7.23 | The Tenderer must provide a full spare set of canisters with each irradiator purchased including necessary canisters for dose mapping. | O Yes O No | No
7.24 | The Tenderer's engineers must be trained and qualified to all relevant standards required to perform work on irradiators. The Tenderer must ensure the PPM is carried out by an engineer trained per of this specification. | O Yes O No | No
7.25 | The Tenderer must define the recommended working life (RWL) for the irradiator. This must be equal to or exceed 10 years. | O Yes O No | No
7.26 | The Tenderer acknowledges and agrees to be able to (continue to) (a) maintain the X-Ray irradiator and (b) supply spare parts for at least ten (10) years after succesful installation and fully operational X-Ray irradiatior including peripheral | O Yes O No | No
7.27 | Tenderer guarantees an uptime for the X-Ray irradiator of 98 percent during each period of twelve (12) months. Downtime hours begin at the time when Erasmus MC reports a fault causing downtime during the business hours and end of the time when the X-Ray irradiator is once again actually available for use by Erasmus MC. | O Yes O No | No
7.28 | The Tenderer must ensure the PPM is supported by a log book or service record compliant with 9.1 of this specification. The Tenderer must ensure the PPM is supported by a log book or service record compliant wiith 9.3 | O Yes O No | Yes | The Tenderer must provide a sample log book.
7.29 | The Tenderer must ensure the PPM includes calibration of both the timing mechanism and the back up timer or timers. | O Yes O No | No
7.30 | The Tenderer must ensure the PPM includes dose rate measurement and mapping (compliant with ISO/ASTM 51939:2016/2017) inclusive of all costs (including repeat testing of any failed test results. Erasmus MC will advise the Tenderer of the due date. The Tenderer can perform the dose map up to 2 weeks prior to the due date but must not exceed the due date. | O Yes O No | No
7.31 | The Tenderer must ensure the absorbed dose rate measurement is performed immediately before and after any work being carried out in order to provide an 'as found' reading. | O Yes O No | No
7.32 | The Tenderer must ensure a new dose measurement and mapping must be performed following any change to the tube, geometry or mechanical alteration of the irradiator which may affect operation or any other parameter identified that may impact dose rate or distribution. | O Yes O No | No
7.33 | The Tenderer must ensure that if adjustment(s) are made to the irradiator that may affect the dose rate, that new dose measurements and mapping are performed in accordance with ISO/ASTM 51939:2016/2017. | O Yes O No | No
7.34 | After (Planned Preventive) Maintainance the X-RAY irradiator is delivered conform specification and a 'Fit for use" specification will be given | O Yes O No | No
Section 8 | Documentation required for preventative and reactive maintenance
Note | All PPM visits will be performed under supervision of an site representive.
8.1 | A log book or service record is mandatory, which specifies; - make and model of the irradiator - serial number of the irradiator or other unique identifier (per 1.2) - location of the irradiator (which site) - the date of the PPM service - the name and signature of the service engineer | O Yes O No | Yes | The Tenderer must provide a sample log book
8.2 | A protocol is mandatory, which specifies; - the procedure used for dose measurement and mapping - the corrective action to be taken for values outside of this specification | O Yes O No | Yes | Sample protocol
8.3 | A worksheet and checklist is mandatory, which specifies the items covered in the PPM visit, to include: - identification details and calibration status of the test equipment and instruments - inspection of irradiator, check components, safety features, engineered controls, warning devices and accessories - adjusting X-ray tube as necessary - checking, lubricating and if necessary adjusting drive mechanism - ensuring proper operation and calibration of timing device and backup timer - specification, acceptable tolerance and measured value or condition for each test - 'as found', 'as left' measured values for critical parameters - checking and adjusting all necessary controls - checking that irradiation levels around equipment are normal and within regulatory limits and as low as reasonably practical - checking power supply, drum rotation and operator controls, switches, GUI where relevant - details of replacement of any defective parts - any corrective action taken and outcome | O Yes O No | Yes | The Tenderer must provide a sample checklist.
8.4 | A service record for the irradiator with a copy of the PPM checklist must be presented to the site representative at completion of visit for verification and signature. | O Yes O No | Yes | The Tenderer must provide a sample service record
8.5 | A written report of the dose mapping results must be provided to the Validated Person of Erasmus MC within 2 days of the service visit. | O Yes O No | No
Section 9 | General
9.1 |  | O Yes O No | No
9.2 | After awarding, the irridator will be delivered, installed and ready for use within 2 months at the latest. | O Yes O No | No
9.3 | The Tenderer must support the Erasmus MC in the correct implementation of the device at the location of the Erasmus MC. | O Yes O No | No
9.4 | There are no costs associated with submitting a Tender and you do not derive any rights (now and in the future) from the fact that you may submit a Tender now. | O Yes O No | No
9.5 | You agree to the entire method described and all conditions as stated in the accompanying Quotation Guidelines, the other tender documents and the appendices. | O Yes O No | No
9.6 | You agree with the agreements and the provisions contained therein, which have been added as appendices to the Tender Guidelines. | O Yes O No | No
9.7 | The provisions of the agreements as included in the appendices to the Quotation Guidelines will prevail. For situations for which the provisions of these (framework) agreements do not provide, the NFU General Purchase Conditions, version 1 September 2016, apply. These general purchasing conditions are attached as an appendix to the Offer Guidelines. | O Yes O No | No
9.8 | Do you accept the applicability of these purchase conditions under the explicit exclusion of your own general terms and conditions? | O Yes O No | No
9.9 | The final agreements resulting from this Quotation Guidance will be the result of all information and documentation provided. The Tenderer guarantees that his organization has the required capacities, skills and resources to be able to (continue to) comply with the Schedule of Requirements and Wishes as stated in this Tender Guide. | O Yes O No | No
9.10 | Erasmus MC must be informed immediately of any substantial change in the Tenderer's situation in order to meet these requirements. | O Yes O No | No
9.11 | The Tenderer hereby declares to have taken note of the (draft) purchase agreement and associated conditions (see Appendix 4) that apply to this assignment and indicates by submitting this PoR that the Tenderer agrees to the contract conditions stated herein. | O Yes O No | No
9.12 | The Tenderer hereby declares to guarantee the correctness and completeness of all data supplied. The Tenderer is aware that, if it appears at a later stage that incorrect and / or incomplete information has been provided, the Tenderer can be excluded from further participation, or agreements already made can be canceled without legal consequences for Erasmus MC. | O Yes O No | No
9.13 | Are the conditions offered by you in Euros? | O Yes O No | No
9.14 | Have you completely filled in the price table as included in Appendix 7 and enclosed it with your offer? | O Yes O No | No
Company name:
Legally valid signature: