## Programma van Eisen
PvE: microCT
| Question | Yes/No
CT System
1 | Gantry based with X-ray tube and flat panel detector
2 | 360-degree continuous gantry rotation
3 | Adjustable X-ray to sample to detector distances without recalibration
4 | Cone beam microfocus X-ray source
5 | High speed flat panel detector (up to 100 fps)
Installation
6 | No special infrastructure required (size < 1500 x 1000 x 1000 mm (H, B, L) weight < 400 kg)
7 | Standard electricity power plug 230 V, 5A
8 | Built-in shielding and interlocks to meet cabinet X-ray systems 21 CFR 1020.40
9 | Noise level below 60 dB measured within 50 cm to minimize occupational noise induced hearing loss
Control unit and user interface
10 | Integrated image acquisition and reconstruction computer with local area network connection
11 | GUI to enable image acquisition, viewing, and analysis
Scan protocols and reconstruction techniques
Scan modes
12 | Preset protocols for high resolution and functional cardio-respiratory imaging
13 | Automated scan/stitching for full transaxial scanning
Protocols in general
14 | Possibility to save user defined protocols
15 | Possibility of using changeable X-ray Al/Cu filters
16 | Calibrated and automated HU conversion included
Ex vivo sample imaging capability
17 | High resolution mode (<5 micron)
Reconstruction
18 | realtime GPU-based cone-beam filtered backprojection (<30 seconds)
Output formats
19 | BMP
20 | TIF
21 | JPEG
22 | PNG
23 | AVI
24 | DICOM
Service and warranty
26 | Warranty period is at least twelve months and starts after use of the system
27 | The entire system is ‘State-of-the-Art’ at delivery (installation). By ‘State-of-the-Art’, Erasmus MC means that all component parts of the installed system – both hardware and software – as well as the system as a whole are of the highest level of technical development available, including any advances that become commercially available between the bid/offer being submitted and the start of installation. By submitting a bid/offer, a company obligates itself to deliver the most advanced version of the technology that is commercially available at the start of installation, even if this is being marketed as a different system rather than as an advanced model of the offered system.
Other
28 | Training of users according to the needs of Erasmus MC is included (see further specifcation below)
29 | De company making a bid will provide training covering all aspects of the microCT system.
30 | The bid must include training of two application specialists in the use of all aspects of the microCT system.
31 | All training sessions shall be given by application specialists and/or scientific personnel.
32 | All costs for training shall be paid by the bidding company, including transportation and accommodation costs for employees of the Erasmus MC as well as for company employees.
33 | The bidding company will provide a basic training for the equipment and training for advanced applications.
34 | The bidding company will add an extensive program for basic and advanced training to the bid.
35 | Basic training shall include all aspects of the microCT system including software and hardware applications, safety considerations, and frequently occurring problems and artefacts.
36 | During or shortly after installation, the bidding company will train 2 or more application specialists to perform quality controls and to monitor output . How to carry out phantom studies will be explicitly addressed. These application specialists will also receive any needed information or instructions to enable them to use 3rd party hardware and software in conjunction with the microCT system
37 | A minimum of 3 application specialists and expert users shall be trained in the use of the postprocessing and analysis software for image analysis, such training to be at least 1 full day in length. The training shall take place before, during or immediately after installation of the equipment and shall be planned in consultation with the Erasmus MC.
38 | Complete system complies with CE and TUV standards
39 | ISO compliant/certified (include number)
40 | Scientific collaboration: Active attitude towards cooperation in the development of application possibilities of the technology; including a quarterly update on new developments in the field and the company's roadmap for new developments.
Firmanaam: Rechtsgeldige ondertekening:

## Programma van Wensen
PvW: microCT
[6 empty] | reply supplier | Points: | Score= Points * weight
| Question | inadequate= 0 points | adequate= 1 point | good= 2 point | weight | (also please explain)
CT System
System start-up & system check
1 | Total time from stand-by mode to scanning (min) | > 3 | 1-3 | <1 | 3
2 | X-ray tube warm-up time from 24 hrs non-use (min) | > 30 | 15-30 | < 15 | 3
3 | System calibration frequency | ≥1 time per day | ≥1 time per week | ≤1 time per month | 3
Gantry
4 | Max gantry 360-degree rotatation speed (sec/revolution) | > 10 | 5-10 | < 5 | 10
5 | Min gantry 360-degree rotation speed (min/revolution) | < 25 | 25-50 | > 50 | 10
6 | Max bore hole diameter (mm) | < 100 | 100-150 | > 150 | 10
7 | Removable and cleanable gantry covers | N | x | Y | 10
X-ray tube
8 | Focal spot size at 4 W (µm) | > 10 | 5-10 | ≤ 5 | 5
9 | Minimum tube voltage (kV) | > 40 | 30-40 | ≤ 20 | 3
10 | Maximum tube voltage (kV) | < 60 | 60-90 | ≥ 90 | 3
11 | Maximum tube current (µA) | < 100 | 100-150 | ≥ 200 | 3
12 | Number of user removable X-ray Al/Cu filters | ≤ 2 | 3-5 | ≥ 6 | 3
flat panel detector
13 | detector size (megapixels) | < 6 | 6 | > 6 | 8
14 | Max detector frame rate (frames per sec) | < 50 | 50-100 | > 100 | 10
stages/animal beds
15 | Heated with temperature control | N | x | Y | 10
16 | Adjustable position in X, Y, Z with real-time fluoroscopy viewing | N | x | Y | 5
17 | Scannable length with table transport (mm) | < 100 | 100-200 | > 200 | 5
18 | Max sample/animal weight (g) | < 2000 | 2000-4000 | ≥ 5000 | 3
Control unit and user interface/ vervangen door multimodality
Communication-animal monitoring
19 | Direct animal observation though a X-ray protective transparent window | N | x | Y | 5
20 | real time animal/sample fluoroscopic display | N | x | Y | 5
21 | Body temperature, respiratory and ECG monitoring | N | x | Y | 10
Acquisition and Reconstruction Computer Specifications
22 | Windows 10 (64 bit) operating system | N | x | Y | 3
23 | Intel Xeon Processor | 1620 | 2104 | 2123 | 3
24 | RAM (GB) | 8 | 16 | 32 | 3
25 | Reconstruction GPU NVIDIA with CUDA cores | GTX770 | GTX980 | GTX1080 | 3
26 | Computer integration into the scanner cabinet | N | x | Y | 3
Local storage & Archiving
27 | Computer data storage capacity (Gbyte) | < 4000 | 4000-7000 | ≥8000 | 3
28 | Network-attached storage (NAS) capacity (Gbyte) | N | 32000-64000 | >64000 | 3
29 | On DVD-RW | N | x | Y | 3
30 | Support of data storage on external USB devices | N | usb 2.0 | usb 3.0 | 3
31 | Support of data storage via local area network (LAN) adapter | N | i211 | i219 | 3
Multi-user and user-friendliness facilitation
32 | Ability of creating, importing, exporting, moving user-specific databases | N | x | Y | 3
33 | Availability of user pre-determined scanning modes | N | x | Y | 3
34 | Animal orientation and stage/animal bed control | N | x | Y | 3
35 | Rapid image acquisition (sec) * | > 60 | 5-60 | < 5 | 3
36 | Rapid image reconstruction with image matrix ≥ 512x512x512 after acquisition (sec) | > 60 | 15 - 60 | < 15 | 3
37 | Rapid raw projection viewing after acquisition (sec) | > 60 | 5-60 | < 5 | 3
38 | Rapid viewing of axial, sagittal, and coronal images after reconstruction (sec) | > 60 | 5-60 | < 5 | 3
39 | Rapid 3D rendering after reconstruction (sec) | > 60 | 5-60 | < 5 | 3
40 | Automated intrinsic retrospective cardio-respiratory gating | N | x | Y | 3
41 | Support for BMP, TIF, PNG, RAW, DICOM data formats | N | x | Y | 3
* Images must have good CNR and spatial resolution to allow qualitative and quantitative image interpretation
Accessories
42 | Ex vivo sample holder for high resolution imaging | N | x | Y | 3
43 | In vivo animal holder mice | N | x | y | 3
44 | In vivo animal holder rat | N | x | y | 3
45 | In vivo animal holder rabbit | N | x | y | 3
46 | FMT cassette holder compatibility | N | x | Y | 3
47 | IVIS cassette holder compatibility | N | x | Y | 3
48 | Isoflurane anaesthesia system | N | x | Y | 3
49 | Injection pump | N | x | Y | 3
50 | HU and BMD calibration phantoms | N | x | Y | 3
Scan protocols and reconstruction techniques
Reconstruction
51 | Reconstruction tasks are executed in the background without delaying image acquisitions | N | x | Y | 7
52 | Auto-viewing upon completion of reconstruction | N | x | Y | 7
53 | Ability of reconstructing portions of scanned volumes | N | x | Y | 10
54 | Ring artifact reduction | N | x | Y | 3
Dose and dose reduction techniques
55 | Dose display for scan modes | N | x | Y | 3
56 | Low dose predetermined scanning modes | N | x | Y | 3
57 | Technique to avoid additional dose due to excessive live viewing | N | x | Y | 3
58 | Real-time dose display during image acquisition | N | x | Y | 3
CT Angiography
59 | Contrast-media monitoring in fluoroscopy mode | N | x | Y | 10
60 | Contrast-media recording in fluoroscopy mode | N | x | Y | 3
Automated intrinsic retrospective respiratory and cardiac 2-phase gating
61 | Temporal resolution, detector speed (ms) | >30 | 20-30 | <20 | 3
62 | Image acquisition time (min) | >10 | 5-10 | <5 | 3
63 | Real-time respiratory cycle length and breaths per minute display | N | x | Y | 3
64 | Gating region positioning | N | x | Y | 3
65 | Optimization of respiratory gating parameters | N | x | Y | 3
66 | Optimization of cardiac gating parameters | N | x | Y | 3
67 | Time to respiratory 2-phase reconstruction with image matrix ≥ 512x512x512 (min) | >5 | 1-5 | <1 | 1
68 | Time to cardiac 2-phase reconstruction with image matrix ≥ 512x512x512 (min) | >5 | 1-5 | <1 | 1
High Resolution Imaging
69 | Ex vivo sample imaging resolution (microns), 10%MTF | >20 | 20-15 | <15 | 2
70 | Ex vivo sample imaging matrix size (number of pixels) | <3000 | 3000-7000 | 8000 | 2
71 | In vivo imaging resolution (microns), 10% MTF | >40 | 40-30 | <30 | 2
72 | In vivo imaging matrix size (number of pixels) | <3000 | 3000-7000 | 8000 | 2
Service
73 | Vendor acquisition and reconstruction software upgrades for the first 5 years included in purchase price. | N | x | Y | 3
74 | Remote system monitoring for potential component failure prediction | N | x | Y | 3
| TOTAL SCORE | [6 empty] | 0
Firmanaam: Rechtsgeldige ondertekening:

## Informatieblad
Information sheet microCT
| Question | Answer supplier (also please explain)
Accessoires en opties
1 | Describe all options, with prices
2 | Describe all available accessories, with prices
Overig
3 | Expected X-ray tube life (in scan seconds)
4 | Total heat output?
5 | Does an IHE integration statement apply to this system? (if yes which one.)
6 | Attach data sheet and brochure of the system.
7 | As a supplier, do you have a mission statement with regard to corporate social responsibility? If so, specify how this mission statement is translated into objective objectives. Please attach the Mission Statement
8 | As a supplier, do you have specific objective (s) to reduce the energy consumption of the relevant equipment?
9 | Reference sites
Firmanaam: Rechtsgeldige ondertekening: