## Appendix A05
Appendix A06, Conformity list 'Programme of Requirements' Guidelines for the multiple limited tender for an E-Beam Evaporator, ref. TNO 2021 FPL-INK 114 dated 26-08-2021
By the following Programme of Requirements TNO will assess if Tenderer meets the relevant requirements as laid down in Section 6 and 8. The Tenderer must indicate for each item, the conformity for the requirements in the table below.
If to any of the requirements, as stated in section 8, are not met by Tenderer, he will be excluded for any further assessment. By answering a question with “compliant” will mean that in the event of awarding the contract, non-conformity to the requirement shall count as failure to meet an obligation from the Agreement to be concluded.
Requirement | Decription | Tenderer is compliant to the requirement | Reference in Tenderer's documentation where compliances with the requirement is described | Remark
Requirement 8.1.1. | The e-beam evaporation system must be a newly built first-hand system or a retrofitted chamber with all critical parts replaced and new: - moving parts and sealings - electronics - Software - Pumping system | Click the drop-down box to select
Requirement 8.1.2 | The evaporator comprises a process chamber, housing e-gun(s). | Click the drop-down box to select
Requirement 8.1.3 | The evaporator comprises a control rack with a windows-controlled PC based user interface. | Click the drop-down box to select
Requirement 8.1.4 | The evaporator must be CE certified and be compatible with Dutch work environment safety regulations (ARBO). | Click the drop-down box to select
Requirement 8.1.5 | The system will be placed in a cleanroom and should be cleanroom class 10.000 (ISO class 7) compatible. | Click the drop-down box to select
Requirement 8.1.6 | Footprint of the tool should be smaller than: 2x3 m2. Please indicate floor space | Click the drop-down box to select
Requirement 8.1.7 | High voltage discharge hook is provided by the tenderer | Click the drop-down box to select
Requirement 8.2.1 | The process chamber is a UHV compatible stainless-steel chamber and metal parts, like screws, bolts also from stainless steel. | Click the drop-down box to select
Requirement 8.2.2 | The process chamber allows access to various parts (e-gun, crucibles, shields) for maintenance and processing. | Click the drop-down box to select
Requirement 8.2.3 | The chamber has a view port to observe the hearth of the E-gun. This view port is equipped with a shutter and a polarization filter (or something similar) for eye protection | Click the drop-down box to select
Requirement 8.2.4 | The chamber must be equipped with at least one spare blind flange outside the deposition area (at least DN40). | Click the drop-down box to select
Requirement 8.2.5 | The chamber is equipped with a quarter inch gas line connection, which enables the addition of oxygen or argon to the evaporation process later, if needed. | Click the drop-down box to select
Requirement 8.2.7 | Stainless steel shields are present to cover the chamber walls. | Click the drop-down box to select
Requirement 8.2.8 | The shields must withstand the cleaning procedure by sandblasting | Click the drop-down box to select
Requirement 8.2.9 | The shields can be mounted and unmounted by TNO engineer within 60 min. | Click the drop-down box to select
Requirement 8.2.10 | (at least one) set of spare shields are supplied. | Click the drop-down box to select
Requirement 8.2.11 | The chamber must accommodate the sample holder(s), specified in section 8.5 | Click the drop-down box to select
Requirement 8.3.12 | The chamber is equipped with a shutter on the e-gun/source, which can be opened and closed in manual and automatic mode. | Click the drop-down box to select
Requirement 8.2.13 | The process chamber houses a crystal monitor system to measure layer thickness from 1 Angstrom to the micron level. | Click the drop-down box to select
Requirement 8.2.14 | The thickness / rate meter is mounted such that the deposition rate can be tuned with the deposition shutter closed. | Click the drop-down box to select
Requirement 8.2.15 | Shutter control is possible on a recipe basis and the shutter control is interfaced with the layer thickness sensor. | Click the drop-down box to select
Requirement 8.3.1 | The base pressure of the system (after 12 hours of pumping) should be lower than 1.0 x 10-7 mbar | Click the drop-down box to select
Requirement 8.3.2 | Pump down time of process chamber to operational pressure (≤1 x10-6 mbar) must not take longer than 60 min. | Click the drop-down box to select
Requirement 8.3.4 | The primary pump is a dry pump. No pumps based on oil to be used | Click the drop-down box to select
Requirement 8.3.5 | Pressure gauges are present to monitor pump down in all vacuum ranges. | Click the drop-down box to select
Requirement 8.3.6 | Atmospheric pressure detection or indication is present. | Click the drop-down box to select
Requirement 8.4.1. | The process chamber houses at least one E-gun with at least 4 pockets. | Click the drop-down box to select
Requirement 8.4.2 | The system must be ready to be extended for a co-deposition process (second e-gun) for future upgrade | Click the drop-down box to select
Requirement 8.4.4 | The E-gun has an arrangement so that the hot cathode does not contaminate the process. A common arrangement is an (approx.) 270° deflecting gun. | Click the drop-down box to select
Requirement 8.4.5 | Evaporation of Al2O3 (10 – 1000 nm), Fe (1 - 50nm), Au (20 - 200 nm) without refilling must be possible, with a minimum pocket size of 6 cc | Click the drop-down box to select
Requirement 8.4.6 | The pocket size should be suitable for the above-mentioned materials and thicknesses | Click the drop-down box to select
Requirement 8.4.7 | E-gun and pockets should not contaminate the processing materials due to heating. An appropriate cooling system must be applied. | Click the drop-down box to select
Requirement 8.4.8 | Pocket indexing must be automated via software. | Click the drop-down box to select
Requirement 8.4.9 | Spare parts for E-gun (like filaments and beam formers) are common parts and easy to order. | Click the drop-down box to select
Requirement 8.4.10 | The system is equipped with a pre-configured flexible beam sweep system. | Click the drop-down box to select
Requirement 8.4.11 | The power of E-gun and power supply should be matching. | Click the drop-down box to select
Requirement 8.4.12 | The power supply can provide sufficient power to evaporate metal oxides and metals with low vapor pressure (mentioned in section 8.6). Power is precisely controllable. Pmax ≥ 6 kW. | Click the drop-down box to select
Requirement 8.5.1 | The substrate holder can accommodate at least four 6” round wafers (thickness: 0.7mm – 2.1mm). | Click the drop-down box to select
Requirement 8.5.2 | An additional substrate holder can accommodate three152mm x 152mm substrates (from 0.7 up to 2.1mm thickness) | Click the drop-down box to select
Requirement 8.5.4 | The sample holder can be continuously rotated, and static mode evaporation is also possible | Click the drop-down box to select
Requirement 8.6.1 | Deposition of various metals and metal oxides (examples: Al2O3, MoOx, Ti, Pd, Au, Pt, Fe) with full process parameters control. | Click the drop-down box to select
Requirement 8.6.2 | The evaporator must produce coatings (100 nm Al2O3 and MoOx, 10 nm Fe and Ti) on all of the (at least) four 6” wafers in a run with a thickness variation within <6 % (maximum thickness – minimum thickness) / average thickness). | Click the drop-down box to select
Requirement 8.6.4 | Deposition rates can be tuned for high-rate (10 Å/s for 100 nm Al2O3) as well as low rate (≤1 Å/s for 1 nm Fe) metals and metal oxides. | Click the drop-down box to select
Requirement 8.6.5 | Sequential depositions of multilayers can be done in situ (without breaking the vacuum) | Click the drop-down box to select
Requirement 8.6.6 | The evaporation process can be performed in automatic mode (closed e-gun/crystal-controller loop) as well as manual control. | Click the drop-down box to select
Requirement 8.6.7 | The tenderer should provide a starting recipe for the depositions of materials (Al2O3, MoOx Ti, Pd, Pt, Fe). | Click the drop-down box to select
Requirement | Sample cleaning prior to deposition may become of future importance and the system then needs to be upgraded. It must be possible to extend the system with an extra cleaning system (ion source or RF cleaning system). Provide best option and price. | Click the drop-down box to select
Requirement 8.7.1. | The system has a control rack, containing doors or removable panels on the sides and back. | Click the drop-down box to select
Requirement 8.7.2 | The control rack comprises a computer screen enabling easy operation by persons of different skill level. This screen must give an overview of the system status, and if the system status is ‘ready for processing’, a recipe-based interface for automatic deposition. The control system defines different levels, e.g.: - ‘student/user’ level to perform basic deposition - ‘researcher/engineer’ level for process development and configuration editing with access to manual operations and a - service/tool owner’ level with full access to the hardware. | Click the drop-down box to select
Requirement 8.7.3 | The control system includes a safety / fault detection subsystem to prevent operations that would risk damaging the system (at the lower operator levels). | Click the drop-down box to select
Requirement 8.7.4 | The control system includes an interlock system that prevents turning on any high voltage when the system is open / the high voltage is exposed. | Click the drop-down box to select
Requirement 8.7.5 | The control system allows for the creation and execution of recipes. | Click the drop-down box to select
Requirements 8.7.7 | The system contains different interlocks that prevents against cooling water supply failures or other failures. | Click the drop-down box to select
Requirement 8.7.8 | Tenderer makes software-bug fixing and necessary software updates available, so that ‘current’ operating systems are supported. These updates are free of charge for the first 5 years. TNO owns the necessary licenses for common operating systems. Such updates are necessary so that the system can be connected to the network. | Click the drop-down box to select
Requirement 8.8.1. | The evaporation system must have a warranty period of at least 24 months. | Click the drop-down box to select
Requirement 8.8.2 | Confirmation of receipt of maintenance within warranty – or service request or complaint within warranty: 24 hours* | Click the drop-down box to select
Requirement 8.8.5 | Operators training program: Tenderer should propose a training program for the machine operators, covering at least: - start-up of the machine, - usage of the software, - safety instructions, - cleaning of the tool. - maintenance that can be carried out by TNO service engineer/tool owner and the operators All the offered items shall be quoted and specified/described separately in appendix A04: Schedule of Prices and Costs." | Click the drop-down box to select
Requirement 8.8.6 | Delivery includes transport of the evaporator, unpacking and installation at the site (HTC 21, 5656 AE, Eindhoven, the Netherlands). | Click the drop-down box to select
Requirement 8.8.7 | The evaporator must be delivered together with all documentation required to operate and maintain the apparatus. This includes a user manual, a system level mechanical drawing, electrical schematics and technical documentation as well as documentation about the control system. The documentation must be supplied in the English language, and, where possible, (also) in digital format. | Click the drop-down box to select
Requirement 8.8.8 | Delivery, including approved SAT, must take place within/before 7 months after agreement and signature of a contract between the parties, exact delivery date should be agreed upon by both parties Note: not achieving the offered delivery time will lead to a penalty discount of 20% of the total order amount and shall be immediately due and payable and eligible for setoff and will replace clause 15.6 of the General Purchasing Conditions TNO - 2014 | Click the drop-down box to select
Requirement 8.8.10 | Tenderer should provide a list of the most critical parts of the system together with argumentation why Tenderer considers these critical parts, as well as delivery time of such critical parts. Tenderer should also provide list of recommended spare parts and consumables to have on stock at TNO. | Click the drop-down box to select
Requirement 8.8.11 | Most service can be done by TNO when supplied with required technical information. The tenderer maintains a telephone number for answering technical questions. | Click the drop-down box to select

## validatie (hidden)
Click the drop-down box to select
Yes, compliant to requirement
No, not compliant to requirement
note: dit blad 'verbergen' indien je dit gaat delen intern/extern (behoudens aan inkoop collega's)