## Appendix A05
Appendix A05, Conformity list 'Programme of Requirements' Guidelines for the European Public Tender for a continuous furnace system, ref. TNO 2021 FPL-INK 151, dated 10-11-2021
By the following Programme of Requirements TNO will assess if Tenderer meets the relevant requirements as laid down in Section 6 and 8. The Tenderer must indicate for each item, the conformity for the requirements in the table below.
If to any of the requirements, as stated in section 8, are not met by Tenderer, he will be excluded for any further assessment. By answering a question with “compliant” will mean that in the event of awarding the contract, non-conformity to the requirement shall count as failure to meet an obligation from the Agreement to be concluded.
Requirement | Decription | Tenderer is compliant to the requirement | Reference in Tenderer's documentation where compliances with the requirement is described | Remark
Requirement 8.1.1 | The continuous furnace system including all its components, should not consume more energy than 180A at 400V at peak capacity. | Click the drop-down box to select | (mandatory field)
Requirement 8.1.3 | The continuous furnace system has be possibility to anneal float glass with the following dimensions: • 1000mm X 1000 mm X 2.0 - 6.0 mm* • 500mm X 500mm X 2.0 - 6.0 mm* • 330mm X 450 mm X 2.0 - 6.0 mm* * thickness can vary between 2.0-6.0 mm | Click the drop-down box to select | (mandatory field)
Requirement 8.1.5 | The continuous furnace system should not have a larger footprint than 11 (l) X 2.5m (w) X 2.0 (h) (including a safe loading and safe unloading part of furnace) If the continuous furnace exceeds this dimensions, it cannot be placed within our infrastructure. Please describe how the safe loading and unloading of the furnace will be implemented/achieved. The entire answer should be no more than two (2) A4 page. If the answer exceeds two (2) A4 pages, the additional pages will be disregarded. | Click the drop-down box to select | (mandatory field)
Requirement 8.1.7 | Nitrogen consumption needed to operate the system under normal operating conditions should not exceed 50m3/h. | Click the drop-down box to select | (mandatory field)
Requirement 8.1.9 | If other utilities and/ or supporting apparatuses are required in addition to the ones specified above they should be listed, With the required flows, sizes and/or dimensions | Click the drop-down box to select | (mandatory field)
Requirement 8.1.10 | The floor loading should not exceed 500kg/m2. The weight of the continuous furnace system can be distributed over a larger area. Please describe your proposed solution. The entire answer should be no more than three (3) A4 pages. If the answer exceeds three (3) A4 pages, the additional pages will be disregarded. | Click the drop-down box to select | (mandatory field)
Requirement 8.1.11 | For installation: Individual segments should fit through a door opening of 1.79m in width, 2.0m height and 2.5m in length. | Click the drop-down box to select | (mandatory field)
Requirement 8.1.12 | If an installation box / E-cabinet is used it should not exceed the following dimensions: 3.0m (length) x 0.75m (width) x 2.5m (height). Requirement 8.1.11 applies also for the E-cabinet. | Click the drop-down box to select | (mandatory field)
Requirement 8.1.14 | Installation of all components of the furnace has to be carried out by the supplier; in a dialogue with TNO The assembly of the continuous furnace with all it's components is the responsibility of the manufacturer at location. | Click the drop-down box to select | (mandatory field)
Requirement 8.2.1 | Homogeneity of the temperature distribution on the surface of the glass should not vary more than 5⁰C from the programmed temperatures during all stages of the curing cycle, except in cooling section H when the temperature is below 200⁰C. (determination during FAT by attaching 9-12 thermocouples equally divided over the surface of the glass). During FAT a 1000mm X 1000mm X 2.0-6.0mm optiwhite, low-iron, soda-lime floatglass. will be used. | Click the drop-down box to select | (mandatory field)
Requirement 8.2.2 | Repeatability; during continuous operation of the oven all programmed and desired parameters should be achieved and be reproducible. | Click the drop-down box to select | (mandatory field)
Requirement 8.2.4 | The atmosphere within the continuous furnace system has to be homogeneous and not deviate more than 5ppm of the set O2 concentration when an O2 concentration (ranging between 10 and 150ppm) is specified in the process (this excludes segment A and B of scheme 1). | Click the drop-down box to select | (mandatory field)
Requirement 8.2.5 | O2 concentration is measurable during operation of the continuous furnace system, with an accuracy ≤1 ppm and should be measured from segment C-G; with read-out. Describe how this preference will be implemented/achieved. The entire answer should be no more than one (1) A4 page. If the answer exceeds one (1) A4 pages, the additional pages will be disregarded. | Click the drop-down box to select | (mandatory field)
Requirement 8.2.6 | Segment A-B from scheme 1; Heating under air from ambient (18-26⁰C) to a set temperature of maximum 300⁰C with a controllable and programmable heating rate of 0.1-0.45⁰C/s with intermittent dwell times. | Click the drop-down box to select | (mandatory field)
Requirement 8.2.7 | Segment C (= transporting from segment B to D), scheme 1; at a set temperature ranging between ambient and ≤ 300⁰C Wherein the time should be shorter than 1 minute; Within this timeframe the atmosphere is changed from oxygen rich to a set atmosphere containing between 10-150ppm of O2 in N2 rich atmosphere (using nitrogen with above-described quality). The O2 should be adjustable (in segment D) to a set ppm at a ranging between 10-150ppm with a tolerance of 5ppm to the set O2-concentration. Requirements 8.2.4 and 8.2.5 do not apply during the duration of segment C. | Click the drop-down box to select | (mandatory field)
Requirement 8.2.9 | Segment D-F from scheme 1; Heating from ambient to a set temperature of maximum 550⁰C at a heating rate programable and measurable between 0.06-0.45⁰C/s. With multiple (> 5) intermittent programable dwell-times in a set O2 atmosphere containing between 10-150ppm of O2 (using nitrogen with above-described quality). The O2 should be adjustable to a set ppm at a ranging between 10-150ppm with a tolerance of 5ppm to the set O2-concentration. | Click the drop-down box to select | (mandatory field)
Requirement 8.2.10 | Segment G from scheme 1: cooling. Samples can be homogenous (requirement 8.2.1) cooled from 550⁰C to 200⁰C under set atmosphere containing between 10-150ppm of O2 in N2 at a cooling rate of at least ≥0.1⁰C/s. Segment H: The cooling from 200⁰C towards ambient conditions can be performed under air (In or outside furnace), whilst maintaining safe operation conditions | Click the drop-down box to select | (mandatory field)
Requirement 8.3.1 | The continuous furnace system must be newly built, first-hand system. | Click the drop-down box to select | (mandatory field)
Requirement 8.3.2 | All electronic components and the continuous furnace system as a whole are CE marked, a declaration of conformity must be provided. | Click the drop-down box to select | (mandatory field)
Requirement 8.3.3 | The operation of all controls are interlocked to prevent damage to the continuous furnace system and for safety of the operator | Click the drop-down box to select | (mandatory field)
Requirement 8.3.4 | All documentation, software commands and read-out must be written in the English or Dutch language. | Click the drop-down box to select | (mandatory field)
Requirement 8.3.5 | Documentation must include at least the following: • Installation documentation (this includes for instance relevant STEP-files of interfaces, fitting diameters, electrical connections, floor load, required space for transportation, EMC fields, hoisting points etc.) • Instruction manual • Maintenance manual including all required technical documentation for Preventive and Corrective Maintenance, including relevant STEP-files and “as built” drawings. • Quick- start guide • P&ID of the system • Safety measures and their implementation • Certificates and manuals for all third-party hardware that is part of the continuous furnace system All the documents should be delivered digitally in English or Dutch and are available during FAT. | Click the drop-down box to select | (mandatory field)
Requirement 8.3.6 | Certificates The supplier shall document all the relevant certificates. TNO requires at least: - Calibration certificates (if applicable) - Specifications and manuals of equipment Delivery of documentation is expected during FAT. | Click the drop-down box to select | (mandatory field)
Requirement 8.3.7 | The warranty on the continuous furnace system and system components has to be at least 24 months as stated in the General Purchasing Conditions TNO (Appendix C03), starting after the SAT (onsite acceptance and test period) of the continuous furnace system by TNO. | Click the drop-down box to select | (mandatory field)
Requirement 8.3.8 | The Tenderer shall maintain an inventory of common spare parts that may otherwise cause system failures. Describe what type of parts will be kept in stock. | Click the drop-down box to select | (mandatory field)
Requirement 8.3.9 | The guaranteed availability of all system parts has to be at least 10 years. | Click the drop-down box to select | (mandatory field)
Requirement 8.3.10 | Instruction of personnel - training A minimum of 3 TNO employees will be trained on the system to operate the system, firstly at delivery and secondly after 3 months of operation. | Click the drop-down box to select | (mandatory field)
Requirement 8.3.11 | Instruction of personnel - maintenance A minimum of 3 TNO employees will be trained on preventive and corrective maintenance. | Click the drop-down box to select | (mandatory field)
Requirement 8.3.12 | The factory acceptance test (FAT) should be carried out at the premises of supplier. | Click the drop-down box to select | (mandatory field)
Requirement 8.3.13 | A test protocol for the factory acceptance test (FAT) will be agreed after Contract Award but before delivery of the continuous furnace system to Brightlands Materials Center (see also paragraph 2.2.23). Tenderer must submit a draft test protocol for the factory acceptance test with its Tender. | Click the drop-down box to select | (mandatory field)
Requirement 8.3.14 | The on-site acceptance test should be carried out at the site of the TNO, after installation and hook-up. | Click the drop-down box to select | (mandatory field)
Requirement 8.3.15 | A test protocol for the on-site acceptance test (SAT) will be agreed during the verification phase of this Tender (see also paragraph 2.2.24). Tenderer must submit a draft test protocol for the factory acceptance test with its Tender. | Click the drop-down box to select | (mandatory field)
Requirement 8.3.16 | EMC directive The system shall be able to operate safely and without interference in a normal environment | Click the drop-down box to select | (mandatory field)
Requirement 8.3.17 | Safety & Ergonomics standards It must be safe to operate the continuous furnace system for a skilled operator without any safety issues. | Click the drop-down box to select | (mandatory field)
Requirement 8.4.1 | Service during warranty period In case of system failure during warranty period TNO can contact the supplier by e-mail or phone on working days. The supplier will respond and propose a solution to TNO within 24 hours (during working days), and guarantees a maximal on-site response time of three (3) working days, for maintenance, troubleshooting, and/or repair in case of system failures. | Click the drop-down box to select | (mandatory field)
Requirement 8.4.2 | Service department Tenderer must have a European service department that has a high level of competence in troubleshooting the supplied equipment, by telephone assistance, or, in the case of complex failures, by a service visit. At least 2 persons in the service organization must be able to carry out such field repairs. To reduce down time, Tenderer keeps a supply of common replacement parts (requirement 8.3.8). Tenderer must be able to supply telephone assistance through a service hotline. Tenderer has a response time of no more than 3 working days outside the warranty period. Provide a description and a reference of the service department so that it is clear that it meets the requirements. | Click the drop-down box to select | (mandatory field)
Requirement 8.6.2 | Operating systems become obsolete and their security updates end. The Tenderer will actively provide TNO with the necessary operating system software updates, to fix, among others, software bugs. These updates shall be provided free of charge to TNO for a period of at least ten (10) years after SAT. | Click the drop-down box to select | (mandatory field)
Requirement 8.7.1 | - 30% at order, and after receipt of a bank guarantee covering 60% of the total price; - 30% after approval by TNO of the Factory Acceptance Test (FAT), in the presence of a TNO employee and approved by TNO Management and TNO Procurement. - 20% after delivery at location TNO Holst Centre of the total system. At this moment TNO will return the bank guarantee. - 20% after installation and commissioning On-site and Site Acceptance Test (SAT) including test periods, approved by TNO Management and TNO Procurement. The Tenderer to whom the Contract will be awarded must provide a bank guarantee to TNO covering at least 60% of the total price. TNO must have the right to draw the bank guarantee in any and all cases where falls short in fulfilling its obligations under the Contract. The validity of the bank guarantee may not expire until approval of the SAT by TNO Management and TNO Procurement. | Click the drop-down box to select | (mandatory field)

## validatie (hidden)
Click the drop-down box to select
Yes, compliant to requirement
No, not compliant to requirement
note: dit blad 'verbergen' indien je dit gaat delen intern/extern (behoudens aan inkoop collega's)