## Appendix A05
Appendix A05, Conformity list 'Programme of Requirements' Guidelines for the European Public Tender for two glovebox systems, ref. TNO 2022 FPL-INK 035, dated 04-03-2022
By the following Programme of Requirements TNO will assess if Tenderer meets the relevant requirements as laid down in Section 6 and 8. The Tenderer must indicate for each item, the conformity for the requirements in the table below.
If to any of the requirements, as stated in section 8, are not met by Tenderer, he will be excluded for any further assessment. By answering a question with “compliant” will mean that in the event of awarding the contract, non-conformity to the requirement shall count as failure to meet an obligation from the Agreement to be concluded.
Requirement | Decription | Tenderer is compliant to the requirement | Reference in Tenderer's documentation where compliances with the requirement is described | Remark
Requirement 8.1.1 | Box size - The required dimensions for the tooling and the assembly space are at least 4300 mm (width) x 725 mm (depth), and the minimum height is 900 mm. The box will be max 5000mm wide, 1000mm deep and 1200mm high. Remark: all measures concern internal dimensions. | Click the drop-down box to select
Requirement 8.1.2 | Seams between panels - The working area is not interrupted by raised seams and the seam widths are kept to a minimum. Especially the bottom panel is smooth as equipment and samples will be moved around it and dirt (which can take up moisture) can build up in seams. Also, the back panel is not interrupted by raised seams, making it possible to place shelves across the seams. The window may be interrupted. | Click the drop-down box to select
Requirement 8.1.3 | Storage - The GB must have shelves at the back of the GB. | Click the drop-down box to select
Requirement 8.1.3.1 | Storage - There are areas where no shelves are allowed. These areas are indicated in the layout in appendix C04. | Click the drop-down box to select
Requirement 8.1.4 | Lighting - The box must have adequate and homogeneous lighting throughout the GB. The industry standard is a lighting strip across the entire top of the box, on the outside in LED or fluorescent lighting. The box will have an equivalent, or better lighting solution. | Click the drop-down box to select
Requirement 8.1.5 | Feedthroughs - The GB will require a number of argon gas feedthroughs, vacuum feedthroughs and 220V wall socket feedthroughs. You will find the GB layout in appendix C04. - Ar feedthrough shall be a 6 mm or ¼” stainless steel tube or a Swagelok connection (bulkhead connection or KF40 flange). - vacuum feedthrough shall be a 12mm or ½” stainless steel tube or Swagelok connection. | Click the drop-down box to select
Requirement 8.1.6 | Glovebox Pressure - The GB has a pressure regulation system. The user can set working pressure, maximum pressure and minimum pressure. | Click the drop-down box to select
Requirement 8.1.7 | Glovebox controls - The GB must have basic operation settings and controls. Basic GB controls can control. | Click the drop-down box to select
Requirement 8.1.7.1 | Glovebox controls - In reference to req. 8.1.6, the GB pressure can be controlled from the GB control panel. | Click the drop-down box to select
Requirement 8.1.7.2 | Glovebox controls - In reference to req. 8.1.4, the GB lighting can be switched on and off from the GB control panel. | Click the drop-down box to select
Requirement 8.1.7.3 | Glovebox controls - In reference to req. 8.3.1, the GB circulation can be switched on and off from the GB control panel, and the automatic regeneration cycle can be started. | Click the drop-down box to select
Requirement 8.1.7.4 | Glovebox controls - In reference to req 8.4.1, the GB audible and visual alarm settings can be controlled from the GB control panel. | Click the drop-down box to select
Requirement 8.1.7.5 | Glovebox controls - In reference to req. 8.2.1, the GB AC and its settings can be controlled from the GB control panel. | Click the drop-down box to select
Requirement 8.1.7.6 | Glovebox controls - In reference to req. 8.4.2, the GB purge can be switched on and off from the GB control panel. | Click the drop-down box to select
Requirement 8.1.7.7 | Glovebox controls - In reference to pref. 8.4.2.1, the GB automatic purge settings can be controlled from the GB control panel. | Click the drop-down box to select
Requirement 8.1.8 | Glovebox controls - The controls will be primary accessible through buttons or a touchscreen outside of the GB. | Click the drop-down box to select
Requirement 8.1.10 | Display - The GB system will have a display. The display will be on the GB or attached to it and will have a maximum diameter of 61 cm (24 inch) and a minimum of 12.7 cm (5 inch). It will not be an extra separate computer or computer screen. | Click the drop-down box to select
Requirement 8.1.10.1 | Display - In relation to pref. 8.1.9, the GB display can graphically represent the logging data. | Click the drop-down box to select
Requirement 8.1.10.2 | Display - In relation to req. 8.3.1, the display can indicate the regeneration status or progress. | Click the drop-down box to select
Requirement 8.1.10.3 | Display - In relation to req. 8.2.1, the display can indicate the AC evacuation status and the evacuation cycle status or progress. | Click the drop-down box to select
Requirement 8.1.10.4 | Display - In relation to reqs. 8.3.1 and 8.4.3, the display can indicate the H2O and O2 levels and alarms. | Click the drop-down box to select
Requirement 8.1.11 | Metric units - The system will use metric units. Remark: preferred units: (m)bar, °C, ltr/min or m3/hr, gr, ppm and kg. | Click the drop-down box to select
Requirement 8.1.12 | Windows - The windows must have a coating to prevent scratches (for instance, a sapphire coated polycarbonate window is sufficient). The GB will not be fitted with bare polymer windows. Remark: To make solutions for slot die coating, particulate matter will be solved in solvent. By the lack of moisture in the GB electrostatic effects are more prevalent. Therefore, we want a coating that has less static charge build-up than PC, like sapphire or glass. | Click the drop-down box to select
Requirement 8.1.13 | Gloves - Standard butyl rubber gloves will be fitted on the GB. | Click the drop-down box to select
Requirement 8.1.14 | Vacuum pump - Any vacuum pumps must be dry scroll pumps, capable of displacing at least 12.5 m3/hr. (Like ACP15G or similar, purge port, corrosion resistance). | Click the drop-down box to select
Requirement 8.2.1 | Antechamber 1.1 – Big antechamber The GB will have one big main antechamber (AC) on the right-hand side. - Size: Internal Diameter > 380mm, depth > 500mm but no deeper that 1000mm. - Temperature: The internal temperature can be set between room temperature and 120C or higher. The temperature homogeneity at any set temperature will be 20% or better. - Slide: The slide will have a ball or needle (or equivalent) bearing system to minimize generation of particles by metal-on-metal friction. - Vacuum lead safety: The AC vacuum lead will be protected from materials like wipes and small parts falling in and damaging the pump/clogging the vacuum lead. - Evacuation and Refill: AC will be used to store and dry the slot die coated samples. Therefore it will be under vacuum for long periods of time and the AC will have the possibility to be evacuated and refilled manually. | Click the drop-down box to select
Requirement 8.2.1.1 | Big antechamber - The big AC must have an automated pressure-controlled evacuation and refill cycle with adjustable parameters: - A lower vacuum limit at which the AC starts refilling the AC. - An upper vacuum limit at which the AC starts evacuating again. - Amount of cycles. - A timer for the evacuation step. - If it takes to long for the AC to reach base pressure, the timer will end the cycle and trigger an (audible) alarm to inform the user that something is wrong. - 3 evacuation and refill cycles must not exceed 20 minutes. The settings are accessible through the control panel. | Click the drop-down box to select
Requirement 8.2.1.2 | Big antechamber - The big AC must have an interlock system so that the AC can only be opened from the inside when it is safe (e.g. when the AC atmosphere is inert). | Click the drop-down box to select
Requirement 8.2.1.4 | Big antechamber – In order to be able to still use the GB and the small AC (see req. 8.2.2), when the big AC is under vacuum (especially for prolonged periods of time), the big AC will have its own vacuum pump. | Click the drop-down box to select
Requirement 8.2.2 | Antechamber 1.2 - Small Antechamber The GB will have a mini antechamber (AC) on the left hand side. - Size: Internal diameter > 150mm, depth > 300mm but no deeper that 600mm. - Temperature: The internal temperature can be set between room temperature and 120C or higher. The temperature homogeneity at any set temperature will be 20% or better. - Slide: AC will have a sliding tray to accommodate introduction of materials into the GB. - Vacuum lead safety: The AC vacuum lead will be protected from materials like wipes and small parts falling in and damaging the pump/clogging the vacuum lead. Remark: AC will be used to dry out consumables. Therefore it will be under vacuum for long periods of time and the AC will have the possibility to be evacuated, refilled and closed manually. | Click the drop-down box to select
Requirement 8.2.3 | The ACs will have their own Ar supply and will not fill with gas from the GB. | Click the drop-down box to select
Requirement 8.2.4 | The ACs are protected against overpressure from the Argon line. | Click the drop-down box to select
Requirement 8.3.1 | Purifier - The GB has a purifier system. | Click the drop-down box to select
Requirement 8.3.1.1 | Purifier - The circulation blower must have a capacity of at least 70 m3 per hour. | Click the drop-down box to select
Requirement 8.3.1.2 | Purifier - The purifier can absorb at least 35L of oxygen and 1250 gr of H2O and will be regenerable. | Click the drop-down box to select
Requirement 8.3.1.3 | Purifier - The purifier can maintain O2 and H2O levels below 1 ppm at a leakrate of 0,03 volume % per hour. Remark: This leakrate is specified to correspond to a large leak like a cut in a glove. | Click the drop-down box to select
Requirement 8.3.2 | Solvents - The GB has a solvent filter system which can remove the solvents listed in the solvent list provided in appendix C05. | Click the drop-down box to select
Requirement 8.3.4 | Nitrogen filter - The GB has a N2 filter, capable of removing N2 from the system at a rate that is enough to maintain <1ppm of N2 in the GB at a leakrate of <0,03 vol% per hour. Nitrogen filter is able to trap at least 500L of N2 (measured value). Remark: This leakrate is specified to correspond to a big leak, like a cut in a glove. | Click the drop-down box to select
Requirement 8.3.4.1 | Nitrogen filter - The Nitrogen filter is protected from exposure to large amounts of N2 through accidental leaks in the GB. Remark: this can be done for example through an automatic box purge or automatic shutdown of circulation over the filter at elevated O2 or H2O levels. | Click the drop-down box to select
Requirement 8.3.5 | Particle filters - The system will have HEPA particle filters to filter incoming and outgoing working gas flow. The HEPA class is H13 or better. | Click the drop-down box to select
Requirement 8.4.1 | Audible alarms - An audible alarm will be available to indicate raised O2 and/or H2O levels. The alarm will not be loud, but clearly audible. Available settings must be at least: silenced on/off, activated/deactivated, the threshold settings are accessible through the controls. | Click the drop-down box to select
Requirement 8.4.2 | Box purge - The glovebox has a purge option that instead of circulating over the purifier supplies the box with inert supply gas. This purge is capable of reducing raised O2 and H2O levels in the box much faster than the purifier. | Click the drop-down box to select
Requirement 8.4.3 | Sensors - The GB system will have an H2O and an O2 sensor capable of a resolution of 1 ppm or better. | Click the drop-down box to select
Requirement 8.5.1 | Box size - The required dimensions for the tooling and the assembly space are at least 2550 mm (width) x 775 mm (depth), and the minimum height is 900 mm. The box will be max 3200mm wide, 1000mm deep and 1200mm high. | Click the drop-down box to select
Requirement 8.5.2 | Seams between panels - The working area is not interrupted by raised seams and the seam widths are kept to a minimum. Especially the bottom panel is smooth as equipment, and samples will be moved around it, and dirt (which can take up moisture) can build up in seams. Also, the back panel is not interrupted by raised seams, making it possible to place shelves across the seams. The window may be interrupted. | Click the drop-down box to select
Requirement 8.5.3 | Storage - The GB must have shelves at the back of the GB. | Click the drop-down box to select
Requirement 8.5.3.1 | Storage - There are areas where no shelves are allowed. These areas are indicated in the layout in appendix C04. | Click the drop-down box to select
Requirement 8.5.4 | Lighting - The box must have adequate and homogeneous lighting throughout the GB. The industry standard is a lighting strip across the entire top of the box, on the outside in LED or fluorescent lighting. The box will have an equivalent, or better lighting solution. | Click the drop-down box to select
Requirement 8.5.5 | Feedthroughs - The GB will require a number of argon gas feedthroughs, vacuum feedthroughs and 220V wall socket feedthroughs. You will find the GB layout in appendix C04. - Ar feedthrough shall be a 6 mm or ¼” stainless steel tube or a Swagelok connection (bulkhead connection or KF40 flange). - vacuum feedthrough shall be a 12mm or ½” stainless steel tube or Swagelok connection. | Click the drop-down box to select
Requirement 8.5.6 | Glovebox Pressure - The GB has a pressure regulation system. The user can set working pressure, maximum pressure and minimum pressure. | Click the drop-down box to select
Requirement 8.5.7 | Glovebox controls - The GB must have basic operation settings and controls. | Click the drop-down box to select
Requirement 8.5.7.1 | Glovebox controls - In reference to req. 8.5.6, the GB pressure can be controlled from the GB control panel. | Click the drop-down box to select
Requirement 8.5.7.2 | Glovebox controls - In reference to req. 8.5.4, the GB lighting can be switched on and off from the GB control panel. | Click the drop-down box to select
Requirement 8.5.7.3 | Glovebox controls - In reference to req. 8.7.1, the GB circulation can be switched on and off from the GB control panel, and the automatic regeneration cycle can be started. | Click the drop-down box to select
Requirement 8.5.7.4 | Glovebox controls - In reference to req. 8.8.1 and 8.8.3, the GB audible and visual alarm settings can be controlled from the GB control panel. | Click the drop-down box to select
Requirement 8.5.7.5 | Glovebox controls - In reference to req. 8.6.1, the GB AC and its settings can be controlled from the GB control panel. | Click the drop-down box to select
Requirement 8.5.7.6 | Glovebox controls - In reference to req. 8.8.2, the GB purge can be switched on and off from the GB control panel. | Click the drop-down box to select
Requirement 8.5.7.7 | Glovebox controls - In reference to pref. 8.8.2.1, the GB automatic purge settings can be controlled from the GB control panel. | Click the drop-down box to select
Requirement 8.5.7.8 | Glovebox controls – In reference to req. 8.6.1, the GB big AC can be set to vacuum, without compromising the usability of the GB. | Click the drop-down box to select
Requirement 8.5.8 | Glovebox controls - The controls will be primary accessible through buttons or a touchscreen outside of the GB. | Click the drop-down box to select
Requirement 8.5.10 | Display - The GB system will have a display. The display will be on the GB or attached to it and will have a maximum diameter of 61 cm (24 inch) and a minimum of 12.7 cm (5 inch). It will not be an extra separate computer or computer screen. | Click the drop-down box to select
Requirement 8.5.10.1 | Display - In relation to pref. 8.5.7, the GB display can graphically represent the logging data. | Click the drop-down box to select
Requirement 8.5.10.2 | Display - In relation to req. 8.7.1, the display can indicate the regeneration status or progress. | Click the drop-down box to select
Requirement 8.5.10.3 | Display - In relation to req. 8.6.1, the display can indicate the AC evacuation status and the evacuation cycle status or progress. | Click the drop-down box to select
Requirement 8.5.10.4 | Display - In relation to reqs. 8.7.1 and 8.8.3, the display can indicate the H2O and O2 levels and alarms. | Click the drop-down box to select
Requirement 8.5.11 | Metric units - The system will use metric units. Remark: preferred units: (m)bar, °C, ltr/min or m3/hr, gr, ppm and kg. | Click the drop-down box to select
Requirement 8.5.12 | Windows - The windows must have a coating to prevent scratches (for instance, a sapphire coated polycarbonate window is sufficient). The GB will not be fitted with bare polymer windows. | Click the drop-down box to select
Requirement 8.5.13 | Gloves - Standard butyl rubber gloves will be fitted on the GB. | Click the drop-down box to select
Requirement 8.5.14 | Vacuum pump - Any vacuum pump on the system will be a dry scroll pump capable of displacing at least 12.5 m3/hr. (Like ACP15G or similar, purge port, corrosion resistance). | Click the drop-down box to select
Requirement 8.6.1 | Antechamber 2.1 – Big antechamber The GB will have one big main antechamber (AC) on the right-hand side. - Size: Internal Diameter > 380mm, depth > 500mm but no deeper that 1000mm. - Temperature: The internal temperature can be set between room temperature and 120C or higher. The temperature homogeneity at any set temperature will be 20% or better. - Slide: The slide will have a ball or needle (or equivalent) bearing system to minimize generation of particles by metal-on-metal friction. - Vacuum lead safety: The AC vacuum lead will be protected from materials like wipes and small parts falling in and damaging the pump/clogging the vacuum lead. - Evacuation and Refill: the big AC will be under vacuum for prolonged periods of time. The AC will have the possibility to be evacuated and refilled. The GB still needs to be usable in a situation where the AC is under vacuum. Therefore, it will have a “close function” (available from the control panel see req. 8.5.7) or a separate pump. | Click the drop-down box to select
Requirement 8.6.1.1 | Big antechamber - The big AC must have an automated pressure-controlled evacuation and refill cycle with adjustable parameters: - A lower vacuum limit at which the AC starts refilling the AC. - An upper vacuum limit at which the AC starts evacuating again. - Amount of cycles. - A timer for the evacuation step. - If it takes to long for the AC to reach base pressure, the timer will end the cycle and trigger an (audible) alarm to inform the user that something is wrong. - 3 evacuation and refill cycles must not exceed 20 minutes. The settings are accessible through the control panel. | Click the drop-down box to select
Requirement 8.6.1.2 | Big antechamber - The big AC must have an interlock system so that the AC can only be opened from the inside when it is safe (e.g. when the AC atmosphere is inert). | Click the drop-down box to select
Requirement 8.6.2 | Antechamber 1.2 - Small Antechamber The GB will have a mini antechamber (AC) on the right hand side. - Size: Internal diameter > 150mm, depth > 300mm but no deeper that 600mm. - Slide: AC will have a sliding tray to accommodate introduction of materials into the GB. - Vacuum lead safety: The AC vacuum lead will be protected from materials like wipes and small parts falling in and damaging the pump/clogging the vacuum lead. | Click the drop-down box to select
Requirement 8.6.3 | The ACs will have their own Ar supply and will not fill with gas from the GB. | Click the drop-down box to select
Requirement 8.6.4 | The ACs are protected against overpressure from the Argon line. | Click the drop-down box to select
Requirement 8.7.1 | Purifier - The GB has a purifier system. | Click the drop-down box to select
Requirement 8.7.1.1 | Purifier - The circulation blower must have a capacity of at least 70 m3 per hour. | Click the drop-down box to select
Requirement 8.7.1.2 | Purifier - The purifier can absorb at least 35L of oxygen and 1250 gr of H2O and will be regenerable. | Click the drop-down box to select
Requirement 8.7.1.3 | Purifier - The purifier can maintain O2 and H2O levels below 1 ppm at a leakrate of 0,03 volume % per hour. Remark: This leakrate is specified to correspond to a large leak like a cut in a glove. | Click the drop-down box to select
Requirement 8.7.2 | Solvents - The GB has a solvent filter system which can remove the solvents listed in the solvent list provided in appendix C05. | Click the drop-down box to select
Requirement 8.7.4 | Particle filters - The system will have HEPA particle filters to filter incoming and outgoing working gas flow. The HEPA class is H13 or better. | Click the drop-down box to select
Requirement 8.8.1 | Audible alarms - An audible alarm will be available to indicate raised O2 and/or H2O levels. The alarm will not be loud, but clearly audible. Available settings must be at least: silenced on/off, activated/deactivated, the threshold settings are accessible through the controls. | Click the drop-down box to select
Requirement 8.8.2 | Box purge - The glovebox has a purge option that instead of circulating over the purifier supplies the box with inert supply gas. This purge is capable of reducing raised O2 and H2O levels in the box much faster than the purifier. | Click the drop-down box to select
Requirement 8.8.3 | Sensors - The GB system will have an H2O and an O2 sensor capable of a resolution of 1 ppm or better. | Click the drop-down box to select
Requirement 8.9.1 | Delivery time – It is required by TNO that delivery takes place within 22 weeks, after TNO submitted a purchase order. Note: not achieving the offered delivery time will lead to a penalty discount of 20% of the total order amount and shall be immediately due and payable and eligible for setoff and will replace clause 8.4 of the General Purchasing Conditions TNO – 2022. | Click the drop-down box to select
Requirement 8.9.3 | Installation and SAT - Installation and SAT of the GB systems is included and will be performed within 2 weeks after delivery of the GB systems (see also paragraph 2.2.23). | Click the drop-down box to select
Requirement 8.9.4 | Critical parts - Tenderer should provide a list of the most critical parts of the system together with argumentation why Tenderer considers these critical parts, as well as delivery time of such critical parts. Tenderer should also provide a list of recommended spare parts and consumables to have on stock at TNO. Spare parts shall be available for at least 10 years after SAT. | Click the drop-down box to select
Requirement 8.9.5 | Training - Operators training program: Tenderer should propose a training program for the machine operators, covering at least: - correct use of the glovebox, - safety instructions, - cleaning of the glovebox. - maintenance that can be carried out by TNO service engineer/tool owner and the operators. | Click the drop-down box to select
Requirement 8.9.6 | Warranty - The GB system must have a warranty period of at least 24 months, as stated in TNO’s General Purchasing Conditions – February 2022. | Click the drop-down box to select
Requirement 8.9.7 | Maintenance/Services - during warranty TNO requires the following response times: - Confirmation of receipt of service request or complaint within one (1) working-day*; - Service or support request online, in the event of a GB system failure, within two (2) working days*; - Service or support request on site, in the event of a GB system failure, within seven (7) calendar-days*. * within regular office hours and workdays. | Click the drop-down box to select
Requirement 8.9.9 | Language - The tenderer declares that all senior and executive members of staff/management of its organisation have a thorough command of the English language. It is a particular requirement that all oral and written communication during the (maintenance) Contract shall be in English for the duration of the agreement. | Click the drop-down box to select
Requirement 8.10.1 | Terms and conditions of payment - 30% at order, after receipt of a bank guarantee covering 80% of the total amount. - 50% after delivery at location TNO. - 20% after installation and commissioning On-site and Site Acceptance Test (SAT) including test periods, approved by TNO’s management. The Tenderer to whom the Contract will be awarded must provide a bank guarantee to TNO, covering at least 80% of the total price. The Bank guarantee and the bank who issues the Bank guarantee have to be approved by TNO (minimal A rating). TNO must have the right to draw the bank guarantee in any and all cases where Tenderer falls short in fulfilling its obligations under the Contract. The validity of the bank guarantee may not expire until approval of the SAT by TNO | Click the drop-down box to select
Requirement 8.11.1 | Documentation - Documentation must include at least the following: - Installation documentation - Instruction manual - Maintenance manual including all required technical documentation for Preventive and Corrective Maintenance. - Quick- start guide - P&ID of the system - Safety measures and their implementation - Certificates and manuals of all third-party hardware that is part of the GB system. All the documents should be delivered digitally in English and are expected during SAT. | Click the drop-down box to select
Requirement 8.11.2 | CE certification - Safety & Ergonomics standards It must be safe to operate the GB systems for a skilled operator without any safety issues. Official CE certification is required. | Click the drop-down box to select

## validatie (hidden)
Click the drop-down box to select
Yes, compliant to requirement
No, not compliant to requirement
note: dit blad 'verbergen' indien je dit gaat delen intern/extern (behoudens aan inkoop collega's)