## Requirements and Wishes
No. | Requirement | Type | Max. points or KO | Short description by applicant
1 | General description
1.1 | The instrument(s) will be purchased by KNMI. | M | KO
1.2 | The instrument(s) must be a Doppler LIDAR (LIght Detection And Ranging) device to measure the 3D wind field | M | KO
1.3 | The instrument(s) must be operating around a wavelength of 1.5 μm and must be eye-safe. | M | KO
1.4 | The instrument(s) must be capable of 24/7 unattended operation. | M | KO
1.5 | The instrument(s) must be able to operate for at least 10 years. | M | KO
2 | Wind measurements
2.1 | The instrument(s) must be able to measure and output range resolved profiles of the radial wind speed. | M | KO
2.2 | The instrument(s) must be able to measure and output range resolved profiles of signal-to-noise (SNR) or carrier-to-noise (CNR) value. | M | KO
2.3 | The instrument(s) must be able to scan the full hemisphere (range azimuthal angle 0° to 360°, range elevation angle 0° to 180°). | M | KO | Add short description how this Mandatory Requirement is met.
2.4 | The instrument(s) must be able to perform different user configurable scanning patterns to measure the 3D wind field: PPI (plan position indicator, "horizontal flow mapping"), RHI (range height indicator, "vertical flow mapping"), staring mode, and scans in which azimuth and elevation are both changed simultaneously. | M | KO | Add short description how this Mandatory Requirement is met.
2.5 | The instrument(s) must be able to measure vertical profiles of the horizontal wind speed and wind direction, and the vertical wind speed using a VAD (velocity azimuth display) and/or DBS (Doppler beam swinging) scan. | M | KO | Add short description how this Mandatory Requirement is met.
2.6 | The instrument(s) must be able to alternate between multiple scan patterns, and this must be user-configurable and schedulable. | M | KO | Add short description how this Mandatory Requirement is met.
2.7 | The instrument is able to measure the range resolved backscatter signal, in terms of the attenuated backscatter coefficient, with a relative measurement accuracy of the attenuated backscatter coefficients of 10% or better. | W C | 3 | Add short description how this wish is met.
3 | Parameter specifications and accuracies
3.1 | The maximum radial wind speed that can be measured must be 30 m/s or higher. | M | KO
3.2 | The measurement accuracy in the radial wind speed must be 0.5 m/s or better (one standard deviation, on basis of 10 minute averaging time). | M | KO | The Tenderer should provide at least 1 piece of information to confirm this.
3.3 | The accumulation time is user-configurable at least from 0.1 to 10 seconds. | M | KO
3.4 | The laser beam can be freely and reproducibly directed in the full hemisphere with a pointing accuracy of 0.1° or better. | M | KO
3.5 | The laser beam can be freely and reproducibly directed in the full hemisphere with an angular resolution of 0.2° or better. | M | KO
3.6 | The instrument(s) must be able to rotate at 6°/s or faster while measuring in azimuth direction and in elevation direction. | M | KO | Add short description how this Mandatory Requirement is met.
3.7 | The raw high frequency data (i.e. radial wind speeds) can be stored on the instrument(s) for at least 24 hours, and is accessible to the user. | M | KO | Add short description how this Mandatory Requirement is met.
3.8 | The pulse repetition frequency (PRF) can be set independently of the range resolution and range. | W C | 5.5 | Add short description how this wish is met.
4 | Set no. 1 of Parameter specifications and accuracies of at least one instrument
4.1 | The instrument must be able to capture the evolution of wake vortices of a medium sized airplane (Maximum Take Off Weight of approx. 50000 kg). Specify the settings (range resolution, PRF, etc). Prove that the instrument fulfills this purpose. | M | KO | The Tenderer should provide the specifications and at least 1 piece of information to confirm this.
4.2 | The minimal radial range of the instrument(s) (for the conditions of 4.1) must be 200 m or lower. | M | KO | Add short description how this Mandatory Requirement is met.
4.3 | The maximum radial range of the instrument(s) (for the conditions of 4.1) must be 1.4 km or higher. | M | KO | The Tenderer should provide at least 1 piece of information to confirm this.
4.4 | The instrument typical range under the measurement conditions of 4.1 is higher than 1.4 km. [for scoring: see tab 'Scoring matrices' 4.4] | W R | 5.5 | The Tenderer should provide at least 1 piece of information to confirm this.
4.5 | The instrument can measure with a range resolution of 25m or better. | M | KO | Add short description how this Mandatory Requirement is met.
4.6 | The range gates of the instrument can be overlapped in a user-configurable way to reduce the range-gate size to 5m or smaller. | M | KO | Add short description how this Mandatory Requirement is met.
4.7 | The instrument can measure with a range resolution without overlapping of smaller than 25m under conditions of 4.1. [for scoring: see tab 'Scoring matrices' 4.7] | W R | 3 | The Tenderer should provide at least 1 piece of information to confirm this.
4.8 | The unambiguous range under the conditions of 4.1 is higher than 7.5 km. [for scoring: see tab 'Scoring matrices' 4.8] | W R | 5.5 | Add separate document to answer this Wish (max. 1 A4 available).
5 | Set no. 2 of Parameter specifications and accuracies of at least one instrument
5.1 | The typical range of at least one instrument must be 10 km or higher (defined as the range at which the data availability has decreased to 80%, for a line-of-sight within the boundary layer, under "standard" atmospheric conditions (i.e. no clouds along the line of sight, no precipitation and visibility over 10 km), for an accumulation time of 1 s or less and a range resolution of 200 m or better, and taking into account a radial wind speed accuracy as specified in 3.2. | M | KO | The Tenderer should provide at least 3 pieces of information to confirm this.
5.2 | The minimal radial range of the instrument(s) (for the typical range and condition specified in 5.1) must be 300 m or lower. | M | KO | Add short description how this Mandatory Requirement is met.
5.3 | The instrument typical range under the measurement conditions of 5.1 is higher than 10 km. [for scoring: see tab 'Scoring matrices' 5.3] | W R | 5.5 | The Tenderer should provide at least 1 piece of information to confirm this.
5.4 | The maximum radial wind speed that can be measured is higher than 30 m/s. [for scoring: see tab 'Scoring matrices' 5.4] | W R | 5.5 | Add separate document to answer this Wish (max. 1 A4 available).
5.5 | The unambiguous range under the conditions of 5.1 is higher than 15 km. [for scoring: see tab 'Scoring matrices' 5.5] | W R | 5.5 | Add separate document to answer this Wish (max. 1 A4 available).
6 | Operational
6.1 | The dimensions per instrument are not greater than 3m x 3m x 3m (LxWxH) | M | KO
6.2 | The instrument(s) must have the opportunity of fine level adjustment of the positioning, particularly the horizontal placing (leveling). | M | KO
6.3 | The instrument(s) must have an easily visually recognizable North arrow. | M | KO
6.4 | The instrument(s) must be able to recover its latest configuration and automatically continue its scheduled measurement program after a power failure. | M | KO | Add short description how this Mandatory Requirement is met.
6.5 | The instrument(s) must have the possibility to implement a UPS or buffered power supply to handle short (5 minutes) power outages. A UPS does not have to be included. | M | KO
6.6 | The instrument(s) must specify a timestamp in Coordinated Universal Time (UTC) for the measurements. | M | KO
6.7 | The instrument(s) must synchronize its time with a configurable external time reference (Network Time Server or Global Positioning System). | M | KO | Add short description how this Mandatory Requirement is met.
6.8 | The instrument(s) must be equipped with an automated cleaning system to overcome issues with solid precipitation or other soiling (leaves, bird droppings) that blocks the window. | M | KO | Add short description how this Mandatory Requirement is met.
6.9 | If the weight of the instrument(s) is less than 2500 kg, the instrument(s) must be easily transportable on a trailer using a standard type vehicle. KNMI will execute the moves on Schiphol airport. If the weight of the instrument is more than 2500 kg, the Tenderer must offer transportation and placement at Schiphol airport, and to a different location at Schiphol airport for 3 times within a 2 year period. The moment in time of the moves is to be agreed with KNMI. | M | KO | Add short description how this Mandatory Requirement is met.
6.10 | In case a specific colour scheme is required by airport authorities to place the instrument at the airport, the Tenderer will provide the colouring of the instrument housing. | M | KO
7 | Interface
7.1 | The instrument(s) must be remotely accessible through an ethernet network connection and through 4G connection. | M | KO | Add short description how this Mandatory Requirement is met.
7.2 | The instrument(s) must have at least one standard ethernet connection based on TCP/IP protocol and provided with an RJ45 socket. | M | KO
7.3 | The instrument(s) must be remotely controllable via a software package for monitoring, troubleshooting, and for the implementation of configuration changes or ad-hoc data recovery. The software package must be included. | M | KO
7.4 | It must be possible to install updates or upgrades remotely. | M | KO
7.5 | Any person or system accessing the instrument should be authenticated in a secure manner. | M | KO | Add short description how this Mandatory Requirement is met.
8 | OS
8.1 | Operating system (OS) inside the instrument(s) must be supported in the coming 10 years, and must be remotely accessible (e.g. Remote Desktop). Updates to compliant OS are provided if OS or software becomes non-compliant. | M | KO | Add short description how this Mandatory Requirement is met.
8.2 | OS inside the instrument(s) must not be Windows XP. | M | KO
8.3 | OS inside the instrument(s) is Windows 10 LTSC or LINUX-based. | M | KO
8.4 | The instrument(s) must be able to internally store the standard measurement data for a period of at least 2 months (in continuous operation). | M | KO | Add short description how this Mandatory Requirement is met.
9 | Data transfer and format
9.1 | The instrument(s) must have the possibility of data transfer via SFTP. | M | KO
9.2 | The instrument should use international (metereological) standards for the exchange of information. For example NetCDF, HDF and/or CSV. | M | KO | Add short description how this Mandatory Requirement is met.
9.3 | Data format is CfRadial2. | W C | 1.5 | Add short description how this wish is met.
10 | Hardware (mechanical)
10.1 | The instrument(s) must be IP65 complaint or better regarding water and dust protection. | M | KO | Add short description how this Mandatory Requirement is met.
10.2 | The instrument(s) must be equipped with cables and connectors for use under exterior conditions according to European standards. These cables and connectors are available on the European market. | M | KO
11 | Hardware (electrical)
11.1 | The instrument(s) must be able to operate on a commercial power supply 230 VAC ± 10%, 45-60Hz single phase or three phases. European standard EN 50160 is applicable. | M | KO | Add short description how this Mandatory Requirement is met.
11.2 | The power consumption of a single instrument including auxilaries does not exceed 10 kW. | M | KO | Add short description how this Mandatory Requirement is met.
11.3 | The instrument(s) must include protection against power bursts and surges. European standards EN61000-4-4 and 61000-4-5 are applicable. | M | KO | Add short description how this Mandatory Requirement is met.
12 | Safety
12.1 | The instrument(s) must be classified as a laser class 1 or class 1M device, as defined in IEC/ EN 60825-1: 2014. | M | KO | Add short description how this Mandatory Requirement is met.
12.2 | The instrument(s) must be complaint with electric safety norm DIN-EN 60204 and the NEN 1010:2015 standard. | M | KO | Add short description how this Mandatory Requirement is met.
12.3 | The instrument(s) must comply with European standard EN 62305 concerning lightning protection and failure of electrical installation. | M | KO | Add short description how this Mandatory Requirement is met.
13 | Environmental (at instrument location)
13.1 | The instrument(s) must be able to operate at a humidity between 0 and 100%. | M | KO
13.2 | The instrument(s) must be able to operate at an ambient temperature between -25° C and +45° C. | M | KO
13.3 | The instrument(s) must be able to operate at a wind speed up to 40 m/s. | M | KO | Add short description how this Mandatory Requirement is met.
13.4 | The instrument(s) must be able to withstand the weather conditions in the Netherlands. | M | KO
13.5 | The instrument(s) must be able to withstand corrosion and UV radiation for at least 10 years (or spare parts are available). | M | KO
13.6 | A suitable air conditioning unit is provided by the Tenderer in case the instrument(s) needs to be air conditioned. If this is not necessary, clarify why not. | M | KO | Add short description how this Mandatory Requirement is met.
13.7 | The instrument(s) must comply with EMI/EMC standards for industrial environments. European standards EN50081-1 and EN50081-2 are applicable. | M | KO | Add short description how this Mandatory Requirement is met.
14 | Software
14.1 | The Tenderer must inform KNMI of each new software or firmware release for the instrument(s). This includes updates with additional features as well as bug fixes. Information is provided on the content of the update, its consequences for the operation of the instrument(s) and whether or not it is a critical update. Critical updates are supplied free of charge during the entire period of operations. | M | KO | Add short description how this Mandatory Requirement is met.
14.2 | Software is provided for configuration and monitoring of hardware. | M | KO | Add short description how this Mandatory Requirement is met.
14.3 | Software is provided to visualize measurement data graphically within 20s during operation. | M | KO | Add short description how this Mandatory Requirement is met.
14.4 | Software is provided to calculate and visualize the 3D wind components from the radial wind measurements near real-time. | M | KO | Add short description how this Mandatory Requirement is met.
14.5 | Software is provided to visualize radial wind of a PPI scan. | M | KO | Add short description how this Mandatory Requirement is met.
14.6 | Internal algorithms for data processing relevant for the wind vector retrieval and uncertainty estimation are documented in scientific publications. | W C | 1.5 | Add separate document(s) to answer this Wish.
15 | Reliability
15.1 | The instrument(s) is able to operate unattended for at least 6 months. | M | KO | Add short description how this Mandatory Requirement is met.
15.2 | The MTBF of the scanning components need to be higher than 6 months for the instrument that will apply the wake-vortex measurement strategy of 4.1 for at least 75% of the time. | M | KO | Add short description how this Mandatory Requirement is met.
15.3 | The operational up-time of the instrument(s) must be at least 95%. In case of repair is needed that exceeds the maximum downtime, it is allowed to replace the instrument with a replacement instrument of the same type. | M | KO | Add short description how this Mandatory Requirement is met. This description should also contain the maximum possible up-time offered by the manufacturer.
16 | Support and maintenance
16.1 | During the intended use (up to 10 years), Tenderer must support and maintain the offered instrument(s). | M | KO | Add short description how this Mandatory Requirement is met.
16.2 | A helpdesk with skilled personell is available, reachable during working hours (9:00 - 17:00 CET) by telephone and e-mail using the English language. | M | KO | Add short description how this Mandatory Requirement is met.
16.3 | The response time to start investigation after a KNMI request for corrective maintenance is 1 working day. | M | KO
16.4 | The response time for measurement related questions and advice (not related to technical problems of the instrument) is 3 working days or less. | M | KO
16.5 | Maintenance and repair is performed by the Tenderer and is included in the contract for a period of 2 years from instrument(s) acceptance by the customer. Replacing cleaning fluid and drying agent can be performed by the Contracting body. | M | KO
16.6 | A yearly price is specified for maintenance and repair for years 3 to 10 of service in Appendix Pricing Sheet of the Tender Documents. | M | KO
17 | Services
17.1 | The instrument(s) must be calibrated by the provider, and calibration documentation must be provided. | M | KO
17.2 | The Tenderer must re-calibrate the instrument in case of repair or maintenance of a critical component. | M | KO | Add short description how this Mandatory Requirement is met.
17.3 | A training (of a duration of 1 or 2 days) on proper handling and usage of the instrument(s) must be provided (for a group of not more than 5 persons KNMI personal). | M | KO | Add short description how this Mandatory Requirement is met.
18 | Delivery
18.1 | The instrument(s) must be delivered within 6 months after signing the contract en price of delivery is included in the Offer. | M | KO
18.2 | The instrument(s) will be delivered within a shorter period of time than 6 months after signing the contract. | W R | 3 | Add short description how this wish is met and what the delivery time will be, in months.
18.3 | The instrument(s) must be delivered to De Bilt, The Netherlands or directly to Schiphol airport, The Netherlands. The location is to be decided by KNMI. | M | KO
18.4 | The instrument(s) must be delivered in a protective, reusable case that prevents damage to the instrument(s) during transport under normal conditions. | M | KO
18.5 | The instrument(s) must be delivered with a complete documentation of all hardware and software components (electronic as pdf files). Language must be English. | M | KO

## Check list
REQUIREMENT CHECK LIST. FILL IN AND ADD TO TENDER.
No. | Requirement | Type | Agreed (Yes/No) | Short description by applicant
1 | General description
1.1 | The instrument(s) will be purchased by KNMI. | M
1.2 | The instrument(s) must be a Doppler LIDAR (LIght Detection And Ranging) device to measure the 3D wind field | M
1.3 | The instrument(s) must be operating around a wavelength of 1.5 μm and must be eye-safe. | M
1.4 | The instrument(s) must be capable of 24/7 unattended operation. | M
1.5 | The instrument(s) must be able to operate for at least 10 years. | M
2 | Wind measurements
2.1 | The instrument(s) must be able to measure and output range resolved profiles of the radial wind speed. | M
2.2 | The instrument(s) must be able to measure and output range resolved profiles of signal-to-noise (SNR) or carrier-to-noise (CNR) value. | M
2.3 | The instrument(s) must be able to scan the full hemisphere (range azimuthal angle 0° to 360°, range elevation angle 0° to 180°). | M |  | Add short description how this Mandatory Requirement is met.
2.4 | The instrument(s) must be able to perform different user configurable scanning patterns to measure the 3D wind field: PPI (plan position indicator, "horizontal flow mapping"), RHI (range height indicator, "vertical flow mapping"), staring mode, and scans in which azimuth and elevation are both changed simultaneously. | M |  | Add short description how this Mandatory Requirement is met.
2.5 | The instrument(s) must be able to measure vertical profiles of the horizontal wind speed and wind direction, and the vertical wind speed using a VAD (velocity azimuth display) and/or DBS (Doppler beam swinging) scan. | M |  | Add short description how this Mandatory Requirement is met.
2.6 | The instrument(s) must be able to alternate between multiple scan patterns, and this must be user-configurable and schedulable. | M |  | Add short description how this Mandatory Requirement is met.
2.7 | The instrument is able to measure the range resolved backscatter signal, in terms of the attenuated backscatter coefficient, with a relative measurement accuracy of the attenuated backscatter coefficients of 10% or better. | W C |  | Add short description how this wish is met.
3 | Parameter specifications and accuracies
3.1 | The maximum radial wind speed that can be measured must be 30 m/s or higher. | M
3.2 | The measurement accuracy in the radial wind speed must be 0.5 m/s or better (one standard deviation, on basis of 10 minute averaging time). | M |  | The Tenderer should provide at least 1 piece of information to confirm this.
3.3 | The accumulation time is user-configurable at least from 0.1 to 10 seconds. | M
3.4 | The laser beam can be freely and reproducibly directed in the full hemisphere with a pointing accuracy of 0.1° or better. | M
3.5 | The laser beam can be freely and reproducibly directed in the full hemisphere with an angular resolution of 0.2° or better. | M
3.6 | The instrument(s) must be able to rotate at 6°/s or faster while measuring in azimuth direction and in elevation direction. | M |  | Add short description how this Mandatory Requirement is met.
3.7 | The raw high frequency data (i.e. radial wind speeds) can be stored on the instrument(s) for at least 24 hours, and is accessible to the user. | M |  | Add short description how this Mandatory Requirement is met.
3.8 | The pulse repetition frequency (PRF) can be set independently of the range resolution and range. | W C |  | Add short description how this wish is met.
4 | Set no. 1 of Parameter specifications and accuracies of at least one instrument
4.1 | The instrument must be able to capture the evolution of wake vortices of a medium sized airplane (Maximum Take Off Weight of approx. 50000 kg). Specify the settings (range resolution, PRF, etc). Prove that the instrument fulfills this purpose. | M |  | The Tenderer should provide the specifications and at least 1 piece of information to confirm this.
4.2 | The minimal radial range of the instrument(s) (for the conditions of 4.1) must be 200 m or lower. | M |  | Add short description how this Mandatory Requirement is met.
4.3 | The maximum radial range of the instrument(s) (for the conditions of 4.1) must be 1.4 km or higher. | M |  | The Tenderer should provide at least 1 piece of information to confirm this.
4.4 | The instrument typical range under the measurement conditions of 4.1 is higher than 1.4 km. [for scoring: see tab 'Scoring matrices' 4.4] | W R |  | The Tenderer should provide at least 1 piece of information to confirm this.
4.5 | The instrument can measure with a range resolution of 25m or better. | M |  | Add short description how this Mandatory Requirement is met.
4.6 | The range gates of the instrument can be overlapped in a user-configurable way to reduce the range-gate size to 5m or smaller. | M |  | Add short description how this Mandatory Requirement is met.
4.7 | The instrument can measure with a range resolution without overlapping of smaller than 25m under conditions of 4.1. [for scoring: see tab 'Scoring matrices' 4.7] | W R |  | The Tenderer should provide at least 1 piece of information to confirm this.
4.8 | The unambiguous range under the conditions of 4.1 is higher than 7.5 km. [for scoring: see tab 'Scoring matrices' 4.8] | W R |  | Add separate document to answer this Wish (max. 1 A4 available).
5 | Set no. 2 of Parameter specifications and accuracies of at least one instrument
5.1 | The typical range of at least one instrument must be 10 km or higher (defined as the range at which the data availability has decreased to 80%, for a line-of-sight within the boundary layer, under "standard" atmospheric conditions (i.e. no clouds along the line of sight, no precipitation and visibility over 10 km), for an accumulation time of 1 s or less and a range resolution of 200 m or better, and taking into account a radial wind speed accuracy as specified in 3.2. | M |  | The Tenderer should provide at least 3 pieces of information to confirm this.
5.2 | The minimal radial range of the instrument(s) (for the typical range and condition specified in 5.1) must be 300 m or lower. | M |  | Add short description how this Mandatory Requirement is met.
5.3 | The instrument typical range under the measurement conditions of 5.1 is higher than 10 km. [for scoring: see tab 'Scoring matrices' 5.3] | W R |  | The Tenderer should provide at least 1 piece of information to confirm this.
5.4 | The maximum radial wind speed that can be measured is higher than 30 m/s. [for scoring: see tab 'Scoring matrices' 5.4] | W R |  | Add separate document to answer this Wish (max. 1 A4 available).
5.5 | The unambiguous range under the conditions of 5.1 is higher than 15 km. [for scoring: see tab 'Scoring matrices' 5.5] | W R |  | Add separate document to answer this Wish (max. 1 A4 available).
6 | Operational
6.1 | The dimensions per instrument are not greater than 3m x 3m x 3m (LxWxH) | M
6.2 | The instrument(s) must have the opportunity of fine level adjustment of the positioning, particularly the horizontal placing (leveling). | M
6.3 | The instrument(s) must have an easily visually recognizable North arrow. | M
6.4 | The instrument(s) must be able to recover its latest configuration and automatically continue its scheduled measurement program after a power failure. | M |  | Add short description how this Mandatory Requirement is met.
6.5 | The instrument(s) must have the possibility to implement a UPS or buffered power supply to handle short (5 minutes) power outages. A UPS does not have to be included. | M
6.6 | The instrument(s) must specify a timestamp in Coordinated Universal Time (UTC) for the measurements. | M
6.7 | The instrument(s) must synchronize its time with a configurable external time reference (Network Time Server or Global Positioning System). | M |  | Add short description how this Mandatory Requirement is met.
6.8 | The instrument(s) must be equipped with an automated cleaning system to overcome issues with solid precipitation or other soiling (leaves, bird droppings) that blocks the window. | M |  | Add short description how this Mandatory Requirement is met.
6.9 | If the weight of the instrument(s) is less than 2500 kg, the instrument(s) must be easily transportable on a trailer using a standard type vehicle. KNMI will execute the moves on Schiphol airport. If the weight of the instrument is more than 2500 kg, the Tenderer must offer transportation and placement at Schiphol airport, and to a different location at Schiphol airport for 3 times within a 2 year period. The moment in time of the moves is to be agreed with KNMI. | M |  | Add short description how this Mandatory Requirement is met.
6.10 | In case a specific colour scheme is required by airport authorities to place the instrument at the airport, the Tenderer will provide the colouring of the instrument housing. | M
7 | Interface
7.1 | The instrument(s) must be remotely accessible through an ethernet network connection and through 4G connection. | M |  | Add short description how this Mandatory Requirement is met.
7.2 | The instrument(s) must have at least one standard ethernet connection based on TCP/IP protocol and provided with an RJ45 socket. | M
7.3 | The instrument(s) must be remotely controllable via a software package for monitoring, troubleshooting, and for the implementation of configuration changes or ad-hoc data recovery. The software package must be included. | M
7.4 | It must be possible to install updates or upgrades remotely. | M
7.5 | Any person or system accessing the instrument should be authenticated in a secure manner. | M |  | Add short description how this Mandatory Requirement is met.
8 | OS
8.1 | Operating system (OS) inside the instrument(s) must be supported in the coming 10 years, and must be remotely accessible (e.g. Remote Desktop). Updates to compliant OS are provided if OS or software becomes non-compliant. | M |  | Add short description how this Mandatory Requirement is met.
8.2 | OS inside the instrument(s) must not be Windows XP. | M
8.3 | OS inside the instrument(s) is Windows 10 LTSC or LINUX-based. | M
8.4 | The instrument(s) must be able to internally store the standard measurement data for a period of at least 2 months (in continuous operation). | M |  | Add short description how this Mandatory Requirement is met.
9 | Data transfer and format
9.1 | The instrument(s) must have the possibility of data transfer via SFTP. | M
9.2 | The instrument should use international (metereological) standards for the exchange of information. For example NetCDF, HDF and/or CSV. | M |  | Add short description how this Mandatory Requirement is met.
9.3 | Data format is CfRadial2. | W C |  | Add short description how this wish is met.
10 | Hardware (mechanical)
10.1 | The instrument(s) must be IP65 complaint or better regarding water and dust protection. | M |  | Add short description how this Mandatory Requirement is met.
10.2 | The instrument(s) must be equipped with cables and connectors for use under exterior conditions according to European standards. These cables and connectors are available on the European market. | M
11 | Hardware (electrical)
11.1 | The instrument(s) must be able to operate on a commercial power supply 230 VAC ± 10%, 45-60Hz single phase or three phases. European standard EN 50160 is applicable. | M |  | Add short description how this Mandatory Requirement is met.
11.2 | The power consumption of a single instrument including auxilaries does not exceed 10 kW. | M |  | Add short description how this Mandatory Requirement is met.
11.3 | The instrument(s) must include protection against power bursts and surges. European standards EN61000-4-4 and 61000-4-5 are applicable. | M |  | Add short description how this Mandatory Requirement is met.
12 | Safety
12.1 | The instrument(s) must be classified as a laser class 1 or class 1M device, as defined in IEC/ EN 60825-1: 2014. | M |  | Add short description how this Mandatory Requirement is met.
12.2 | The instrument(s) must be complaint with electric safety norm DIN-EN 60204 and the NEN 1010:2015 standard. | M |  | Add short description how this Mandatory Requirement is met.
12.3 | The instrument(s) must comply with European standard EN 62305 concerning lightning protection and failure of electrical installation. | M |  | Add short description how this Mandatory Requirement is met.
13 | Environmental (at instrument location)
13.1 | The instrument(s) must be able to operate at a humidity between 0 and 100%. | M
13.2 | The instrument(s) must be able to operate at an ambient temperature between -25° C and +45° C. | M
13.3 | The instrument(s) must be able to operate at a wind speed up to 40 m/s. | M |  | Add short description how this Mandatory Requirement is met.
13.4 | The instrument(s) must be able to withstand the weather conditions in the Netherlands. | M
13.5 | The instrument(s) must be able to withstand corrosion and UV radiation for at least 10 years (or spare parts are available). | M
13.6 | A suitable air conditioning unit is provided by the Tenderer in case the instrument(s) needs to be air conditioned. If this is not necessary, clarify why not. | M |  | Add short description how this Mandatory Requirement is met.
13.7 | The instrument(s) must comply with EMI/EMC standards for industrial environments. European standards EN50081-1 and EN50081-2 are applicable. | M |  | Add short description how this Mandatory Requirement is met.
14 | Software
14.1 | The Tenderer must inform KNMI of each new software or firmware release for the instrument(s). This includes updates with additional features as well as bug fixes. Information is provided on the content of the update, its consequences for the operation of the instrument(s) and whether or not it is a critical update. Critical updates are supplied free of charge during the entire period of operations. | M |  | Add short description how this Mandatory Requirement is met.
14.2 | Software is provided for configuration and monitoring of hardware. | M |  | Add short description how this Mandatory Requirement is met.
14.3 | Software is provided to visualize measurement data graphically within 20s during operation. | M |  | Add short description how this Mandatory Requirement is met.
14.4 | Software is provided to calculate and visualize the 3D wind components from the radial wind measurements near real-time. | M |  | Add short description how this Mandatory Requirement is met.
14.5 | Software is provided to visualize radial wind of a PPI scan. | M |  | Add short description how this Mandatory Requirement is met.
14.6 | Internal algorithms for data processing relevant for the wind vector retrieval and uncertainty estimation are documented in scientific publications. | W C |  | Add separate document(s) to answer this Wish.
15 | Reliability
15.1 | The instrument(s) is able to operate unattended for at least 6 months. | M |  | Add short description how this Mandatory Requirement is met.
15.2 | The MTBF of the scanning components need to be higher than 6 months for the instrument that will apply the wake-vortex measurement strategy of 4.1 for at least 75% of the time. | M |  | Add short description how this Mandatory Requirement is met.
15.3 | The operational up-time of the instrument(s) must be at least 95%. In case of repair is needed that exceeds the maximum downtime, it is allowed to replace the instrument with a replacement instrument of the same type. | M |  | Add short description how this Mandatory Requirement is met. This description should also contain the maximum possible up-time offered by the manufacturer.
16 | Support and maintenance
16.1 | During the intended use (up to 10 years), Tenderer must support and maintain the offered instrument(s). | M |  | Add short description how this Mandatory Requirement is met.
16.2 | A helpdesk with skilled personell is available, reachable during working hours (9:00 - 17:00 CET) by telephone and e-mail using the English language. | M |  | Add short description how this Mandatory Requirement is met.
16.3 | The response time to start investigation after a KNMI request for corrective maintenance is 1 working day. | M
16.4 | The response time for measurement related questions and advice (not related to technical problems of the instrument) is 3 working days or less. | M
16.5 | Maintenance and repair is performed by the Tenderer and is included in the contract for a period of 2 years from instrument(s) acceptance by the customer. Replacing cleaning fluid and drying agent can be performed by the Contracting body. | M
16.6 | A yearly price is specified for maintenance and repair for years 3 to 10 of service in Appendix Pricing Sheet of the Tender Documents. | M
17 | Services
17.1 | The instrument(s) must be calibrated by the provider, and calibration documentation must be provided. | M
17.2 | The Tenderer must re-calibrate the instrument in case of repair or maintenance of a critical component. | M |  | Add short description how this Mandatory Requirement is met.
17.3 | A training (of a duration of 1 or 2 days) on proper handling and usage of the instrument(s) must be provided (for a group of not more than 5 persons KNMI personal). | M |  | Add short description how this Mandatory Requirement is met.
18 | Delivery
18.1 | The instrument(s) must be delivered within 6 months after signing the contract en price of delivery is included in the Offer. | M
18.2 | The instrument(s) will be delivered within a shorter period of time than 6 months after signing the contract. | W R |  | Add short description how this wish is met and what the delivery time will be, in months.
18.3 | The instrument(s) must be delivered to De Bilt, The Netherlands or directly to Schiphol airport, The Netherlands. The location is to be decided by KNMI. | M
18.4 | The instrument(s) must be delivered in a protective, reusable case that prevents damage to the instrument(s) during transport under normal conditions. | M
18.5 | The instrument(s) must be delivered with a complete documentation of all hardware and software components (electronic as pdf files). Language must be English. | M

## Values (hidden)
mandatory | M
wish (closed) | W C
wish (open) | W O

## Value
Yes | M
No | W C
W R

## Abbreviations
Abbreviations
μm | Micrometer
m/s | Meters per second
°/s | Degrees per second
kg | Kilogram
km | Kilometers
s | Seconds
° | Degrees
° C | Degrees Celcius
kW | KiloWatt
% | Percent
24/7 | 24 hours per day, 7 days a week
LxWxH | Length x Width x Height
CET | Central European Time