## PvE
PvE: CT RT Dordrecht
| Requirement | Answer supplier | Clarification / statement of proof supplier
Equipment Requirements
1 | Bore hole diameter (cm) ≥ 80
2 | Multipurpose Patient Table with flat RTP table top: Fully TG-66 compliant over the scan range
3 | Positioning accuracy with maximum table load (>200kg) according to IEC 60601
4 | Patient table is exchangable with Mastercouch, Omnicouch and Omniboard, that are currently used at the treatment machines
5 | Cool medium gantry is air
6 | Medrad Double Head contrast injector is included/integrated (full functional interface) and flexible to position
Functional Requirements
7 | Volumetric, helical scanning technique
8 | Sequential scanning technique
9 | Minimum rotation time ≤ 350 ms
10 | Maximum helical scan speed for diagnostic images ≥ 200 mm/s
11 | Minimal slice thickness ≤ 1 mm
12 | The scanning field of view is ≥ 60 cm
13 | Total beam collimation at minimum slice thickness ≥ 35 mm
14 | Best isotropic resolution in helical scan mode ≤ 0,40mm
15 | Automatic exposure control (tube current modulation)
16 | Respiratory gating and triggering (without fixation belt)
Reconstructions
17 | Image matrix reconstruction sizes at least 512x512
18 | Model based iterative reconstruction technique
19 | FBP reconstruction speed ≥ 30 images/s
20 | Iterative reconstruction speed ≥ 20 images/s
21 | CT reconstruction algorithm to suppress artefacts due to metal parts
22 | Respiratory 4D CT sequence scan mode that adapts to the patient’s breathing
23 | 4D reconstruction methods (both phase- and amplitude based) that automatically selects the correct data (bins)
Dicom
24 | Communications to other systems is possible according to DICOM Standards Committee part 1-16, with:
24a | DICOM Storage (CT image storage SOP class)
24b | DICOM query/retrieve as an SCP (C-FIND and C-MOVE)
24c | DICOM modality worklist (including automatic retrievement of patient names and appointments)
24d | DICOM MPPS
24e | DICOM storage commitment
25 | Please include the DICOM conformance statement
Levering, Service en garantie
26 | Instruction for use (gebruikershandleiding) and CE quality mark (CE keurmerk) will be included in the offer
27 | Training of two Erasmus MC medical technicians is included in the offer, including travel and accommodation costs
28 | After completing the service training, Erasmus MC hospital technicians receive the same rights and can perform the same work as the service engineer of the system manufacturer. Software and/or hardware for this will be provided after completion of the training
29 | Warranty period is at least twelve months and starts after clinical commissioning of the system
30 | Two complete sets of service documentation, theory of operation, drawings and partlists are included. Software for ordering information, partlists or exploded views are considered part of the service documentation package.
31 | Documentation need to be accessible at all times, without any restrictions (e.g. passwords that change in time)
32 | The supplier guarantees that the equipment will function for ten years in accordance with the functional and technical specifications (see purchase agreement article 8.1)
Overig
33 | Training of users according to the needs of Erasmus MC is included in the offer
34 | Complete system is compatible with room standard NEN1010-A7/3 class 2
35 | Complete system complies with CE, IEC 60601 and IEC 1217 standards
36 | Supplier provides an clear overview of specifications of the offered product
37 | Supplier agrees an complies with attachment 4 - Concept overeenkomst inclusief algemene inkoopvoorwaarden NFU
38 | Supplier agrees an can comply with attachment 6 - Aansluitvoorwaarden Erasmus MC Netwerk
Handtekening rechtsgeldig vertegenwoordiger
Bedrijfsnaam:
Handtekening:
Naam rechtsgeldig vertegenwoordiger:
Functie:
Datum:

## PvW
PvW: CT RT Dordrecht
Score= points * weight
| Questions | insufficient = 0 | sufficient=1 | good = 2 | weight | Answer supplier | Clarification / statement of proof supplier | Points
Equipment Requirements
1 | No need for a separate technical room | N | x | Y | 4
2 | Integrated laser system into the CT gantry, including all necessary items | N | x | Y | 10
3 | The system must be user friendly and controls must work intuitively, straight forward, and minimize amount of clicks pre- and post scan. Please describe a workflow, including # mouse clicks, for a Head-and-Neck scan | N | x | Y | 6
4 | Shortcut (keypad/tablet) options for most common (mouse) actions. Keypad with numeric block. | N | x | Y | 4
5 | (multiple) Tablets are provided for remote control and can be easily charged and set aside using an accessory/docking station | N | x | Y | 4
6 | Standard acquisition/reconstruction protocols are user configurable | N | x | Y | 4
7 | The modified protocols can easily be synchronized between different CT scanners from the same manufactory (safety issue) | N | x | Y | 8
8 | Adjustment of individual couch positions and gantry tilt is possible from both sides of the gantry | N | x | Y | 4
9 | Patient table foot switch is included | N | x | Y | 4
10 | Gantry tilt can be blocked easily | N | x | Y | 4
11 | Multi Purpose patient table with flat RTP tabletop is included, allowing the use of a lock system to mount accessories to the table | N | x | Y | 6
12 | Position accuracy with load of >200 kg in accorcance with IEC 60601: within 5mm and 0.5 degree | N | x | Y | 4
13 | The couch has a longitudinal travel of 200 cm | < 150 cm | 150 - 200 cm | > 200 cm | 6
Functional Requirements
14 | Number of detector rows | ≤32 | <64 | ≥64 | 6
15 | Acquired slices per rotation | ≤32 | <64 | ≥64 | 6
16 | Reconstructed slices per rotation | ≤64 | <128 | ≥128 | 8
17 | For both Helical and Sequence scan modes different collimatation are possible and easily / freely adjustable. Please specify all scan modes available | N | x | Y | 4
18 | In Helical scan mode the pitch factor is freely adjustable | N | x | Y | 4
19 | CT scan range of at least 150 cm | < 150 cm | 150 cm | > 150 cm | 4
20 | The scanning field of view is at least 65 cm | < 65 cm | 65 | > 65 cm | 6
21 | The extended field of view is at least 80 cm | < 80 cm | 80 cm | > 80 cm | 8
22 | In the FOV ring between 50 and 65 cm the HU accuracy is better than +-100 HU. | N | x | Y | 6
23 | HU accuracy (+- max deviation): - HU(water) = 0 +- 4 HU - HU(air) = -1000 +- 10 HU | N | x | Y | 4
24 | Scanning, reconstructing and displaying slices are possible simultaneously | N | x | Y | 4
Reconstructions
25 | Raw CT data can be stored and used for various reconstruction afterwards | N | x | Y | 4
26 | Reconstruction matrix of at least 512x512 | N | 512x512 only | > 512x512 | 8
27 | Reconstruction of at least 20 slices/sec is possible | < 15 | 15-20 | > 20 | 4
28 | Extended HU range reconstructions | N | x | Y | 4
29 | Various parameter settings including slice thickness, slice spacing and slice location, can be adjusted in advance of reconstructing the raw data | N | x | Y | 6
30 | The system time as used for time stamping, can also be read out via an I/O interface during scanning, e.g., in case of respiration-gated CT scanning (see section Respiration) | N | x | Y | 3
31 | For all scanning and reconstruction methods, slices are time-stamped on a system time scale, in such a way that the time interval in which the projections used for reconstruction of the slice are acquired can precisely be determined | N | x | Y | 3
32 | A CT reconstruction algorithm to suppress artefacts due to metal parts. Please provide examples (different metal sizes, patient sizes, anatomical locations) with/without correction | No good examples | One good example | Multiple good examples | 8
33 | Solutions dedicated for challenging metal implants resulting in reduced metal artifacts, show examples if possible | No good examples | One good example | Multiple good examples | 6
34 | Slower rotation times (> 5 seconds) can be selected to average the motion trajectory of moving anatomical structures in a single multi-slice set (slow-scan) | N | x | Y | 4
Respiration-gated CT scanning, 4D data analysis
35 | Respiratory gated, signal detection is without a fixation belt (due to hygien regulatory), and is comfortable for the patient | N | x | Y | 8
36 | The 4D workflow must be user friendly and controls must work intuitively, straight forward, and minimize amount of clicks pre- and post scan. Please describe the 4D workflow, and # mouse clicks for a lung 4D-scan | N | x | Y | 6
37 | To minimize artefacts in 4D scans, scan parameters are automatically based/updated on patients individual breathing (measured prior to scanning/real time during scanning) | N | x | Y | 10
38 | 4D scan duration is adapted to patients breathing | N | x | Y | 6
39 | For 4DCT-scanning a minimum breathing frequency of 4 per min is possible | min > 6 | min 4-6 | min < 4 | 4
40 | The latency time between trigger and scan start is less than 0.5 seconds | N | x | Y | 2
41 | Optimal data (bins) are automatically selected for automated 4D reconstruction | N | x | Y | 6
42 | Reconstruction of 4DCT scan takes less than 5min | >5 | 5 | <5 | 6
43 | Dataset compression methods are implemented to decrease the size of 4D datasets | N | x | Y | 2
44 | The available optical or electronic I/O interface for respiration-gated CT scanning can be controlled with standard commercially available non-proprietary hard- and software | N | x | Y | 2
45 | Externally registered signals (e.g., breathing) can be stored with an identical time stamp, and on the same system, as the CT dataset | N | x | Y | 4
46 | All slices and phases from 4D datasets can be exported in DICOM format, including their time stamps and the corresponding externally registered signals | N | x | Y | 3
47 | The respiration-gated trigger system will be included in the delivery and made fully operational in conjunction with CT scanning | N | x | Y | 6
IT requirements
48 | The CT control system runs on a commercially available operating system, currently like one of the following: Windows 10 Long Time Stable, Linux or Unix (HP-UX or SunOS). If an operating system is updated, the CT control system will be updated accordingly. | N | x | Y | 8
49 | Network communication is based on TCP/IP only | N | x | Y | 4
50 | Network interface card supports at least 10 Gbit full-duplex | N | x | Y | 4
51 | Possible to store both the raw data and the reconstructed images on USB/DVD | N | x | Y | 2
Technical issues and service
52 | System is designed of separate modules | N | x | Y | 3
53 | Digital control of all system parameters (using EPROM's) is required | N | x | Y | 3
54 | Tools to verify the constancy and quality of machine parameters are integrated in the system | N | x | Y | 4
55 | All relevant service tools and test equipment (phantom) are included | N | x | Y | 4
56 | Delivered phantoms can be used for the calibration of the internal lasers | N | x | Y | 4
57 | When in service mode, it is possible to bypass control systems | N | x | Y | 3
58 | The customer is allowed to use parts that are freely on the market for replacements | N | x | Y | 2
59 | Any updates or service bulletins available for the CT are announced to the hospital by means of speedinfo, E-mail or conventional mail and addressed to the hospital service engineer | N | x | Y | 2
Other requirements, specify extra costs if applicable
60 | Software for image guided surgery and virtual guided needle implantation is included | N | x | Y | 4
61 | Bolus tracking = CT acquisition can be triggered by an input bolus. That is, - Continuous scanning on a specific ROI, typically an artery. - Some time after contrast injection the increased HU in the ROI is detected. - This triggers the scanning of the CT scan. | N | x | Y | 2
62 | Automatic selection of optimal kV setting based on patient size and/or clinical task, including automatic reconstruction with appropriate calibration curves. Please provide examples to show improved CT image quality in patient and clincal specific cases that would benefit for radiotherapy purposes | N | x | Y | 4
63 | Dual energy CT acquisition, show potential benefits for radiotherapy purposes | N | x | Y | 4
Handtekening rechtsgeldig vertegenwoordiger
Bedrijfsnaam:
Handtekening:
Naam rechtsgeldig vertegenwoordiger:
Functie:
Datum:

## Informatieblad
Informatieblad CT RT Dordrecht
| Vraagstelling | Antwoord leverancier
Accessoires en opties
1 | Beschrijf alle opties, met prijzen
2 | Beschrijf alle beschikbare accessoires, met prijzen
Overig
3 | Verwachte levensduur röntgenbuis (in scanseconden)
4 | Warmte-afgifte totaal?
5 | Is er een IHE integration statement van toepassing op dit systeem? (zo ja, welke.)
6 | Datasheet en brochure van het systeem bijvoegen.
7 | Heeft u als leverancier een mission statement m.b.t. maatschappelijk verantwoord ondernemen? Zo ja, geef concreet aan hoe deze mission statement vertaald wordt in objectieve doelstellingen. Gaarne het Mission Statement bijvoegen
8 | Heeft u als leverancier concrete doelstelling(en) om het energieverbruik van de desbetreffende apparatuur te verminderen? Zo ja, welke?
9 | Referentiesites
10 | Voeg CT-data bij van de aangeboden scanner, gemaakt in een klinische setting, bijvoorbeeld abdomen, hoofd-hals, obese patiënten
11 | Specificeer het aantal detectoren per rij en het aantal rijen + afmetingen
12 | Specificeer het geluidsniveau van de scanner en geef aan hoe dit gemeten is
13 | Geef het aantal "mouse clicks"dat nodig is voor een head-neck scan (inclusief alle stappen van voorbereiding tot uitvoering)
Handtekening rechtsgeldig vertegenwoordiger
Bedrijfsnaam:
Handtekening:
Naam rechtsgeldig vertegenwoordiger:
Functie:
Datum: