Annex A Program of Requirements (PoR)

# General information about location and installation aspects

Astron, LOFAR distributed Radio-Telescope with central location in Exloo (Drenthe)

The White Rabbit switches will be used to provide a central clock system for the Dutch part of the LOFAR telescope.

The Low Jitter Switches will be placed in a Central location, distributing the clock signal to all stations where the standard White Rabbit switches will provide the clock to the stations

Figure 1: Lofar station Rack with a WR switch installed

# Conformity matrices

|  | Function | Minimum Requirement |  |  |  |  |  |  | Confirm Yes/no |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| R1 | Allan deviation | 10 sec 1.1·10-11 |  |  |  |  |  |  |  |
|  | (from clock coherence time) |  |  |  |  |  |  |  |  |
| R2 | Drift | <0.125ns / hour |  |  |  |  |  |  |  |
| R3⋆ | Clock inputs | 1PPS |  |  |  |  |  |  |  |
|  |  | 10MHz |  |  |  |  |  |  |  |
|  |  | SMA type prefered |  |  |  |  |  |  |  |
| R4 | Management inputs | USB terminal connection |  |  |  |  |  |  |  |
|  |  | RJ45 1000BASE-T |  |  |  |  |  |  |  |
| R5⋆ | Outputs | 1pps |  |  |  |  |  |  |  |
|  |  | 10MHz |  |  |  |  |  |  |  |
|  |  | SMA type prefered |  |  |  |  |  |  |  |
| ⋆ R3 and R5 are no requirements but are wishes. (Information Annex B) |  |  |  |  |  |  |  |  |  |
|  | Function | Minimum Requirement |  |  |  |  |  |  | Confirm Yes/no |
| R6 | Synchronization input, in order to be able to synchronize to i.e. GPS | PPS timing |  |  |  |  |  |  |  |
|  |  | Amplitude |  | Minimal 1.5 Vpeak-peak |  |  |  |  |  |
|  |  | Rise/fall time |  | <3nsec |  |  | <1usec |  |  |
|  |  | Jitter |  | <40psec |  |  | <1usec |  |  |
|  |  | Pulse width |  | 20-100usec |  |  | >100nsec |  |  |
|  |  | Alignment with 10MHz phase |  | <0.2ns (+/- 7 degrees) |  |  |  |  |  |
|  |  | output |  | 10MHz |  |  |  |  |  |
|  |  | Level (50 Ohm) |  | 13 | +/- 1 |  |  | dBm |  |
|  |  | Phase Noise [dBc/Hz] |  |  |  |  |  |  |  |
|  |  |  | WR switch LJ |  |  | WR switch |  |  |  |
|  |  | 1Hz | -95 |  |  | -80 |  |  |  |
|  |  | 100Hz | -115 |  |  | -100 |  |  |  |
|  |  | 100kHz | -138 |  |  | -130 |  |  |  |
| R7 | No. of SFP ports | Minimal 18 |  |  |  |  |  |  |  |
| R8 | Master availability | at least 99.9% |  |  |  |  |  |  |  |
| R9 | Slave availability | at least 99.5% |  |  |  |  |  |  |  |
| R10 | Monitoring and control | Ethernet Lan HTTP |  |  |  |  |  |  |  |
|  |  | SNMP protocol |  |  |  |  |  |  |  |
|  |  | Remote control via ethernet |  |  |  |  |  |  |  |
| R11 | Firmware version | V6.01 or newer |  |  |  |  |  |  |  |
|  | Function | Minimum Requirement |  |  |  |  |  |  | Confirm Yes/no |
| R12 | Hardware version | Based on CERN open hardware design, compatible with CERN hardware. Minimal hardware version V3.4 |  |  |  |  |  |  |  |
| R13 | PTP support | IEEE 1588-2019 with support for external equipment |  |  |  |  |  |  |  |
| R14 | Warranty | Warranty at least 2 years |  |  |  |  |  |  |  |
| R15 | Service | - In case of a fault the supplier shall reply within 48 hours - Supplier shall make a service proposal description of the provided service during the life time (maintenance and malfunctions) - Manuals and (electrical) schematics shall be provided by supplier |  |  |  |  |  |  |  |
| R16 | Service | Replacement equipment should be available within 48 hours for the first 5 years |  |  |  |  |  |  |  |
| R17 | EMC compliance | IEC EN 61000-6-2 Electromagnetic compatibility (EMC). Generic standards. Immunity standard for industrial environments. |  |  |  |  |  |  |  |
| R18 | Electro Magnetic Compatibility Information Technology Equipment | Compliant to the CISPR 22 / EN 55022 Class B standard |  |  |  |  |  |  |  |
| R19 | CE approval | The switches shall comply with Conformité Européenne (CE) standard |  |  |  |  |  |  |  |
| R20 | Size | Size should be 1U of 19 inch rack space and a maximum depth of 40 cm. (Figure 1 shows the position of a WR switch in a LOFAR station) |  |  |  |  |  |  |  |
| R21 | Firmware | Firmware version should be based on CERN open hardware |  |  |  |  |  |  |  |
| R22 | Operational environment conditions | 5-50 Celsius, 5-85% humidity |  |  |  |  |  |  |  |