## Appendix A05
Appendix A05, Conformity list 'Programme of Requirements and Preferences' Guidelines for the European Public Tender for a Scanning Electron Miscroscope, ref. TNO 2022 FPL-INK 103, dated feb 23, 2023.
By the following Programme of Requirements TNO will assess if Tenderer meets the relevant requirements as laid down in Section 6 and 8. The Tenderer must indicate for each item, the conformity for the requirements in the table below.
If to any of the requirements, as stated in section 8, are not met by Tenderer, he will be excluded for any further assessment. By answering a question with “compliant” will mean that in the event of awarding the contract, non-conformity to the requirement shall count as failure to meet an obligation from the Agreement to be concluded.
Requirement | Decription | Tenderer is compliant to the requirement | Reference in Tenderer's documentation where compliances with the requirement is described | Remark
8.1 | General System
Requirement 8.1.1 | The SEM system comprises an emitter, analysis chamber, with a sample stage and detectors, vacuum system and control system. | Click the drop-down box to select
Requirement 8.1.2 | The SEM must be CE certified and be compatible with Dutch work environment safety regulations (ARBO). | Click the drop-down box to select
Requirement 8.1.3 | It must be possible to transfer the samples from the argon glovebox to the ion milling equipment and the SEM for imaging without exposure to air. This is an essential requirement, and to be fully supplied by the tenderer. The tenderer is expected to provide a detailed description of the proposed solution, including the components and transfer procedures. | Click the drop-down box to select
Preference 8.1.4 | Provide full integration of sample preparation unit, airtight transport system and SEM&EDX system. | Click the drop-down box to select
Preference 8.1.5 | Provide a reference of similar system with airtight sample handling (line 10) provided earlier to a customer. | Click the drop-down box to select
8.2 | Emitter
Requirement 8.2.1 | The emitter must be a Schottky or cold field emitter. | Click the drop-down box to select
Requirement 8.2.2 | The emitter must have a guaranteed lifetime of at least 3 years | Click the drop-down box to select
Preference 8.2.3 | The emitter will have a guaranteed lifetime of more than 3 years | Click the drop-down box to select
Requirement 8.2.4 | The sample preparation must be capable to prepare air-senitive samples with the following sample size: | Click the drop-down box to select
• Cross sectioning typical sample sizes of:
o Width minimum of 10 mm
o Length of the sample: min 5 mm
o Thickness of min 5 mm
8.3 | Electron optics and beam current
Requirement 8.3.1 | The acceleration voltage (AV) range must be at least between 0,5 and 30 kV without decelaration or boosting technology | Click the drop-down box to select
Note: For some samples decelaration or booster technology is not recommended because of possible electrical effects like arching which may impair resolution in different ways.
Requirement 8.3.2 | It must be possible to perform automatic beam alignment. | Click the drop-down box to select
Requirement 8.3.3 | It must be possible to perform electronical and/or mechanical alignment for fine tuning by TNO tool owners/experienced users | Click the drop-down box to select
Requirement 8.3.4 | The system must include a beam current detector | Click the drop-down box to select
Requirement 8.3.5 | The beam current is stable within 10% or better at 1, 3 and 15 kV | Click the drop-down box to select
Requirement 8.3.6 | The field of view must be at least 3x3 mm at analytical working distance | Click the drop-down box to select
Requirement 8.3.7 | It must be possible to perform simultaneous viewing and imaging of all available detectors | Click the drop-down box to select
Requirement 8.3.8 | It must be possible to randomly mix different detector signals | Click the drop-down box to select
Preference 8.3.9 | The AV can be 100V or less, with or without deceleration. Specify attainable resolution on standard samples. | Click the drop-down box to select
8.4 | Detectors
Requirement 8.4.1 | Side SE detector. | Click the drop-down box to select
Requirement 8.4.2 | (semi) in-lens (upper) detector(s) for LA and HA electron detection | Click the drop-down box to select
Requirement 8.4.3 | Retractable segmented BSE detector | Click the drop-down box to select
Preference 8.4.4 | Retractable BSE with a high sensitivity allowing for live image viewing. | Click the drop-down box to select
Preference 8.4.5 | Place to retrofit detector for Li-ion detection in the future (> 3 years) | Click the drop-down box to select
8.5 | Resolution and samples evaluation
Requirement 8.5.1 | The SEM system must have a lateral resolution of 1 nm or better | Click the drop-down box to select
referring to a standard e.g. Au on C at different kV (0.5, 1, 2, 5, 10, 15 and 30kV)
Requirement 8.5.2 | The SEM system will have a lateral resolution of at least 0.7 nm referring to a standard e.g. Au on C at different kV (1, 2, 5, 10, 15 and 30kV) | Click the drop-down box to select
Preference 8.5.3 | The SEM system with a better lateral resolution than 0.7 nm referring to a standard e.g. Au on C at different kV (1, 2, 5, 10, 15 and 30kV) | Click the drop-down box to select
Requirement 8.5.4 | The SEM system must have a height resolution of 10 nm or better at different kV | Click the drop-down box to select
Note: The supplier must show this with their standard sample.
Preference 8.5.5 | The SEM system will have a height resolution of 5 nm or better referring to a standard. | Click the drop-down box to select
Preference 8.5.6 | Samples evaluation; Provided cross section and top view images will be assessed by the evaluation team on sharpness and resolution (Appendix A07) | Click the drop-down box to select
8.6 | Chamber & stage requirements
Requirement 8.6.1 | The SEM chamber must be equipped with an IR camera for viewing the sample(s) and position with respect to sensors in the SEM during operation with SE detectors. | Click the drop-down box to select
Requirement 8.6.2 | The SEM chamber has air lock operation for inserting samples | Click the drop-down box to select
Note: This is required for air-tight insertion of air sensitive samples
Requirement 8.6.3 | The system must be able to reach an analysis vacuum level within 5 minutes after loading a new non-porous sample. | Click the drop-down box to select
Requirement 8.6.4 | The SEM has a full stage navigation system using in-stand colour imaging and/or external colour image for stage calibration for navigating to areas of interest | Click the drop-down box to select
Requirement 8.6.5 | The SEM can perform stage navigation using acquired SEM images | Click the drop-down box to select
Requirement 8.6.6 | A sample of at least 100x100 mm2 can be loaded onto the sample stage. Stage navigation must accommodate full investigation of this sample size. | Click the drop-down box to select
Requirement 8.6.7 | It must be possible to load 5 2x2 cm2 samples | Click the drop-down box to select
Note: TNO requires this for fast investigation of samples
Requirement 8.6.8 | Minimum allowed sample height: 30 mm | Click the drop-down box to select
Note: The sample is loaded into a sample holder supplied by tenderer.
Requirement 8.6.9 | Stage relocaton reproducibility is at least within +/- 5 µm | Click the drop-down box to select
Requirement 8.6.10 | Stage tilting is minimum 70°, without pre-tilt sample holder. | Click the drop-down box to select
Preference 8.6.11 | Stage tilting is possible to 90° without pre-tilt sample holder | Click the drop-down box to select
Requirement 8.6.12 | The system must have eucentric tilt stage | Click the drop-down box to select
Requirement 8.6.13 | The control system must have tilt correction | Click the drop-down box to select
8.7 | Analysis system
Requirement 8.7.1 | The EDX system must be able to detect the presence and location of elements down to lithium | Click the drop-down box to select
Note: Main elements of interest are:
Al, Ni, Ir, Pt, Ta, Ti, Ru, Gd, Ce, Sn, O, N, C, Cl, F, Zn, In, Mn, Co and Fe.
Requirement 8.7.2 | The EDX system must be able to perform automated element mapping over an area of at least 10x10mm2 | Click the drop-down box to select
Requirement 8.7.3 | The EDX system must be able to perform EDX-line scans over a distance of up to at least 10mm, both directly on the sample and as extracted spectra from acquired mapping | Click the drop-down box to select
Requirement 8.7.4 | It must be possible to define the stepsize of both linescans and EDX mapping | Click the drop-down box to select
Requirement 8.7.5 | The EDX software must be coupled to the SEM and be able to take control of the SEM. | Click the drop-down box to select
Preference 8.7.6 | The EDX system must be able to detect the presence and location of elements down to lithium | Click the drop-down box to select
Note: Main elements of interest are: Li
Preference 8.7.7 | It is possible to perform stitching of the element mapping. | Click the drop-down box to select
Preference 8.7.8 | Fast EDX acquisition for live EDX images should be possible | Click the drop-down box to select
Preference 8.7.9 | For ‘high resolution’, EDX acquisition a short working distance for simultaneous ‘high resolution’ imaging should be possible. Specify min WD for analysis. | Click the drop-down box to select
8.8 | Aspect Control system and software
Requirement 8.8.1 | The SEM must be equipped with such SEM and sample protection measures that during power loss or elimination of cooling systems, the SEM will be safely shut down, without loss of vacuum within 12 hrs. | Click the drop-down box to select
Requirement 8.8.2 | The SEM must be equipped with protection and warning systems/sensors to prevent damaging or hitting detectors during normal operation. This to prevent both damage to SEM detectors and samples. Please describe the protection and warning systems. | Click the drop-down box to select
Requirement 8.8.3 | The SEM will be provided with: | Click the drop-down box to select
• PC
• Keyboard
• Mouse
• Minimum of two monitors minimal 21” (wide screen) with high gray contrast (e.g. EIZO)
Requirement 8.8.4 | The operation software must be in English. | Click the drop-down box to select
Requirement 8.8.5 | The operating system must be the latest Windows version, with possiblity to upgrade (future proof) | Click the drop-down box to select
Requirement 8.8.6 | All software should be updated throughout the lifetime of the SEM, but at least 10 years after instrument installation. | Click the drop-down box to select
Requirement 8.8.7 | It must be possible to perform Windows safety updates at any time. | Click the drop-down box to select
Requirement 8.8.8 | The system can be connected with the internet. | Click the drop-down box to select
8.9 | Location
Requirement 8.9.1 | With respect to functional operation within specifications the suppliers will perform a check of the intended location of the SEM with respect to e.g. vibrations, magnetic fields, etc. The supplier will then guarantee SEM and EDX operation within specifications. | Click the drop-down box to select
Requirement 8.9.2 | The supplier will perform a Site Acceptance Test (SAT) as final stage of the delivery. The SEM shall be delivered with all the installation software required to (re-)install the control system of the SEM. Remark: - This includes all the programs installed in the PC. - The installation process shall be documented in a manual that is part of the data pack to be delivered. | Click the drop-down box to select
8.10 | Sample preparation system
Requirement 8.10.1 | Ion milling device for cross sectional investigation of thin layers on substrates, NMC’s, CNT’s is required. It must be capable of deformation free and distortion free preparation. | Click the drop-down box to select
Requirement 8.10.2 | Area of analysis must be at least 1mm wide | Click the drop-down box to select
Requirement 8.10.3 | Location for cross-section can be attained within 1 µm accuracy. | Click the drop-down box to select
Requirement 8.10.4 | It must be possible to prepare air-sensitive samples and air tight mounting, preparation and transport to the SEM is required. | Click the drop-down box to select
Requirement 8.10.5 | The sample preparation system must be capable to prepare samples with the following sample size for non-air sensitive samples: | Click the drop-down box to select
• Cross sectioning typical sample sizes of:
o Width of the sample: up to at least 20 mm
o Thickness up to at least 5 mm
o Length of up to at least 5mm
Preference 8.10.6 | In situ sample cooling/conditioning during milling/preparation | Click the drop-down box to select
Preference 8.10.7 | The samples preparation system can perform post-polishing surface treatments | Click the drop-down box to select
Preference 8.10.8 | Describe the procedure for loading the sample into the sample preparation system, user friendliness, etc. etc. | Click the drop-down box to select
8.11 | Delivery, after-sales & payment conditions
Requirement 8.11.1 | The system will be delivered with the following spare parts: | Click the drop-down box to select
- List to be added
- CE certification
- Set of standard consumable for minimal 1 year
- Documentation will be delivered in English
Requirement 8.11.2 | Tenderer will provide a standard maintenance contract for the SEM supplemented with maturities for several years which states the maintenance activities and costs. Tenderer describes its calamity procedure within and outside the warranty periode including response time and complaint registration. | Click the drop-down box to select
Requirement 8.11.3 | The payment schedule will be as follow: | Click the drop-down box to select
- 20% at order, after receipt of a bank guarantee covering 40% of the total amount;
- 20% after approval Critical Design Review (CDR);
- 30% after delivery at location TNO Holst Centre of the total system.
- 30% after installation and commissioning On-site and Site Acceptance Test (SAT) including test periods, approved by TNO Management and TNO Procurement. At this moment TNO will return the bank guarantee.
Note: The Bank guarantee and the bank who issues the Bank guarantee have to be approved by TNO (minimal A rating). TNO must have the right to draw the bank guarantee in any and all cases where Tenderer falls short in fulfilling its obligations under the Contract. The validity of the bank guarantee may not expire until approval of the SAT by TNO Management and Procurement.
Requirement 8.11.4 | Supplier provides Basic (after SAT) and Advanced Operator Training later in time and adds information about this to this tender. | Click the drop-down box to select
Preference 8.11.5 | Supplier with shortest delivery time in days counted from date contract signing until completed SAT. | Click the drop-down box to select
8.12 | Export Control
Requirement 8.12.1 | The Supplier will immediately notify TNO if it is prescribed by export control laws and regulations within the context of the Agreement that an export or import licence is required from a government or any other government body or if the execution of the Agreement is otherwise restricted or prohibited by export control laws and regulations. 'Export control laws and regulations' as referred to in this article should be interpreted as all EU and USA laws and regulations applicable at the time for the regulation of export and import of military goods and dual-use goods (together also referred to as 'strategic goods') and chemical substances and also countries or persons against which sanctions have been imposed by the Security Council of the United Nations (SCUN), by the European Union (EU) or the Dutch government. | Click the drop-down box to select

## validatie (hidden)
Click the drop-down box to select
Yes, compliant to requirement
No, not compliant to requirement
note: dit blad 'verbergen' indien je dit gaat delen intern/extern (behoudens aan inkoop collega's)