## Appendix A05
Appendix A05 Conformity list 'Programme of Requirements' Applicable to the European Public Tender for: the Realization and Delivery of Möbius2 Modules DSM and ASM, ref. 2023 FPL/INK 001
By the following Programme of Requirements TNO will assess if Tenderer meets the relevant requirements as laid down in Chapter 6 and 8. The Tenderer must indicate for each item, the conformity for the requirements in the table below.
If any of the requirements stated in Chapter 8 are not met, Tenderer will be excluded from any further assessment. Answering a question with “compliant” will mean that in the event of awarding the contract, non-conformity to the requirement shall count as breach of contract.
Requirement | Description | Tenderer is compliant to the requirement | Reference in Tenderer's documentation where compliances with the requirement is described | Remark
8.1.1 | The DSM and ASM will each have their own control cabinet that is able to communicate with one central control unit. | Click the drop-down box to select
8.1.2 | The DSM and ASM shall clearly notify the operator of its current operation status, including alarms and warnings and communicate to main central control unit. | Click the drop-down box to select
8.1.3 | The DSM and ASM shall be designed for a lifetime of 10 years. Components with shorter lifetime will be documented and marked. | Click the drop-down box to select
8.1.4 | The flow rate of liquid solvents into the DSM and ASM shall be between 1-10l/h. These liquids have a density of 0.7 – 1 kg/l and a viscosity range between acetone and water. | Click the drop-down box to select
8.1.5 | The DSM and ASM shall be capable of safely handling flammable (cat. II) organic solvents. | Click the drop-down box to select
8.1.6 | The DSM and ASM shall be able to operate in a temperature range from 25(ambient)- 150 degC. | Click the drop-down box to select
8.1.7 | The DSM and ASM shall be able to operate from 1 bar(g), to a pressure of 15 bar(g) at the maximum operating temperature. | Click the drop-down box to select
8.1.8 | The DSM and ASM shall be skid mounted. | Click the drop-down box to select
8.1.9 | The DSM and ASM skid will have a maximum height of 2 m. | Click the drop-down box to select
8.1.10 | One of the dimensions (either the width or length) of the DSM and ASM skid will be at most 1 m. | Click the drop-down box to select
8.1.11 | One of the dimensions (either the width or length) of the DSM and ASM skid will be at most 1.7 m. | Click the drop-down box to select
8.1.12 | The DSM and ASM shall comply with the maximum floor loading of 10 kN/m2. | Click the drop-down box to select
8.1.13 | The DSM and ASM shall be liftable using a standard pallet truck pushed through the narrowest dimension of the module. | Click the drop-down box to select
8.1.14 | The DSM and ASM shall each be powered by a standard three phase electrical connection (400V). The connector should be 3P+N+E with a maximum current rating of 16 or 32 A. | Click the drop-down box to select
8.1.15 | The DSM and ASM skids shall be equipped with 2 suction points, 250 mm diameter, to be connected to the RSCG ventilation. | Click the drop-down box to select
8.1.16 | The DSM and ASM skids shall be designed in such a way that they minimize solvent exposure to the operator during filling, cleaning, draining and operation. | Click the drop-down box to select
8.1.17 | Those parts in the DSM and ASM that need it, shall be easily accessible for maintenance, replacement and cleaning. | Click the drop-down box to select
8.1.18 | The DSM and ASM shall meet all relevant general Dutch laws for experimental equipment, given that the site has all relevant permits. | Click the drop-down box to select
8.1.19 | The DSM and ASM shall each measure the concentration of organic vapors in one of the ventilation connection with a LEL sensor for safety reasons. The value is send to the central control unit for control and safety. | Click the drop-down box to select
8.1.20 | The DSM and ASM shall each measure the flow (m3/h) in the ventilation connection for safety reasons. This value shall be a read-out electronically. | Click the drop-down box to select
8.1.21 | The equipment and piping inside the DSM and ASM shall be able to be purged with nitrogen via an automated process. This is not intended for the cupboard itself. | Click the drop-down box to select
8.1.22 | The hot equipment and piping inside the DSM and ASM shall be insulated to minimize the loss of heat. | Click the drop-down box to select
8.1.23 | The DSM and ASM shall be placed inside closed setup frame, equiped with doors and windows for easy access to the setup. See photo of current Möbius1.0 set-up for reference of what TNO envisions. | Click the drop-down box to select
8.1.24 | The DSM and ASM shall both have one inlet for Nitrogen and two outlets. One should be used as an inlet for Nitrogen, a second should be an outlet towards the following module and the third at an outlet towards the 250mm ventilation outlet. The connections shall be made by Swagelok and all piping should be in imperial measurements. | Click the drop-down box to select
8.1.25 | The DSM and ASM shall have sensors that measure temperature(accuracy=0.1 degC), pressure (accuracy=+/-0.25%) of and flow rate pressure (accuracy=+/-0.5%) where it is important in the module for safety and operation of the module. The temperature and pressure should be controlled remotely. | Click the drop-down box to select
8.1.26 | The DAC of the DSM and ASM shall read-out the various sensors and other inputs at a sample rate of 1s. | Click the drop-down box to select
8.1.27 | The DAC of the DSM and ASM shall have configurable alarms and configurable procedures regarding those alarms. | Click the drop-down box to select
8.1.28 | Data acquisition and control (DAC) system be located in a separate cabinet mounted to the frame with locable door. The box should be mounted in such a way that each module can be right next to another module on both sides. | Click the drop-down box to select
8.1.29 | The DAC of the DSM and ASM shall have 4-20mA output signals for all installed sensors (waivers are possible for the type of output signal after good motivation and approval by TNO) | Click the drop-down box to select
8.1.30 | The DSM and ASM skids should be fitted out with wheels for easy maneuvering in the RSCG lab but also be able to be stationary for years without losing maneuverability. | Click the drop-down box to select
8.1.31 | All equipment and piping in the DSM and ASM that are in direct contact with the solvent or plastic feed shall be made of SS 316L, and for the entire system there is a high preference for items to be of item systems profile 8. | Click the drop-down box to select
8.1.32 | Delivery time – It is preferred by TNO that delivery takes place within 20 weeks, after TNO submitted a purchase order. Tenderer is requested to provide the fastest, most reliable delivery time that Tenderer can guarantee, without any additional costs for TNO. Note: not achieving the offered delivery time will lead to a penalty discount of 20% of the total order amount and shall be immediately due and payable and eligible for setoff and will replace clause 8.4 of the General Purchasing Conditions TNO – 2022. | Click the drop-down box to select
8.1.33 | Terms and conditions of payment: - 30% at order, after receipt of a bank guarantee covering 80% of the total amount. - 50% after delivery at location TNO. - 10% after succesfull Site Acceptance Test (SAT) - 10% after the test periods and approvement by TNO’s management. The Tenderer to whom the Contract will be awarded must provide a bank guarantee to TNO, covering at least 80% of the total price. The Bank guarantee and the bank who issues the Bank guarantee have to be approved by TNO (minimal A rating). TNO must have the right to draw the bank guarantee in any and all cases where Tenderer falls short in fulfilling its obligations under the Contract. The validity of the bank guarantee may not expire until approval of the SAT by TNO. | Click the drop-down box to select
8.1.34 | The DSM and ASM shall be delivered with all relevant documentation. This documentation must include at least the following: - Installation documentation - Design Package, including the P&ID of the system - List of components/equipment - Instruction manuals of the components and equipment where applicable - Maintenance manual including all required technical documentation for Preventive and Corrective Maintenance of the components and equipment where applicable. - Description of the safety features implementated - Certificates and manuals of all third-party hardware that is part of the system. All the documents should be delivered digitally in English and are inspected during FAT. | Click the drop-down box to select
8.1.35 | The system must be fully automated during normal operation and safely shutdown without operator intervention in case of a safety deviation for instance. Also, all critical components ( to be discussed with TNO at a later stage) shall be CE certified. | Click the drop-down box to select
8.1.36 | The DAC of the DMS and ASM shall be LabVIEW software based and Modbus TCP or equivalent serial communication protocol. All units shall be read-out and logged using SI based units such as °C, barg, l/h, m3/h, g/min, %LEL, . | Click the drop-down box to select
8.1.37 | A Factory Acceptance Test must be performed by the tenderer at the location of the tenderer. As described on page 18, as part of the tender bid, we require the Tenderer to submit a proposal for the FAT plan. TNO states that the minimum requirements for checkpoints that must be incorporated in the FAT plan are: • Both modules will be checked for damages; • The proper functioning of the modules will be checked; o Mechanical functions, o Electrical, - and switching functions, o User interface(s), including software, controls and data storage, o Safety functions o A mock experiment to check the combined equipment in each module. • The FAT includes testing all functions and the requirements as listed in the schedule of requirements section 8, and all other functions offered by the tenderer. • The presence and quality of the required documentation will be checked. The details of the FAT shall be defined and agreed between TNO and the final Tenderer. | Click the drop-down box to select
8.1.38 | A full Site Acceptance Test on-site at TNO, RCSG Rijswijk, shall be performed by the tenderer in cooperation with TNO. As part of the tender bid, we require the Tenderer to submit a proposal for the SAT plan TNO states that the minimum requirements for checkpoints that must be incorporated in the SAT plan are: • Both modules will be checked for damages; • Any open issues from the FAT will be checked • A limited repeat of the proper functioning check of the modules during the FAT; Execution of a mock experiment to check the combined equipment in each module; • The presence of the required documentation will be checked. | Click the drop-down box to select
8.2.1 | The DSM shall be built around two CSTR type autoclaves. | Click the drop-down box to select
8.2.2 | The operator shall be able to choose between each of the autoclaves, or to operate them in parallel or both of them in series. . This should be done with actuated valves. | Click the drop-down box to select
8.2.3 | The DSM shall be capable of handling various forms of plastic feed by changing the type of the stirrer and its rotational speed. | Click the drop-down box to select
8.2.4 | It shall be possible to load the DSM with various forms of plastic feed. Flakes with maximum size of up to 5mm x 3mm, roundish pellets and chunks of up to 10 mm diameter. | Click the drop-down box to select
8.2.5 | Solids, including the plastic feed, shall be able to be added to the each autoclave of the DSM under pressure with overpressure. | Click the drop-down box to select
8.2.6 | Each of the autoclaves shall have a size of 7.5l. | Click the drop-down box to select
8.2.7 | The DSM shall be able to maintain the temperature of the dissolution process in each autoclave at a set temperature up to 150 degC by using a jacket. The heating fluid will be supplied by another module supplied by TNO. | Click the drop-down box to select
8.2.8 | Each autoclave of the DSM shall be able to stir the plastic feed in the CSTR at a speed of 1 to 200 RPM. The content of the autoclaves can be assumed to be either a slurry with the particles as described in Requirement 8.2.4 in the solvent described in Requirement 8.1.5 or a solution with a viscosity range off 0.7-10 mPa*s. | Click the drop-down box to select
8.2.9 | Each autoclave of the DSM shall be able to be pressurized with nitrogen up to the maximum operating pressure. | Click the drop-down box to select
8.2.10 | Apart from the main polymer solution exit, each autoclave of the DSM shall have a secondary exit for cleaning and flushing purposes. | Click the drop-down box to select
8.2.11 | Each autoclave of the DSM shall have a transparent section at the exit of or in the dissolution section. | Click the drop-down box to select
8.2.12 | The DSM should be fitted with a sampling port. | Click the drop-down box to select
8.2.13 | All piping shall be traced. | Click the drop-down box to select
8.2.14 | The DSM should be fitted with a coarse filter before the polymer solution leaves the DSM. | Click the drop-down box to select
8.3.1 | The ASM shall contain two CSTR type autoclavlse. | Click the drop-down box to select
8.3.2 | The autoclaves ASM shall have a 10l. | Click the drop-down box to select
8.3.3 | The ASM shall be able to maintain the temperature of the precipitation process in the autoclave at a set temperature from 35(ambient) to 150 degC by using a jacket. The heating fluid will be supplied by another module supplied by TNO. | Click the drop-down box to select
8.3.4 | The ASM shall measure, read-out and log the temperature and pressure inside the autoclaves. | Click the drop-down box to select
8.3.5 | The ASM shall be able to measure, read-out and log the amount of liquid in the autoclaves. | Click the drop-down box to select
8.3.6 | The ASM shall be able to stir the feed in the autoclaves at a speed of 1 to 200 RPM. | Click the drop-down box to select
8.3.7 | The autoclaves of the ASM shall be able to be pressurized with nitrogen up to the maximum operating pressure. | Click the drop-down box to select
8.3.8 | Apart from the main polymer solution exit, the autoclave of the ASM shall have a secondary exit for cleaning and flushing purposes. | Click the drop-down box to select
8.3.9 | The ASM shall have a transparent section at the exit of or in the autoclaves. | Click the drop-down box to select
8.3.10 | The ASM should be fitted with a sampling port on the autoclaves. | Click the drop-down box to select
8.3.11 | All piping that is upstream of the autoclaves that contains the solvent solution, anti solvent or heating, as stated in the P&ID shall be traced. | Click the drop-down box to select

## validatie (hidden)
Click the drop-down box to select
Yes, compliant to requirement
No, not compliant to requirement
note: dit blad 'verbergen' indien je dit gaat delen intern/extern (behoudens aan inkoop collega's)