TITLE: METH044_Nanopore dual barcoding with kit 14

| STATUS | Under review |
| --- | --- |
| LAST UPDATE |  |
| DOCUMENT OWNER | Maarten van der Salm |
| APPROVER / GOEDKEURDER |  |

- INTRODUCTION

  - SCOPE

This protocol describes how to carry out dual barcoding of amplicons using the Native Barcoding Kit 24 V14 (SQK-NBD114.24). There are 24 unique barcodes available, allowing the user to pool up to 24 different plates in one sequencing experiment.

For applications that differ from this protocol, please refer to chapter 7 or the available procedures on https://nanoporetech.com.

  - OBJECTIVE

This protocol allows massively parallel barcode sequencing of up to 2,304 samples (amplicons) on a single flow cell. It uses both the PCR Barcoding Primers and the Native Barcoding kit: . First, up to 96 barcodes are added to the DNA ends with PCR. Up to 24 pools of 96 samples can then have a secondary barcode added through the ligation of one of 24 native barcodes. All primary pools can then be combined into a single library and sequenced.

  - DEFINITIONS AND ABBREVIATIONS

Definitions and abbreviations used in this procedure

  - STATUS

The status and management of this document are regulated in the AODocs platform, the foundation wide accessibility is provided on N=Info.

- RESPONSIBILITIES & PEOPLE

SHEET025_WORKFLOW_OVERVIEW_SOPs_SD

  - RESPONSIBILITIES

Mention the responsibilities for every function involved in this procedure

| FUNCTION | ROLES and RESPONSIBILITIES |
| --- | --- |
| (Senior) Research Technician | When trained the technician should be able to execute the workflow, give support, and train new staff. |
| Lab workers (students, guests) | Should familiarise with the protocols and technique before start of activities regarding Nanopore sequencing. And make a reservation for a flow cell. details will be noted in the “Inventory list ONT materials and reagents” |
| Head Laboratories | Responsible for declaration of the material costs of the workflow. |

  - LIST OF PEOPLE ADDRESSED

SHEET010_OVERVIEW_ROOM_MANAGERS_&_TELEPHONE NUMBERS

Optionally include specific people e.g. apparatus managers.

FLOW CHART

1. DNA extraction and first PCR (PCR 1)

2. Barcoding PCR (PCR 2: BC01-BC96)

3. End-prep: end-repair / dA-tailing

4. Native Barcode Ligation (NB01-24)

5. Ligation of Native Adapter (NA)

6. Priming and loading the SpotON flow cell (up to 24 plates)

- ADDITIONAL AIDS AND RESOURCES

  - REAGENTS, SOLUTIONS, KITS, CHEMICALS AND WASTE

    - Buffers

- Nuclease-free water / MilliQ

- Freshly prepared 80% ethanol in nuclease-free water/MilliQ

    - Plastics

- 1.5 ml Eppendorf DNA LoBind tubes

- 0.2 ml thin-walled PCR tubes

- Corning® Round Bottom plates - 96 well (VWR, 734-1641)

- VWR skirted 96 well PCR plate - 96 well (VWR, 211-0297)

### PCR oligo’s

- PCR1: Primers including ONT adapter sequence

| Name | Direction | Sequence 5’ - 3’ |
| --- | --- | --- |
| ONT_adapter Forward | F | TTTCTGTTGGTGCTGATATTGC[primer] |
| ONT_adapter Reverse | R | ACTTGCCTGTCGCTCTATCTTC[primer] |

- PCR2: PCR Barcoding Primers BC01-BC96, link to technical information

    - Bioke

- NucleoMag NGS Clean-up and Size Select beads (Bioke 744970)

- LongAmp Taq 2X Master Mix (Bioke, M0287)

- NEB Blunt/TA Ligase Master Mix (Bioke, M0367)

- NEBNext Ultra II End repair/dA-tailing Module (Bioke, E7546)

- NEBNext Quick Ligation Module (Bioke, E6056)

    - Nanopore items/kits

- Native Barcoding Kit 24 V14 (SQK-NBD114.24)

- Flow Cell (R10.4.1),  FLO-MIN114

    - Optional expansion kits

- Flow Cell Priming Kit V14,  EXP-FLP004

- Sequencing Auxiliary Vials V14, EXP-AUX003

- Flow Cell Wash Kit, EXP-WSH004

- Native Barcoding Expansion (EXP-NBA114)

- SFB Expansion, EXP-SFB001

- Additional information: https://store.nanoporetech.com/eu/expansion-packs.html

  - EQUIPMENT

- PCR machine/Thermo cycler

- C.wash (bead clean up after PCR1)

- Opentrons OT2 (pooling after PCR2)

- Tapestation (D5000)

- Qubit (flex) fluorometer (HS & BR  dsDNA kits)

- MinION Mk1B, GridION Mk1 or MinION Mk1C

- METHOD

#### Before starting the method

- Acquire a flow cell by contacting the head of laboratories or a technician.

- Make sure that sufficient reagents are available. If not, always inform a technician when reagents are low or out of stock.

- Create a folder under the correct year in the NGS projects folder → Ask a technician to create a shared project folder if not already available.

- Make a copy of the following documents and fill in the sheets according to your project information :

  - Data analysis: Nanopore data input sheet

  - Workflow: “Template ONT Dual Barcoding checklist”

Tips before starting the method

- After flow cell wash the number of pores still available differs between the in-run flow cell check vs the regular flow cell check. This is due to a different voltage being used.

- Small fragments of DNA (100-250bp) cause the Flow Cell to deteriorate faster, potentially because small fragments are more strenuous on the Nanopores. The higher exchange rate of motor proteins (hypothetical) causes more damage to the nanopores.

- An additional bead clean-up might be performed in order to wash away small fragments. Always perform a final QC on the tapestation.

- Try to minimise the vortexing of components used in the ligation process.

- Do not vortex any enzymes used in the process.

- Do not vortex the buffers used in loading the Flow Cell, only mix by pipetting. This will minimise the introduction of micro-bubbles/air onto the sensor array. Which in turn forms bubbles in the sensor array.

#### Sample prep and first PCR (PCR 1)

- Extract the DNA using a valid method. - Used extraction dependent on the sample matrix and/or project. - Perform the first PCR. - Always use primers with the ONT adapters on the 5’ ends (4.1.3.) - See METH008 (PCR and Gel electrophoresis) for further instructions. - use the PCR template for the PCR setup. - Check amplification success of your PCR products using agarose gel electrophoresis. - PCR products have to be checked on an E-Gel, according to: WI_SP0004_E-GEL - Or use a handmade gel (METH008) - Perform a Bead Clean-up using the MN beads* and elute in MilliQ water. - *Bead Clean-up ratio is dependent on the amplicon size: Adjust if necessary. - Perform Bead Clean-Up using either the manual Stamp method (METH030 from 4.3) or the C.wash (WI_C.wash). - tip: leave the eluate in the round-bottom plate to reduce plastic waste. Clean PCR products can be stored at 4°C together with the magnetic beads. Use a magnetic separator before pipetting the clean PCR products. - Optional: check whether the primer dimers are successfully removed by using the tape station D5000 kit. Test approximately 2 or 3 samples.

#### Barcoding PCR (PCR 2)

#### Barcoding PCR (PCR 2)

- The pre-mixed Barcodes (BC01 - BC96) are located in the Pre-PCR Lab: L03.01, FRI-01. To prevent cross contamination, spin down the plate in a plate centrifuge for about 2 minutes at max speed.

- Contact a technician if the pre-mixed barcodes are not available.

- Barcode layout:

Set up a barcoding PCR plate for each sample:

- Add the following reagents per well

| Component | Volume | - Mix by pipetting, seal and spin down - Add the PCR product in the Post PCR lab - The PCR reactions have a final volume of 10µL * One of BC1-BC96, ready-to-use (0,4 µL of the fw and Rev primer at 10µM + 1,7 µL MQ) |
| --- | --- | --- |
| PCR Barcodes * | 2,5 µL |  |
| LongAmp Taq 2x master mix | 5 µL |  |
| PCR product (add in postlab) | 2,5 µL |  |

- Use the following PCR program for the 2nd PCR:

| Step | Temperature | Time (min:sec) | Cycles |
| --- | --- | --- | --- |
| Initial denaturation | 95°C | 03:00 | 1 |
| Denaturation | 95°C | 00:15 | #12 - 15x |
| Annealing | 62°C | 00:15 |  |
| Extension | 65°C | 00:50* |  |
| Final extension | 65°C | 03:00 | 1 |
| Hold | 15°C | ∞ | 1 |

* Determine the extension time based on the amplicon size, use 50 seconds per kb.

# Cycles: The number of cycles is dependent on the amount of PCR product used, however, in most cases 12 cycles is sufficient. If necessary the number of cycles could be lowered.

- Pool 1 µL per sample and proceed to the bead-clean-up.If the samples are not quantified/checked using the Fragment Analyzer.

  - Tip: This step can be performed using the  opentrons OT2 liquid handler.

  - Target plate for OT2 liquid handler: Corning® Round Bottom plates - 96 well.

- Purify the DNA pools with a MN bead-clean-up*

  - *bead-clean-up ratio is dependent on the amplicon size: Ratio overview

  - Allow the DNA to bind to beads for 5 minutes at Room Temperature.

  - Put the tubes on a magnet and discard supernatant.

  - Wash beads twice with 200 µL fresh 80% ethanol without disturbing the pellet.

  - Spin down and place the tube back on the magnet. Pipette off any residual ethanol.

  - Allow to dry for ~ 30 seconds, but do not dry the pellet to the point of cracking.

  - Remove the tube from the magnetic rack and resuspend the pellet in Nuclease-free water. Incubate for 2 minutes at RT.

    - Elution volume is dependent on the starting volume.

  - Pellet the beads on a magnet until the eluate is clear and colourless, for at least 1 minute.

  - Remove and retain the eluate into a clean 1.5 ml Eppendorf DNA LoBind tube or  0.2 ml thin-walled PCR tube.

- Use the Qubit dsDNA BR assay to quantify the pool (QUBIT FLEX)

  - Use 2 µL per DNA pool for the dsDNA BR assay. Use the build in calculator to determine the nM concentration (1nM = 1fmol/µL)

- Take forward 200 fmol amplicon DNA in 12.5 μL (130 ng for 1 kb amplicons) MilliQ to the next step.

  - Example: If your pool is 142nM = 142 fmol/µL, the amount to take forward is: 200/142 = ~1.4µL

## End-prep: end-repair / dA-tailing

#### Step specific reagents

- NEBNext Ultra II End Prep Enzyme Mix, thaw and mix by pipetting, and place on ice/cooling block.

- NEBNext Ultra II End Prep Reaction Buffer, thaw and mix by vortexing.

    - Method

- Transfer 200 fmol (0.2pmol) DNA of the sample into a 0.2 ml thin-walled PCR tube.

  - Mix thoroughly by pipetting up and down several times.

  - Spin down briefly in a microfuge.

- In a 0.2 ml thin-walled PCR strip-tube, mix the following:

| Component | Volume |
| --- | --- |
| Pool of barcoded amplicon DNA | 12.5 µL |
| Ultra II End-prep reaction buffer | 1,75 µL * |
| Ultra II End-prep enzyme mix | 0.75 µL* |

*Make a mastermix of the enzyme mix and reaction buffer for the total amount of pools, and use 2,5µL per pool/tube.

- Ensure the components are thoroughly mixed by pipetting.

- Using a thermal cycler, incubate at 20°C for 5 minutes and 65°C for 5 minutes.

- Transfer the pools to a clean column of a Corning® Round Bottom plate.

Purify DNA with a MN 1x bead clean up

- Resuspend MN beads by vortexing.

- Add 15µL of resuspended MN beads to the end-prep reaction and mix by pipetting.

- Allow DNA to bind to beads for 5 minutes at Room temperature.

- Put the tubes on a magnet and discard supernatant.

- Wash beads twice with 200 µL fresh 80% ethanol without disturbing the pellet.

- Spin down and place the tube back on the magnet. Pipette off any residual ethanol. Allow to dry for ~ 30 seconds, but do not dry the pellet to the point of cracking.

- Remove the tube from the magnetic rack and resuspend the pellet in 12µL Nuclease-free water. Incubate for 2 minutes at RT.

- Pellet the beads on a magnet until the eluate is clear and colourless.

- Remove and retain 12 µL of eluate into a clean 1.5 ml DNA LoBind tube or 0.2 ml thin-walled PCR tube strip.

- Quantify 2 μl of each eluted sample using a Qubit flex using the dsDNA 1x HS assay

- Take forward an equimolar mass of samples to be barcoded and pooled forward into the native barcode ligation step.

- At this point it is also possible to store the sample at 4°C overnight.

#### Native Barcode Ligation

#### Step specific reagents

- NEB Blunt/TA Ligase Master Mix (M0367).

- Native Barcoding Expansion kits: Select a unique barcode for every sample to be run together on the same flow cell, from the provided 24 barcodes.

- EDTA: used to stop the ligation reaction.

- AMPure XP Beads (AXP).

    - Start protocol:

- Thaw the EDTA at RT, mix by vortexing, spin down and place on ice or  on a cooling block.

- Thaw the Native Barcodes (NB01-24) required for your number of samples at RT. Individually mix the barcodes by pipetting, spin down, and place them on ice/cooling block

- In clean 0.2 ml PCR-tubes or a 96-well plate, add the reagents in the following order per well:

| Reagent | Volume |
| --- | --- |
| End-prepped DNA | 7,5 μL |
| Native Barcode (NB01-24) | 2,5 μL |
| Blunt/TA Ligase Master Mix | 10μL |

- Ensure the reaction is thoroughly mixed by gently pipetting and spin down briefly.

- Incubate for 20 minutes at RT.

- Add 2 μl of EDTA to each well and mix thoroughly by pipetting and spin down briefly.

  - EDTA is added at this step to stop the reaction.

- Pool all the native barcoded sample pools in a 1.5 ml Eppendorf DNA LoBind tube.

| Number of dual barcoded pools | 1 Pool | 6 Pools | 12 Pools | 24 Pools |
| --- | --- | --- | --- | --- |
| Volume | 22 µL | 132 µL | 264 µL | 528 µL |

Purify DNA with a AMPure XP 0.4x bead-clean-up

- Resuspend the AMPure XP Beads (AXP) by vortexing.

- Add AMPure XP Beads (AXP) to the pooled Library, and mix by pipetting.

  - Determine the correct volume of AXP beads for a 0.4X clean-up.

- Allow DNA to bind to beads for 10 minutes at room temperature.

- Spin down the sample and pellet on a magnet for 5 minutes. Keep the plate on the magnetic rack until the eluate is clear and colourless, and pipette off the supernatant.

- Keep the tube on the magnetic rack and wash the beads with 700 μl of freshly prepared 80% ethanol without disturbing the pellet. Remove the ethanol using a pipette and discard.

- Repeat the previous step.

- Spin down and place the tube back on the magnet. Pipette off any residual ethanol. Allow to dry for ~ 30 seconds, but do not dry the pellet to the point of cracking.

- Remove the tube from the magnetic rack and resuspend the pellet in 35 μl Nuclease-free water by gently flicking or by pipetting up and down.

- Incubate for 10 minutes at RT. Every 2 minutes, agitate the sample by gently flicking for 10 seconds to encourage DNA elution.

- Pellet the beads on a magnet until the eluate is clear and colourless.

- Remove and retain 35 µL of eluate into a clean 1.5 ml DNA Lobind tube.

- Quantify 2 μl of each eluted sample using a Qubit flex using the dsDNA 1x HS assay

  - Shake the bottle of buffer

  - Optional quality control: use 1 µL of end-prepped DNA on the Tape station (D5000 kit).

- Take forward the barcoded DNA library to the adapter ligation and clean-up step. However, at this point it is also possible to store the sample at 4°C overnight.

  - Native Adapter ligation

    - Step specific reagents

- Long Fragment Buffer (LFB) → for DNA fragments of 3kb or longer

- Short Fragment Buffer (SFB) → for DNA fragments of all sizes

- Elution Buffer from the Oxford Nanopore kit (EB)

- Native Adapter (NA)

- NEBNext Quick T4 DNA Ligase

- NEBNext Quick Ligation Reaction Buffer (5X)

Important notes:

- The Native Adapter (NA) used in this kit and protocol is not interchangeable with other sequencing adapters.

- Depending on the wash buffer (LFB or SFB) used, the clean-up step after adapter ligation is designed to either enrich DNA fragments of >3 kb, or purify all fragments equally.

- From this point on the clean-up steps uses Long Fragment Buffer (LFB) or Short Fragment Buffer (SFB) rather than 80% ethanol to wash the beads. The use of ethanol will be detrimental to the sequencing reaction.

    - Method

- Thaw the reagents at RT.

- Spin down the reagent tubes for 5 seconds.

- Ensure the reagents are fully mixed by performing 10 full volume pipette mixes.

- Spin down the Native Adapter (NA) and Quick T4 DNA Ligase, pipette mix and place on ice.

- Thaw the Elution Buffer (EB) at RT, mix by vortexing and spin down.

- Thaw either Long Fragment Buffer (LFB) or Short Fragment Buffer (SFB) at RT, mix by vortexing, spin down and place on ice.

- In a 1.5 ml Eppendorf DNA LoBind tube, mix in the following order:

| Reagent | Volume |
| --- | --- |
| Pooled barcoded sample/library | 30 μl |
| Native Adapter (NA) | 5 μl |
| NEBNext Quick Ligation Reaction Buffer (5X) | 10 μl |
| NEBNext Quick T4 DNA Ligase | 5 μl |

- Ensure the reaction is thoroughly mixed by gently pipetting and spin down briefly.

- Incubate the reaction for 20 minutes at RT.

- Resuspend the AMPure XP Beads (AXP) by vortexing.

- Add 20 μl of resuspended AMPure XP Beads (AXP) to the reaction and mix by pipetting up and down.

- Incubate for 10 minutes at RT

- Spin down the sample and pellet on a magnet. Keep the tube on the magnet, and pipette off the supernatant.

- Wash the beads by adding either 125μl Long Fragment Buffer (LFB) or 125μl Short Fragment Buffer (SFB). Flick the beads to resuspend, spin down, then return the tube to the magnetic rack and allow the beads to pellet. Remove the supernatant using a pipette and discard.

- Repeat the previous step.

- Spin down and place the tube back on the magnet. Pipette off any residual supernatant.

- Remove the tube from the magnetic rack and resuspend the pellet in 15 µL Elution Buffer (EB). Incubate for 10 minutes at RT.

  - Every 2 minutes, agitate the sample by gently flicking for 10 seconds to encourage DNA elution.

- Pellet the beads on a magnet until the eluate is clear and colourless, for at least 1 minute.

- Remove and retain 15 μl of eluate containing the DNA library into a clean 1.5 ml Eppendorf DNA LoBind tube.

- The tapestation D5000 kit can be used for a final assessment on the library quality.

- Use the Qubit dsDNA 1x HS assay to quantify the DNA library.

- Use 2 µL per DNA pool for the dsDNA 1x HS assay on the Qubit Flex.

- Ensure that the library is set up to 12 μl at 10-20 fmol.

- The prepared library is used for loading into the flow cell. Store the library on ice until ready to load.

#### Priming and loading the SpotON flow cell

Important notes

- Please note, this kit is only compatible with R10.4.1 flow cells (FLO-MIN114).

- For optimal sequencing performance and improved output on MinION R10.4.1 flow cells (FLO-MIN114), add Bovine Serum Albumin (BSA) to the flow cell priming mix at a final concentration of 0.2 mg/ml. The use of recombinant BSA is not recommended.

- Perform a flow cell check before starting the loading of the spotON flow cell, see step 3 on how to insert the flow cell.

- Avoid the introduction of air bubbles, this will damage the nanopores.

- Thaw the Sequencing Buffer (SB), Library Beads (LIB), Flow Cell Tether (FCT) and one tube of Flow Cell Flush (FCF) at RT. Mix by vortexing and spin down.

- To prepare the flow cell priming mix with BSA, add the following reagents directly to the tube of Flow Cell Flush (FCF), and mix by inverting the tube and pipette mix at RT:

- One tube of Flow Cell Flush (FCF) (1170 μl)

- 25 μl Bovine Serum Albumin (BSA) at 10 mg/mL

- 30 μl Flow Cell Tether (FCT)

| - Open the MinION or GridION device lid and slide the flow cell under the clip. Press down firmly on the flow cell to ensure correct thermal and electrical contact. - Slide the flow cell priming port cover clockwise to open the priming port (red arrow) to the position in the next step. |  |
| --- | --- |
| - After opening the priming port, check for a small air bubble under the cover. Draw back a small volume to remove any bubbles (a few μl): - Set a P1000 pipette to 200 μl - Insert the tip into the priming port - Turn the wheel until the dial shows 220-230 μl, or until you can see a small volume of buffer entering the pipette tip Note: Visually check that there is a continuous buffer from the priming port across the sensor array. Important: Do not remove more than 20-30 μl! |  |
| - Load 800 μl of the priming mix into the flow cell via the priming port, avoiding the introduction of air bubbles. Wait for 5 minutes. During this time, prepare the library for loading by following the steps below. - Thoroughly mix the contents of the Library Beads (LIB) by pipetting. - The Library Beads (LIB) tube contains a suspension of beads. These beads settle very quickly. It is vital that they are mixed immediately before use. - In a new tube, prepare the library for loading as follows: - 37.5 μl Sequencing Buffer (SB) - 25.5 μl Library Beads (LIB) mixed immediately before use - 12 μl of the prepared DNA library (10 -20 fmol) |  |
| - Complete the flow cell priming: - Gently lift the SpotON sample port cover to make the SpotON sample port accessible (red arrow). - Carefully load 200 μl of the remaining priming mix into the flow cell priming port (not the SpotON sample port), avoiding the introduction of air bubbles. |  |
| - Mix the prepared library gently by pipetting up and down just prior to loading. - Add 75 μl of sample to the flow cell via the SpotON sample port in a dropwise fashion. Ensure each drop flows into the port before adding the next. - Gently replace the SpotON sample port cover, making sure the bung enters the SpotON port, close the priming port and replace the MinION or GridION device lid. |  |
| Overview of a correctly closed flow cell |  |

  - Starting the run

- In the MinKNOW UI go to the “Start” screen.

- Select “Start Sequencing”.

- This will take you to the run setup screen.

- Positions

  - Create an experiment name and sample ID, flow cell ID and flow cell type will be recognized automatically.

  - ARISE experiment name: ARISE_<YYYY>

  - ARISE Sample ID: “<YYYY>-<mm>-<DD>_BATCH-<batch number>_<amount of plates>P”

- Kit selection

  - The preferred kit can be selected here.

- Run options

  - Here changes can be made to the “run duration” and “minimum read length”.

  - “Adaptive sampling” and “Advanced options” should only be used/changed if necessary.

- Analysis

  - “Basecalling” should be turned “ON”.

  - Under “Basecalling” change the options from “High accuracy basecalling” to “Super High accuracy basecalling”.

  - "Barcoding" should automatically be enabled if barcodes are used, and the correct Kit has been selected.

- Output

  - No changes need to be made to this section.

- Review & start

  - Shows an overview of the used settings.

  - Once all settings are correct and the flow cell is loaded, the user can begin the run.

Ending the experiment

- Runs automatically stop at the end of the duration set in "Run duration" or can be stopped manually if enough reads have been collected or there are no more active pores in the flow cell.

- After the sequencing experiment is complete, if you would like to reuse the flow cell, please follow the Wash Kit instructions (EXP-WSH004) and store the washed flow cell at 2-8°C, OR go to step 3.

  - Checklist for the Flow Cell Wash

  - Full protocol for the Flow Cell Wash

  - After washing the flow cell with the wash kit, use 1mL of miliQ to flush the waste channel and tape shut waste port 1 and waste port 2 to prevent air drying and crystallisation.

- Follow the returns procedure to flush out the flow cell ready to send back to Oxford Nanopore.

  - The procedure is described on the “return box” present in the lab near the GridION.

## RELATED/ASSOCIATED DOCUMENTS

Mention important documents referred to in this procedure.

Mention all the appendices (work instructions, forms) referred to in this procedure.

SHEET010_OVERVIEW_ROOM_MANAGERS_&_TELEPHONE NUMBERS

SHEET025_WORKFLOW_OVERVIEW_SOPs_SD

### Ratios beads vs DNA sample, DNA fragment recovery

Plate layout Barcodes - 2nd PCR

|  | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| A | BC01 | BC09 | BC17 | BC25 | BC33 | BC41 | BC49 | BC57 | BC65 | BC73 | BC81 | BC89 |
| B | BC02 | BC10 | BC18 | BC26 | BC34 | BC42 | BC50 | BC58 | BC66 | BC74 | BC82 | BC90 |
| C | BC03 | BC11 | BC19 | BC27 | BC35 | BC43 | BC51 | BC59 | BC67 | BC75 | BC83 | BC91 |
| D | BC04 | BC12 | BC20 | BC28 | BC36 | BC44 | BC52 | BC60 | BC68 | BC76 | BC84 | BC92 |
| E | BC05 | BC13 | BC21 | BC29 | BC37 | BC45 | BC53 | BC61 | BC69 | BC77 | BC85 | BC93 |
| F | BC06 | BC14 | BC22 | BC30 | BC38 | BC46 | BC54 | BC62 | BC70 | BC78 | BC86 | BC94 |
| G | BC07 | BC15 | BC23 | BC31 | BC39 | BC47 | BC55 | BC63 | BC71 | BC79 | BC87 | BC95 |
| H | BC08 | BC16 | BC24 | BC32 | BC40 | BC48 | BC56 | BC64 | BC72 | BC80 | BC88 | BC96 |

Fw + rv

## LITERATURE

https://protocols.opentrons.com/protocol/normalization

C.wash manuals

- C.wash user manual

- C.wash software manual

- SOP

Nanopore

- Checklist dual barcoding kit v14

- Full protocol

- Technical information