## Lot 3 - 12+ ch datalogger
Response form Lot 3 - Program of Requirements and awarding criteria
Awarding criteria - Lot 3: 12+ ch datalogger for geophone string
Type | # | Datalogger quality consisting of: | Max. points | Answer tenderer | Score
KO | D1 | The provided datalogger will have simultaneous sampling on all channels | Yes/No |  | Not applicable
KO | D2 | The provided datalogger will have a resolution of at least 24 bit @ 200 sps | Yes/No |  | Not applicable
KO | D3 | The provided datalogger will have at least 12 channel input | Yes/No |  | Not applicable
KO | D4 | The provided datalogger is capable of continuous recording with MiniSEED support | Yes/No |  | Not applicable
KO | D5 | The provided datalogger will be able to continuously stream data, seismic and state-of-health, supporting the SeedLink protocol | Yes/No |  | Not applicable
KO | D6 | The provided datalogger provides automatic data backfill for when communication is interrupted | Yes/No |  | Not applicable
KO | D7 | The provided datalogger will haev a timing accuracy < 5 µs via internal GNSS receiver and external antenna | Yes/No |  | Not applicable | more than 3 extra channel
KO | D8 | The provided dataloffer wil have maximum dimensions of 25 x 50 x 40 cm (lxwxh) | Yes/No |  | Not applicable | 2 or 3 extra channel
KO | D9 | The provided datalogger will have an Internal storage of at least 25 GB | Yes/No |  | Not applicable | 1 extra channel
KO | D10 | The provided datalogger support file (incidental) file transfer through internet | Yes/No |  | Not applicable | no extra channel
KO | D11 | The provided datalogger has remote self diagnostics, configuration, tests and calibration capabilities through internet, in combination with the offered sensor | Yes/No |  | Not applicable
KO | D12 | The provided datalogger has an operating temperature range of -20°C to 60°C | Yes/No |  | Not applicable
KO | D13 | The provided datalogger has an IP67 equivalent or higher rating | Yes/No |  | Not applicable
KO | D14 | Full datasheet with specifications of the datalogger is submitted | Yes/No |  | Not applicable
KO | D15 | Including connection on the datalogger for (bare-wire) geophone channels, to be connected in field conditions (i.e. break-out box/terminal board) | Yes/No |  | Not applicable
D | D16 | Extra channels with support for some or all of the following (see D3): 3C surface seismometer, Geophone sensors, Strong motion accelerometers, Microbarometers, Meteorological Sensors more than 3 extra channel - 50 points 2 or 3 extra channel - 35 1 extra channel - 15 points no extra channel - 0 points | 50 |  | 0
|  | Total score D: (if all knockout criteria are met) | 50 | 0 | 0
Type | # | Other criteria: | Max. points | Answer tenderer | Score
KO | O1 | Delivery time up to 4 months faster is better (see #O5), longer is a knockout | Yes/No |  | Not applicable
KO | O2 | Warranty Free of charge replacment of faulty components for (at least) 12 months a component is deemd faulty if it causes the station to deviate from the knockout criteria | Yes/No |  | Not applicable | ≤60 days
DO | O3 | Environmental Awareness and sustainability The Contracting Authority attaches great value to sustainable business practices and considers this to be a joint responsibility. Tenderer is requested to substantiate the ways it's business, the manufacturing of the required product and usage of the product including service contributes to sustainability. The more it can be demonstrated that sustainability is also pursued and realized by Tenderer in this procurement, the more points can be achieved for this criterion. (Please also see Descriptive Document 6.3.1) Please add you substantiation in max. 1 A4 in a MS Word- or PDF - file to your Tender (clearly marked with "D O2 Environmental awareness and sustainability" in the file name). | Is the substantiation provided with the Tender? Yes/No (Max points: 100) |  | Max score is 100 The score for this criterion will be assessed by the assessment team on the basis of the substantiation provided | >60 days ≤90 days
DO | O4 | Service and added value Tenderer is requested to substantiate it's Service and Added value for Contracting Authority. The more added value Tenderer can provide in terms of service, support, troubleshooting or extra feautures in the product, the better the score will be. (Please also see Descriptive Document 6.3.1) Please add you substantiation in max. 2 A4's in a MS Word- or PDF - file to your Tender (clearly marked with "D O4 Service and added value" in the file name). | Is the substantiation provided with the Tender? Yes/No (Max Points: 150) |  | Max score is 150 The score for this criterion will be assessed by the assessment team on the basis of the substantiation provided | >90 ≤ 120 days | more than 36 months
D | O5 | Faster delivery upto max. 60 days - 50 points up to max. 90 days - 25 points up to max. 120 days - 0 points more than 120 days - 0 points (Also a K.O. Please also see requirement # O1) Please note; the given term will be included as a contractual agreement. | 50 |  | 0 | >120 days | more than 12 up to 36 months
D | O6 | Longer warranty period more than 36 months - 50 points more than 12 up to 36 months - 25 points upto 12 months - 0 points Please note; the given term will be included as a contractual agreement. | 50 |  | 0 |  | up to 12 months
|  | Total score O: (if all knockout criteria are met) | 350 | 0 | 0
|  | Total score (D + O) | 400 |  | 0

## Blad1 (hidden)
Yes | ≤
No | ≥
S12
0-1.5 | 10 points
1.5-3 | 5 points
>3 | 0 points
S13 (Hz)
<0.01->50 | 20 points
>0.01-<50 | 10 points
0.01 Hz-40 H | 0 points
S14
>10 Hz | 25 points
>0.08<10 | 12 points
0.1-5 Hz | 0 points
S15
>150 | 10 points
140≤150 | 5 points
135 dB to 140 dB | 0 points
D14
Resolution > 24 bit @ 100 sps
≤24 | 0 points
>24 | 10 points
D15
>3 | 30 points
2-3 | 10 points
1 | 5 points
0 | 0 points
D16
Yes | 5 points
No | 0 points
D17
Yes | 5 points
No | 0 points
D18
Yes | 10 points
No | 0 points
D19
1.5-3 | 5 points
>3 | 0 points
W05
0≤30 | 20 points
31≤60 | 10 points
61≤90 | 0 points
W06
≥36 | 10 points
>12<36 | 5 points
≤12 | 0 points