## Instructions and Explanations
Instructions regarding the Programme of Requirements and Desirables - TN-349307
Introduction
This Program of Requirements and Desirables forms an integral part of TN-349307. It details the technical Requirements and Desirables of the new present weather sensors that the KNMI intends to buy. Every Requirement and Desirable is of a specific type, where the possible types are listed below in the Requirement and Desirable Types section. Every type has a specific answer format that must be adhered to. Some types are associated with a substantiation that will be given a score by the assessment team. Every Desirable can receive a score between 0 and 5. This score will be multiplied by the weight of the particular Desirable to yield the points awarded to the Tenderer for that specific Desirable. Note that the weights are not integer numbers. A scaling factor is applied to the score of the Desirables to KNMI in order to ensure the maximum achievable points for the Quality equals 750. For more information on the award criteria, see the Descriptive Document. In general, the assessment is done using the guidelines described in the Descriptive Document. Additional focus points are mentioned in the Evaluation Criteria next to each Desirable, indicating the specific aspects that are of interest to the KNMI assessment team. Read these carefully, because for the D KO type of Desirables these evaluation criteria can contain guidelines on what is considered necessary to achieve the minimum score.
Fill-in instructions IMPORTANT INFORMATION
This document also functions as the formal declaration of the Tenderer to comply with the Requirements and to indicate in what extent the Tenderer complies with the Desirables. For this purpose, fill in the Requirements and Desirables sheet of this excel file for all yellow cells. These cells typically contain a drop-down menu that can be used to select the proper option. Tenderer may not change any other cells within this excel file even if technically possible. For the purpose of answering the open Desirables with a substantiation, Annex 12 is provided with a numbering of the Desirables that require substantiations. Tenderer may use this annex to structure the substantiations. Tenderer may also use an own template or file to structure the substantiations. In all cases, the text in this Program of Requirements is leading for all Requirements. Note that specifically Mandatory #36 and Desirable #117 call for attachment of separate files to be fulfilled. These must be included with the tender and must be clearly marked in the file name to contain the required files. If multiple files are required, these may be bundled in a single zip file per Mandatory or Desirable, again clearly marked in the file name. The amount of space available for the substantiations is given in terms of A4 pages. This is done deliberately to allow for the inclusion of figures & graphs which tenderer may use to support the substantiation. The text in these substantiations must be in "Calibri" font with font size 11pt or larger. Any text within graphics must be of similar (or larger) size and readability. Any substantiation text or graphics which are outside the allowed size will be ignored during the assessment.
Requirement and Desirable Types
Every Requirement and Desirable in this Program of Requirements and Desirables is of a specific type. These types are listed here and includes an explanation of what is expected for each type.
Type | Sub-type | Abbreviation | Desired answer | Explanation | Scoring
Mandatory | Closed | M C | Yes or No | A singular Yes or No answer is requested, without any extra information. | n/a
Mandatory | Information | M I | Yes or No with extra information | A singular Yes or No answer is requested that is supplemented with particular objective information (such as a number or list). | n/a, the supplementary information is not used in the scoring.
Desirable | Closed | D C | Yes or No | A singular Yes or No answer is requested, without any extra information. | Score 0 for No. Score 5 for Yes.
Desirable | Scale | D S | Information | A specific piece of objective information is requested (such as a number or list). | Scoring according to the criteria given.
Desirable | Open | D O | Substantiation | A substantiation for the Desirable is requested, in accordance with the guidelines given. | Scoring by assessment team based on Assessment table (see below). Score of 0 to 5 is given.
Desirable | Open with Knock-Out | D KO | Substantiation with minimum score | A substantiation for the Desirable is requested, in accordance with the guidelines given. A knock-out criteria applies to this substantiation if the Tenderer does not achieve the minimum score stated. | Scoring by assessment team based on Assessment table (see below). Score of 0 to 5 is given.
Assessment levels
Level | Description |  |  |  | Score
Excellent | The substantiation is very complete and highly consistent, is clearly relevant, is very convincing and/or the solution offers extensive added value. |  |  |  | 5
Good | The substantiation is mostly complete and consistent, is mostly relevant, is convincing and/or the solution offers clear added value. |  |  |  | 4
Sufficient | The substantiation is sufficient/consistent, is partly relevant, is reasonably convincing and/or the solution is sufficient and offers some added value. |  |  |  | 3
Poor | The substantiation is only barely sufficient/consistent, is only somewhat relevant, is barely convincing and/or the solution is barely sufficient and offers no added value. |  |  |  | 2
Insufficient | The substantiation is insufficient or not consistent, is not relevant, is not convincing and/or the solution is insufficient. |  |  |  | 1
Terrible/No information | The substantiation is either completely missing or too fragmentary to meaningfully judge for its value. |  |  |  | 0
Definitions
Short | Definition
WMO requirements | WMO-No. 8, Guide to Instruments and Methods of Observation, 2018 Edition
ICAO requirements | ICAO Annex 3, Meteorological Service for International Air Navigation, Twentieth Edition, July 2018
EASA airport requirements | Certification Specifications and Guidance Material for Aerodromes Design Issue 6 - CHAPTER Q — VISUAL AIDS FOR DENOTING OBSTACLES
1st line maintenance | See descriptive document section 3.1. This maintenance is performed by KNMI or 3rd party personnel in the field.
2nd line maintenance | See descriptive document section 3.1. This maintenance is performed by KNMI personnel in De Bilt.
3rd line maintenance | See descriptive document section 3.1. This maintenance is performed by Tenderer.

## Requirements and Desirables
KNMI - PWS Tender 2023 - Requirements and Desirables | [6 empty] | Notice: Not all mandatories are completed.
# | M/D | Description | Score Type | Weight | Evaluation Criteria | Answer | Outcome | Points
1 Functional Requirements
1.1 General Requirements
1. | <M> | The PWS shall be capable of measuring and reporting the: • Visibility (MOR: Meteorological Optical Range); • Background luminance; • Precipitation type; • Precipitation intensity. | M C | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
2. | <D> | Tenderer should substantiate why the offered PWS is considered to be suitable for the intended purpose of KNMI (see section 3.1 through to 3.4 of the Descriptive Document) and the advantages of the offered PWS. | D O | 8.0214 | Please provide a substantiation of maximum 1 A4 page. The substantiation will be assessed and a score is given between 0 to 5 maximum. The assessment is based on expert judgement with special attention given to the following assessment criteria: - Suitability and advantages of the PWS w.r.t. the intended purpose as described in the Descriptive Document. - Clarity, completeness and level of detail of the substantiation. - Higher score for better performing PWS in general. | Select applicable option: | The score will be assigned by the KNMI assessment team. | #VALUE!
1.2 Visibility (MOR) Requirements
3. | <M> | The PWS shall meet the requirements for visibility (MOR, aeronautical visibility and Runway Visual Range) according to the WMO and ICAO requirements, see Definitions in the "Instructions & Explanations" sheet. Please note: One exception is that the WMO requirement of MOR up to 100 kilometers is not required for this Tender. Instead, the MOR range shall be at least 10 meters to 50 kilometers. Also, note that since the PWS is used for synoptic, climatological and aeronautical applications it must meet the requirements for RVR, aeronautical visibility, and synoptical MOR. Therefore, the strictest requirement is applicable, more specifically, RVR at low end and synoptical MOR at high end. | M C | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
4. | <M> | The MOR reported by any two sensors of the offered PWS placed next to each other shall agree within the WMO and ICAO measurement uncertainty over the entire measurement range. | M C | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
5. | <D> | Tenderer should substantiate that Requirements 3 and 4 are met. The substantiation should include the measurement uncertainty according to the "Guide to the Expression of Uncertainty in Measurement" (GUM, JCGM_100_2008)". This should include the effects of e.g. long and short term stability, linearity, hysteresis, reproducibility, temperature dependency, calibration and lens contamination on the measurement uncertainty. It should also include the quality and measurement uncertainty of the MOR reported over the entire measurement range of the PWS. | D KO | 14.4385 | Please provide a substantiation of maximum 2 A4 pages. The substantiation will be assessed and a score is given between 0 to 5 maximum. A minimum score of 3 must be achieved; otherwise, the Tender is disqualified (knock-out). The assessment is based on expert judgement with special attention given to the following assessment criteria: - Clarity, completeness and level of detail of the substantiation. - The results of a field test with at least two identical sensors next to each other must be included; otherwise, the Tender is disqualified. This field test must span all seasons and the entire measurement range of the MOR. - Inclusion of the results of a field test with a PWS versus a transmissometer and spanning all seasons is desired. - Detail with which the measurement uncertainty is stated in accordance with the GUM. - Absolute accuracy/quality of the MOR measurement. | Select applicable option: | The score will be assigned by the KNMI assessment team. | #VALUE!
6. | <D> | The range of MOR should be at least 10 meters to 100 kilometers. | D S | 1.6043 | 50 kilometers range = score 0 51 to 63 kilometers = score 1 63 to 75 kilometers = score 2 75 to 88 kilometers = score 3 88 to 99 kilometers = score 4 ≥100 kilometers = score 5 No additional score for minimum range <10 meters. | Please provide the actual range of the PWS: | Score | #VALUE!
1.3 Background Luminance Requirements
7. | <M> | The PWS shall meet the WMO and ICAO requirements for background luminance. | M C | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
8. | <M> | The background luminance reported by any two sensors of the offered PWS placed next to each other shall agree within the WMO and ICAO measurement uncertainty over the entire measurement range. | M C | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
9. | <D> | Tenderer should substantiate that Requirements 7 and 8 are met. The substantiation should include a specification of the measurement uncertainty according to the "Guide to the Expression of Uncertainty in Measurement" (GUM, JCGM_100_2008). This substantiation should include the effects of e.g. long and short term stability, linearity, hysteresis, reproducibility, temperature dependency, calibration and lens contamination on the measurement uncertainty. It should also include the quality and measurement uncertainty of the background luminance reported over the entire range of the PWS. | D KO | 4.8128 | Please provide a substantiation of maximum 2 A4 pages. The substantiation will be assessed and a score is given between 0 to 5 maximum. A minimum score of 3 must be achieved; otherwise, the Tender is disqualified (knock-out). The assessment is based on expert judgement with special attention given to the following assessment criteria: - Clarity, completeness and level of detail of the substantiation. - Inclusion of the results of a field test with at least two sensors next to each other and spanning all seasons is desired. - Detail with which the measurement uncertainty is stated in accordance with the GUM. - Accuracy/quality of the background luminance measurement. | Select applicable option: | The score will be assigned by the KNMI assessment team. | #VALUE!
1.4 Precipitation Type Requirements
10. | <M> | Each precipitation type that the PWS can report has been verified during field evaluations. The Probability of Detection (POD) of liquid precipitation for conditions in the Netherlands shall exceed 80% and the False Alarm Rate (FAR) shall be less than 25%. For solid precipitation, the POD shall exceed 65% and the FAR shall be less than 25%. | M C | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
11. | <D> | Tenderer should substantiate that Requirement 10 is met. References to field evaluations, for example against certified human observers using WMO and ICAO reporting practices or via an alternative method (details of this method should be provided), shall be provided and the level of agreement shall be stated for each type expressed in Probability of Detection (POD) and False Alarm Rate (FAR). These field evaluations should cover all seasons and as many precipitation types as possible, and should be representative for the climate conditions in the Netherlands. | D KO | 14.4385 | Please provide a substantiation of maximum 2 A4 pages. The substantiation will be assessed and a score is given between 0 to 5 maximum. A minimum score of 3 must be achieved; otherwise, the Tender is disqualified (knock-out). The assessment is based on expert judgement with special attention given to the following assessment criteria: - The FAR and POD values. - The precipitation types present during the evaluation period. - The results of a field test covering all seasons must be included; otherwise, the Tender is disqualified. The field test is graded on its representativeness. | Select applicable option: | The score will be assigned by the KNMI assessment team. | #VALUE!
12. | <M> | The precipitation types that the PWS can report shall include at least drizzle, rain, snow, ice pellets (sleet) and hail. | M C | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
13. | <D> | Tenderer should specify the precipitation types that the PWS can report and should specify and substantiate the quality of the precipitation types reported by the PWS. | D O | 6.4171 | Please provide a substantiation of maximum 1 A4 page. The substantiation will be assessed and a score is given between 0 to 5 maximum. The assessment is based on expert judgement with special attention given to the following assessment criteria: - Clarity, completeness and level of detail of the substantiation. - Quality of the detection of precipitation. - Accuracy with which the precipitation types are reported, preferably in substantiated FAR and POD numbers. Note that no score is awarded for the number of precipitation types that can be reported. | Select applicable option: | The score will be assigned by the KNMI assessment team. | #VALUE!
14. | <D> | Tenderer should specify and substantiate the agreement between the reported precipitation type of any two sensors of the offered PWS placed next to each other. | D O | 4.8128 | Please provide a substantiation of maximum 1 A4 page. The substantiation will be assessed and a score is given between 0 to 5 maximum. The assessment is based on expert judgement with special attention given to the following assessment criteria: - Clarity, completeness and level of detail of the substantiation. - Level of agreement between any two sensors. | Select applicable option: | The score will be assigned by the KNMI assessment team. | #VALUE!
1.5 Precipitation Intensity Requirements
15. | <M> | The PWS shall meet the WMO and ICAO requirements for precipitation detection. Detection shall either be based on the precipitation intensity threshold for detection of WMO or on certified observers and a Probability of Detection (POD) exceeding 80% and a False Alarm Rate (FAR) below 25%. | M C | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
16. | <D> | Tenderer should substantiate that Requirement 15 is met. References to field evaluations must be included. These field evaluations should be validated by certified human observers using WMO and ICAO reporting practices and the level of agreement should be stated and expressed in Probability of Detection (POD) and False Alarm Rate (FAR). Alternatively, a field evaluation of the precipitation detection against the precipitation intensity threshold should be substantiated. The field evaluations should cover at least all seasons and be representative for the climate conditions in the Netherlands. | D KO | 8.0214 | Please provide a substantiation of maximum 2 A4 pages. The substantiation will be assessed and a score is given between 0 to 5 maximum. A minimum score of 3 must be achieved; otherwise, the Tender is disqualified (knock-out). The assessment is based on expert judgement with special attention given to the following assessment criteria: - Clarity, completeness and level of detail of the substantiation. - Quality of the precipitation detection (FAR and POD numbers). - The results of a field test covering all seasons must be included; otherwise, the Tender is disqualified. The field test is graded on its representativeness. | Select applicable option: | The score will be assigned by the KNMI assessment team. | #VALUE!
17. | <D> | Tenderer should specify and substantiate the quality and measurement uncertainty of the precipitation intensity reported by the offered PWS for the precipitation intensity range 0.02 to 500 mm/h. The substantiation should include a specification of the measurement uncertainty according to the "Guide to the Expression of Uncertainty in Measurement" (GUM, JCGM_100_2008). This substantiation should include the effects of e.g. long and short term stability, linearity, hysteresis, reproducibility, temperature dependency, calibration and lens contamination on the measurement uncertainty. It should also substantiate that the precipitation intensity reported by any two sensors of the offered PWS placed next to each other agree within the WMO measurement uncertainty over the entire measurement range. | D O | 6.4171 | Please provide a substantiation of maximum 1 A4 page. The substantiation will be assessed and a score is given between 0 to 5 maximum. The assessment is based on expert judgement with special attention given to the following assessment criteria: - Clarity, completeness and level of detail of the substantiation. - Detail with which the measurement uncertainty is stated in accordance with the GUM. - The quality of the precipitation intensity. - The results of testing of two sensors next to each other. | Select applicable option: | The score will be assigned by the KNMI assessment team. | #VALUE!
1.6 Temperature and Humidity Requirements
18. | <M> | Either (case 1) the PWS shall not contain an integrated temperature and/or humidity probe that is used for the discrimination of the precipitation type, or (case 2) it shall be possible to verify and calibrate the integrated temperature and/or humidity probe of the PWS. Please note: KNMI envisions the possibility of using temperature and/or humidity data for better discrimination of freezing and solid precipitiation. | M I | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
19. | <M> | Either (case 1) the PWS shall not contain an integrated temperature and/or humidity probe that is used for the discrimination of the precipitation type, or (case 2) the measurement values of the integrated temperature and/or humidity probe shall be available in the data message of the PWS. Please note: KNMI envisions the possibility of using temperature and/or humidity data for better discrimination of freezing and solid precipitiation. | M I | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
1.7 Environmental Requirements
20. | <M> | The PWS shall have a MOR filter for removing or mitigating the signal caused by insects interacting with the measurement volume. For background information regarding how the visibility reported by a forward scatter PWS can be affected by flying insects, see TECO, 2012_paper P1(28). | M C | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
21. | <D> | Tenderer should substantiate that the design of the PWS is such that the effect of flying insects on the measured and reported visibility is within the required measurement uncertainty. The effectiveness of the insect filter should be substantiated and illustrated by examples of real-life cases. The cases should include situations showing an unfiltered MOR below 1 kilometer and filtered MOR above 8 kilometer. It is allowed but not necessary to include situations obtained from external parties that are using the offered PWS. | D KO | 16.0428 | Please provide a substantiation of maximum 2 A4 pages. The substantiation will be assessed and a score is given between 0 to 5 maximum. A minimum score of 3 must be achieved; otherwise, the Tender is disqualified (knock-out). The assessment is based on expert judgement with special attention given to the following assessment criteria: - Performance and quality of the filter. - Clarity, completeness and level of detail of the substantiation. - The results of a field test showing the effect of the insect filter must be included; otherwise, the Tender is disqualified. | Select applicable option: | The score will be assigned by the KNMI assessment team. | #VALUE!
22. | <D> | Tenderer should substantiate that the design of the PWS is such that the effect of disturbances or contamination on the quality of the measurements is within the required measurement uncertainty (when applicable after compensation) or covered by the maintenance interval. Contamination is caused by, but not limited to, insects/spiders, salt water spray and salt deposits, (blowing) precipitation, dust, pollen, exhaust or other environmental factors. | D KO | 14.4385 | Please provide a substantiation of maximum 2 A4 pages. The substantiation will be assessed and a score is given between 0 to 5 maximum. A minimum score of 3 must be achieved; otherwise, the Tender is disqualified (knock-out). The assessment is based on expert judgement with special attention given to the following assessment criteria: - Clarity, completeness and level of detail of the substantiation. - Robustness against the mentioned contamination. - All three categories of contamination (insects, spiders, and other contamination) must be covered in the substantiation; otherwise, the Tender is disqualified. | Select applicable option: | The score will be assigned by the KNMI assessment team. | #VALUE!
23. | <D> | Tenderer should substantiate that the design of the PWS is such that the effect of wind on the measurements or measurement area is within the required measurement uncertainty. The substantiation should include arguments based on geometrical design and/or relevant field testing of the offered PWS. If possible, the wind speed at which the required measurement uncertainty is no longer guaranteed should be specified and substantiated by experiments. | D O | 1.6043 | Please provide a substantiation of maximum 1/2 A4 page. The substantiation will be assessed and a score is given between 0 to 5 maximum. The assessment is based on expert judgement with special attention given to the following assessment criteria: - Availability of relevant experimental results. - Maximum wind speed at which measurements are still within the required measurement uncertainty. - Clarity, completeness and level of detail of the substantiation. | Select applicable option: | The score will be assigned by the KNMI assessment team. | #VALUE!
24. | <M> | The offered PWS shall be suitable for use within the Netherlands, including North Sea platforms and the Caribbean Netherlands. | M C | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
25. | <M> | The offered PWS shall be equipped with the means to keep the windows unobstructed by condensation, snow and ice during continuous operation for climate conditions in the Netherlands including North Sea platforms and the Caribbean Netherlands. | M C | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
26. | <M> | The offered PWS shall not be affected, within the required measurement uncertainty, by (reflected) sunlight or other light sources, such as runway/traffic lights or lights from passing aircraft/cars. | M C | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
27. | <M> | The offered PWS is suitable for usage above the land and sea conditions occurring in the Netherlands including North Sea platforms and Caribbean Netherlands, within the required measurement uncertainty and maintenance interval. These conditions include, but are not limited to, surfaces of grass, dirt, sand, rock and water including the state of the grass covered ground, the state of the sea and the presence of snow on the ground. | M C | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
2 Maintainability Requirements
2.1 Visibility Requirements
28. | <M> | The Tenderer shall include a "visibility calibrator" that ensures that the MOR of the PWS can be verified, adjusted and calibrated by KNMI staff. This calibration procedure must be clearly documented. | M C | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
29. | <M> | The "visibility calibrator” shall be traceable to a standard reference. | M C | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
30. | <D> | Tenderer should substantiate the verification and calibration of the MOR reported by the PWS using the “visibility calibrator”. | D O | 8.0214 | Please provide a substantiation of maximum 1 A4 page. The substantiation will be assessed and a score is given between 0 to 5 maximum. The assessment is based on expert judgement with special attention given to the following assessment criteria: - Clarity, completeness and level of detail of the substantiation. - Correct traceability to a standard measurement reference. - Completeness of the verification and calibration procedure. - Ease of use to perform the calibration. | Select applicable option: | The score will be assigned by the KNMI assessment team. | #VALUE!
2.2 Background Luminance Requirements
31. | <M> | The Tenderer shall include a "BL calibrator" that ensures that the background luminance of the PWS can be verified, adjusted and calibrated by KNMI staff. This calibration procedure must be clearly documented. | M C | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
32. | <M> | The “BL calibrator” shall be traceable to a standard reference. | M C | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
33. | <D> | Tenderer should substantiate the verification and calibration of the background luminance reported by the PWS using the "BL calibrator". | D O | 3.2086 | Please provide a substantiation of maximum 1 A4 page. The substantiation will be assessed and a score is given between 0 to 5 maximum. The assessment is based on expert judgement with special attention given to the following assessment criteria: - Clarity, completeness and level of detail of the substantiation. - Correct traceability to a standard measurement reference. - Completeness of the verification and calibration procedure - Ease to perform the calibration. | Select applicable option: | The score will be assigned by the KNMI assessment team. | #VALUE!
2.3 Precipitation Type Requirements
34. | <D> | The Tenderer should include a "calibrator" that ensures that the physical quantities from which the precipitation type is derived by the PWS can be verified, adjusted and calibrated by KNMI staff. This calibration procedure must be clearly documented. This procedure should also ensure that the precipitation type reported by the PWS is reproducible with a “calibrator” in a documented manner. Tenderer should substantiate the verification and calibration of these physical quantities and the reproducibility of the precipitation type reported by each PWS. | D O | 6.4171 | Please provide a substantiation of maximum 1/2 A4 page. The substantiation will be assessed and a score is given between 0 to 5 maximum. The assessment is based on expert judgement with special attention given to the following assessment criteria: - Clarity, completeness and level of detail of the substantiation. - Completeness of the verification and calibration procedure - Ease to perform the calibration. | Select applicable option: | The score will be assigned by the KNMI assessment team. | #VALUE!
2.4 Precipitation Intensity Requirements
35. | <D> | The Tenderer should include a "calibrator" that ensures that the physical quantities from which the precipitation intensity is derived by the PWS can be verified, adjusted and calibrated by KNMI staff. This calibration procedure must be clearly documented. This procedure should also ensure that the precipitation intensity reported by the PWS is reproducible with a “calibrator” in a documented manner. Tenderer should substantiate the verification and calibration of these physical quantities and the reproducibility of the precipitation intensity reported by each PWS. | D O | 1.6043 | Please provide a substantiation of maximum 1/2 A4 page. The substantiation will be assessed and a score is given between 0 to 5 maximum. The assessment is based on expert judgement with special attention given to the following assessment criteria: - Clarity, completeness and level of detail of the substantiation. - Completeness of the verification and calibration procedure - Ease to perform the calibration. | Select applicable option: | The score will be assigned by the KNMI assessment team. | #VALUE!
2.5 Verification protocol
36. | <M> | Tenderer shall provide a verification protocol for the PWS and/or its modules directly with their submitted Proposal in response to this Tender. This protocol shall allow complete verification of the correct operation and the required quality of the PWS measurements in the laboratory and on the test field by KNMI before operational deployment of the PWS. | M I | n/a | Include the verification protocol in a separate file. Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
3 Technical Requirements
3.1 General Requirements
37. | <M> | A unique ID/serial number of the PWS, as well as the firmware version shall be available in the data message or is accessible through specific commands. | M C | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
38. | <M> | Commands shall be available to retrieve, change and store the parameter and configuration settings of the PWS. | M C | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
39. | <M> | The parameters/settings configured and stored in the PWS shall be retained after a re-start or shutdown due to power failure. | M C | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
40. | <M> | The time between switching on the PWS (power on) and the first output of a data message containing valid observational data shall be 5 minutes or less. | M C | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
41. | <M> | The measured quantities of the PWS shall be reported and updated at least every minute and with an output averaging time of at most 1 minute. | M C | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
42. | <D> | The measured MOR should be reported and updated every 12 seconds with an output averaging time of 12 seconds. | D C | 1.6043 | Y/N for a reporting and updating time every 12 seconds for MOR. Yes = score 5 No = score 0 | Select your answer: | No answer selected. | #VALUE!
43. | <M> | The measured quantities reported by the PWS shall contain the values generated during the most recent measurement window (according to Requirement 41 and Desirable 42). | M C | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
3.2 Communication Requirements
44. | <M> | The PWS and all associated sensors shall be equipped with one serial interface for all data communications through either (case 1) RS422 full-duplex using 4 data wires (receive and transmit over separate wires); or (case 2) RS485 half-duplex using 2 data wires (receive and transmit over the same wires). | M I | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
45. | <D> | The PWS should be equipped with a second interface for data and service communication that uses and conforms to a common protocol (e.g. serial, network, USB). The interface and functionality available through the interface should be specified. | D O | 2.4064 | Y/N for presence of a specified second interface. Yes = score 5 No = score 0 | Select your answer: | The score will be assigned by the KNMI assessment team. | #VALUE!
In this cell, write the specification of the interface and communication protocol used.
46. | <M> | Handshake shall not be required for communication with the PWS, neither in hardware nor in software. | M C | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
47. | <M> | The PWS shall support at least one of the listed options for each of the following transmission characteristics: • Speed: 1200, 2400, 4800, 9600 or 19200 bits per second; • Parity: Even, Odd, or None; • Wordsize: 7 or 8 bits; • Stopbits: 1 or 2. | M C | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
48. | <M> | The PWS shall support operation in polling mode. | M C | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
49. | <M> | The PWS shall prevent blocking of further communication in case of transmission errors by ignoring any command which is not received completely within 10.000 milliseconds. | M C | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
50. | <M> | The response of the PWS to any command at any time shall start within 200 milliseconds. The response to any valid or invalid command shall be completely transmitted within 1500 milliseconds using the maximum requested and available transmission rate. | M C | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
3.3 Sensor Output Requirements
51. | <M> | The format of the PWS response to any command shall be specified. This specification shall include the length of any response. | M C | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
52. | <M> | The PWS response shall contain a data integrity check, for example a checksum. | M C | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
53. | <M> | The PWS response shall be enclosed by unique predefined delimiters, for example STX (start of text) and ETX (end of text). | M C | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
54. | <M> | The PWS response shall be fixed, i.e. all data is located at fixed positions in the data message. For example, the information for a specific quantity is always given between byte i and byte j or as the n-th element using a unique delimiter to divide the data message into elements. | M C | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
55. | <M> | The PWS response shall be in readable ASCII. | M C | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
56. | <D> | It should be is possible within the PWS to create a user-defined output data message. | D C | 0.80214 | Yes = score 5 No = score 0 | Select your answer: | No answer selected. | #VALUE!
3.4 Status Requirements
57. | <M> | The PWS shall report status information on the operational state of the PWS and clearly identify any problem, either within the regular data message or through documented diagnostic commands. | M C | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
58. | <M> | The PWS shall perform internal checks to ensure correct functioning of each module. Status information shall be reported for each module. | M C | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
59. | <M> | The PWS shall report information that indicates if instrument parameters or window contamination levels exceed thresholds that affect the quality of the measurements or the correct operation of the PWS. These indicators shall have at least two severity classes. For example: warnings and errors. These indicators shall be included in the data message which shall be clearly described in the documentation. Suitable warning and error levels are to be provided by Tenderer. Please note: During warnings the PWS shall still meet the measurement uncertainty requirements. The PWS shall issue an error when the reliability of measurement data no longer meets the required measurement uncertainty. | M C | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
4 Hardware Requirements
4.1 General Requirements
60. | <M> | The weight and dimensions of the PWS shall be such that it can be easily transported and installed by a maximum of two persons. | M C | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
61. | <M> | Each PWS shall be delivered in a reusable protective case that prevents damage to the instrument during transportation under normal conditions. A cardboard box with suitable protection can be used. | M C | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
62. | <D> | Tenderer should substantiate that the weight and dimensions of the PWS and its enclosure are such that it can be transported and installed by one person. Dutch legislation for occupational health and safety (ARBO) is applicable. The substantiation must state the PWS weight, dimensions and installation description. | D O | 3.2086 | Please provide a substantiation of maximum 1/4 A4 page. The substantiation will be assessed and a score of 0 or 5 will be given. Score of 5 if transportable by car and handling and installation by one person is possible within ARBO regulations. Score of 0 in all other cases. Note: the KNMI assessment team will seek internal advice from a specialist in Dutch occupational health and safety (ARBO) regulations when assigning the score for this Desirable. | Select applicable option: | The score will be assigned by the KNMI assessment team. | #VALUE!
4.2 Electrical Requirements
63. | <M> | The Tenderer shall guarantee electrically safe operation of the delivered PWS including its components according to Dutch and EU regulations. Compliance with electric safety standard DIN-EN 60204 and the IEC 60364 standard is preferred. | M C | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
64. | <M> | The PWS and related parts shall comply with the CE and EMC requirements, specifically European standards IEC 61000-6-2 (immunity) and IEC 61000-6-4 (emission) for industrial environments. Tenderer shall supply, upon request, the EMC test report of the offered PWS issued by an independent accredited organisation. | M C | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
65. | <M> | The PWS and related parts shall comply with European standard IEC 62305 concerning lightning protection and failure of electrical installations. | M C | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
66. | <M> | The PWS shall be able to operate using either (case 1) the public electricity in the Netherlands: 230 VAC 50Hz (NEN-EN-50160) with a maximum power consumption of the PWS, including heating, of 1 kiloWatt or less, or (case 2) a power supply voltage of 24 VDC ± 20% with a maximum power consumption of the PWS, including heating, of 150 Watt or less. | M I | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
67. | <D> | The PWS should be able to operate on a power supply 24 VDC ± 20% and the maximum power consumption of the PWS, including heating, should be 100 Watt or less. The maximum power consumption shall be provided. | D S | 1.6043 | Score given on maximum power consumption: <100 Watt = score 5 100 -113 Watt = score 4 113 -125 Watt = score 3 125 -138 Watt = score 2 138 -149 Watt = score 1 ≥150 Watt = score 0 Score of 0 is given if the PWS is unable to operate on 24 VDC ± 20%, regardless of actual maximum power. | Please provide the maximum power consumption of the PWS: | Score | #VALUE!
68. | <D> | The averaged power consumption of the PWS, including heating, should be specified by Tenderer. The offered solution will be evaluated based on the lowest power consumption. | D S | 0.80214 | Score given on average power consumption: <10 Watt = score 5 10-33 Watt = score 4 33-55 Watt = score 3 55-78 Watt = score 2 78-100 Watt = score 1 >100 Watt = score 0 | Please provide the average power consumption of the PWS: | Score | #VALUE!
4.3 Mechanical Requirements
69. | <M> | The PWS and related parts shall have water and dust protection IP65 or better according to norm IEC 60529. | M C | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
70. | <M> | The PWS shall be resistant to shock and vibration in normal operating conditions in all locations, including transport. | M C | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
71. | <D> | The PWS should comply with European standard IEC 60068-2 concerning shock and vibration resistance. | D C | 0.80214 | Yes = score 5 No = score 0 | Select your answer: | No answer selected. | #VALUE!
72. | <M> | The colour of the PWS and mast shall be in accordance with the EASA airport requirements. Specifically, the PWS shall be traffic red (RAL 3020), and the mast shall be both traffic red (RAL 3020) and traffic white (RAL 9016). | M C | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
73. | <M> | The material of the PWS and related parts shall be suitable for use in maritime and industrial environments. | M C | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
74. | <D> | The PWS should comply with European standard IEC 60068-2-52 concerning suitability for use in marine environments. | D C | 0.80214 | Yes = score 5 No = score 0 | Select your answer: | No answer selected. | #VALUE!
4.4 Installation Requirements
75. | <M> | It shall be possible either (case 1) to mount the PWS on the existing tiltable and frangible mast owned by KNMI by using an adapter and/or shortening of the mast and/or suitable mounting materials or (case 2) Tenderer shall include a tiltable and frangible mast in the offer. | M I | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
Notes for case 1: The top of the existing tiltable aluminium mast with a 76 mm diameter is at 2.060 or 2.675 meter (measured with respect to the concrete foundation) and the measurement sample area for visibility shall be at 1.885 and 2.50 meter, respectively. Technical drawings of the mast head, on which the PWS is to be fitted, are included in the sheet "Technical Drawings Mandatory-75" of this Annex 3. In this case, the Tenderer is not responsible for the colour of the masts as stated in Requirement 72. Note that the bolts mounting the existing mast (owned by KNMI) to the concrete foundation provide the required frangibility.
Notes for case 2: The mast supplied by Tenderer shall be tiltable and frangible and the colour of the setup shall be according to Requirement 72. It shall be suitable for installation on a concrete foundation with dimensions 42 x 46 cm. Furthermore, the measurement sample area for visibility shall be placed at 1.885 or 2.50 meter.
76. | <M> | It shall be possible, at locations where background luminance needs to be measured, to mount the background luminance sensor on the PWS or on the mast in such a way that it points north with an elevation of 20 to 30 degrees. | M C | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
77. | <M> | It shall be possible to mount the PWS in a flexible way on the mast such that it can be oriented optimally to avoid obstacle interferences. The mounting of the PWS shall be such that the orientation of the PWS is kept fixed when replacing the sensor. | M C | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
78. | <D> | Tenderer should specify the requirements of the PWS concerning the observation site and its vicinity. Tenderer should specify the optimal orientation and the allowed orientation range of the PWS to ensure that the measurements remain within the required measurement uncertainty. The Tenderer should at least address the requirements when using multiple PWS's close together for parallel measurements, as well as required distances from reflecting obstacles and surfaces including nearby PV installations. | D O | 3.2086 | Please provide a substantiation of maximum 1/2 A4 page. The substantiation will be assessed and a score is given between 0 to 5 maximum. The assessment is based on expert judgement with special attention given to the following assessment criteria: - The minimum distance to any obstacle, both reflecting and non-reflecting. - The constraints and requirements on the siting/surroundings. - The minimum distance between two sensors that allows operation without interference. | Select applicable option: | The score will be assigned by the KNMI assessment team. | #VALUE!
4.5 Sensor Connections Requirements
79. | <M> | The PWS shall be equipped with cables and connectors for use under exterior conditions according to European standards IEC 60204 and/or IEC 60364. These cables and connectors are available on the European market. | M C | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
80. | <M> | The PWS shall be provided with both the power cable and serial connection cable, and with connectors attached. Note that both power and serial connection is allowed to be combined in one cable. The length of the cables and type of the connectors shall be specified by KNMI in consultation with Tenderer before delivery. The cable length will not exceed 3 meters. | M C | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
4.6 Operating Environment Requirements
81. | <M> | The PWS shall be able to operate in all environmental conditions in The Netherlands, including the North Sea platforms and the Caribbean Netherlands. Specifically, it shall be able to operate according to the Requirements in the relative humidity range of 0 to 100%, the ambient temperature range of –40 to 50 °C and wind speeds up to 50 m/s. | M C | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
82. | <M> | The PWS shall not be damaged by environmental conditions in The Netherlands, including North Sea platforms and the Caribbean Netherlands. Damage caused by extreme weather events such as hurricanes and the debris associated with these events is exempt. | M C | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
5 Software Requirements
83. | <M> | Tenderer shall inform KNMI of each new software/firmware release for the PWS during the entire Contract period. This includes updates with additional features as well as bug fixes. Information shall be provided on the contents of the update, its consequences for the operation of the PWS, and whether or not the update is critical. | M C | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
84. | <M> | Tenderer shall supply new software/firmware updates in a format that can be used by KNMI. The installation of updates or upgrades can be performed by adequately trained KNMI staff (according to section 8 of this Annex 3 - Programme of Requirements and Desirables) during 2nd line maintenance and has no effect on warranty. | M C | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
85. | <M> | Tenderer agrees that KNMI shall be free to choose not to install any updates or upgrades without loss of warranty or 3rd line maintenance. In the case that an update is not installed, KNMI accepts the shortcomings reported by the manufacturer which are solved in the new software/firmware version. | M C | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
86. | <M> | Tenderer shall safeguard the PWS to avoid lockup situations in which components or the whole system becomes inoperative or unresponsive. | M C | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
87. | <D> | It should be possible to upgrade the PWS firmware or configuration remotely through the serial or network interface. The PWS should automatically revert to the previous firmware in case of a failure that occurs during the firmware upgrade. The settings and calibration of the PWS should remain intact during upgrade of the firmware. Downgrading to a previous firmware version should also be possible. | D C | 0.80214 | Yes to all = score 5 No to any = score 0 | Select your answer: | No answer selected. | #VALUE!
88. | <D> | Tenderer should state and substantiate whether any new firmware releases are planned that will lead to improvements of the PWS. | D O | 0.80214 | Please provide a substantiation of maximum 1/4 A4 page. The substantiation will be assessed and a score is given between 0 to 5 maximum. The assessment is based on expert judgement with special attention given to the following assessment criteria: - Quality and potential of expected enhancements with respect to the filtering algorithm to counteract the influence of flying insects and/or spiders and contamination. - Quality and potential of expected enhancements with respect to an upgraded precipitation type discrimination algorithm. | Select applicable option: | The score will be assigned by the KNMI assessment team. | #VALUE!
89. | <D> | Tenderer should substantiate how the PWS reports qualitative information about the reported precipitation types. For example, a probability distribution which indicates that precipitation is of a specific type. | D O | 1.6043 | Please provide a substantiation of maximum 1/4 A4 page. The substantiation will be assessed and a score is given between 0 to 5 maximum. The assessment is based on expert judgement with special attention given to the following assessment criteria: - Clarity, completeness and level of detail of the substantiation. - Level of detail of the quality information (e.g. probability of correct detection of the type). | Select applicable option: | The score will be assigned by the KNMI assessment team. | #VALUE!
90. | <D> | The PWS should make “raw” sensor data that can be used to verify the calibration of the PWS available. Raw sensor data is, for example, information on particle size, falling velocity, shape, water content as well as the information from which visibility and background luminance is derived. The availability of “raw” sensor data should be substantiated. The format of the raw data should be disclosed to KNMI. | D O | 1.6043 | Please provide a substantiation of maximum 1/4 A4 page. The substantiation will be assessed and a score is given between 0 to 5 maximum. The assessment is based on expert judgement with special attention given to the following assessment criteria: - Clarity, completeness and level of detail of the substantiation. - The amount, level of detail and type of raw sensor data available. | Select applicable option: | The score will be assigned by the KNMI assessment team. | #VALUE!
91. | <D> | Tenderer should substantiate the possibility to reprocess archived “raw” sensor data offline by either Tenderer or KNMI in order to optimize the algorithms that derive the meteorological quantities. | D O | 0.80214 | Please provide a substantiation of maximum 1/4 A4 page. The substantiation will be assessed and a score is given between 0 to 5 maximum. The assessment is based on expert judgement with special attention given to the following assessment criteria: - Clarity, completeness and level of detail of the substantiation. - The level of optimization that is facilitated/possible to the algorithms with archived raw data. | Select applicable option: | The score will be assigned by the KNMI assessment team. | #VALUE!
6 Reliability Requirements
92. | <M> | The PWS shall be COTS (Commercial Off The Shelf). Please note: The colour of the PWS is not considered part of this requirement (see Requirement 72). | M C | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
93. | <M> | The PWS shall be designed and able to operate continuously 24 hours a day, 7 days a week. | M C | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
94. | <M> | The life time of the PWS and spare parts shall be at least 10 years. | M C | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
95. | <M> | The Mean Time Between Failure (MTBF) of the PWS shall be at least 43 830 hours (=60 months), based on all environmental conditions and continuous 24/7 operation. | M C | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
96. | <D> | The Mean Time To Repair (MTTR) during 2nd line maintenance of the PWS should be specified and should be 4 person-hours or less including diagnosis, calibration and verification. | D S | 6.4171 | Score shall be given on the MTTR: <4 hours = score 5 4-7 hours = score 4 7-10 hours = score 3 10-13 hours = score 2 13-16 hours = score1 >16 hours = score 0 | Please specify the MTTR: | Score | #VALUE!
97. | <M> | The calibration interval of the PWS concerning MOR, background luminance, and the physical quantities from which the precipitation type and intensity are derived, shall be at least 1 year. | M C | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
7 Support and Maintenance Requirements
7.1 Support Requirements
98. | <M> | A helpdesk with skilled personnel fluent in English shall be available and reachable during working hours (9:00-17:00 CET) by telephone and e-mail. The helpdesk shall reply to questions within two working days. | M C | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
99. | <D> | Tenderer should specify the type of help and support that is made available to KNMI during the contract period. In particular, the following aspects are considered important: • An overview of corrective actions to rectify hardware/software bugs and performance limiting factors; • Specification of response time and the levels of support available during standard/helpdesk telephone inquiries inside and outside of office hours; • Specification of response time and levels of available support for detailed response to more complex inquiries (telephone/e-mail/written); • Escalation procedures for hardware/software problems which cannot be fixed in a timely manner by onsite personnel. | D O | 0.80214 | Please provide a substantiation of maximum 1/2 A4 page. The substantiation will be assessed and a score is given between 0 to 5 maximum. The assessment is based on expert judgement with special attention given to the following assessment criteria: - Clarity, completeness and level of detail of the substantiation. - The response times to both standard and complex inquiries relating to the offered PWS. - Clarity and practical usability of the escalation procedures for both hardware and software problems. | Select applicable option: | The score will be assigned by the KNMI assessment team. | #VALUE!
7.2 Maintenance Requirements
100. | <M> | The design of the PWS and the PWS setup shall be such that the cleaning and 1st line maintenance can be performed by one person without the use of a ladder. | M C | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
101. | <M> | Support and maintenance of the offered PWS and parts shall be available for the period of validity of the contract, with a minimum of 10 years and a maximum of 16 years. | M C | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
102. | <M> | The offered PWS, or compatible products meeting the same Requirements that can replace the original PWS, shall be available for the period of validity of the Contract, with a minimum of 10 years and a maximum of 16 years. | M C | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
103. | <M> | PWS sensors or parts that have been sent to the Tenderer for repair shall be returned after repair or replaced within 2 months (including FAT and transport). Tenderer shall specify the typical time required for a PWS repair. | M I | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
[6 empty] | Specify the typical repair time in this field. | No repair time specified.
104. | <M> | The serviceability of the PWS shall be based on the concept of Line Replaceable Units (LRU), which means that it is designed to be repaired quickly at KNMI by either (case 1) replacement of modular components, or (case 2) the PWS as a whole is considered the LRU. For case 1, Tenderer shall provide a list with recommended LRU items including recommended amounts. For case 2, the number of recommended spare PWSs shall be specified. Note that these spare PWS are not included in the 10% spares that KNMI requires for performing 2nd line maintenance. | M I | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
[6 empty] | Specify the requested list or # of additional spares here. | No answer specified.
105. | <M> | The PWS shall be able to operate in the field at rural sites for at least 6 months in between 1st line maintenance instances. | M C | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
106. | <M> | The PWS shall be able to operate on airfield locations and on North Sea platform locations for at least 2 months in between 1st line maintenance instances. | M C | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
107. | <M> | Tenderer shall offer, upon request after a provisional award of the Contract, a Maintenance Activities Plan (MAP) for the PWS. The MAP addresses maintenance activities that shall be performed to safeguard and maintain the correct functioning of the PWS and to ensure that the PWS meets the Requirements. The MAP shall include preventive maintenance activities and the required spare parts. The MAP shall establish Key Performance Indicators (KPI's) for monitoring and reporting the PWS performance. | M C | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
7.3 Delivery Requirements
108. | <M> | Tenderer shall commit to the delivery schedule and support required during the verification test as described in Annex 10. Verification Phase and Delivery. | M C | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
8 Training Requirements
109. | <M> | Training shall be made available to KNMI staff (approximately 20 persons in total) in order to enable them to perform 1st and 2nd line of maintenance on the offered PWS independently from the Tenderer. Organizational details of this training program are to be decided jointly by Tenderer and KNMI. | M C | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
110. | <M> | Tenderer shall train five staff members of KNMI as core instructors which shall be certified to train other KNMI staff to perform 1st and 2nd line maintenance in accordance with the training requirements of Tenderer. | M C | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
111. | <M> | The aforementioned training shall guarantee that the maintenance activities plan (MAP) can be carried out on the PWS by trained KNMI staff without impact on warranty and 3rd line maintenance during the Contract period. | M C | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
112. | <D> | Tenderer should provide a description of the contents of the training program. | D O | 0.80214 | Please provide a substantiation of maximum 1 A4 page. The substantiation will be assessed and a score is given between 0 to 5 maximum. The assessment is based on expert judgement with special attention given to the following assessment criteria: - Applicability of the training curriculum to the 1st line and 2nd line maintenance work that needs to be done by KNMI personnel - Clarity and logical structure of the training program - Completeness of the topics covered during the training | Select applicable option: | The score will be assigned by the KNMI assessment team. | #VALUE!
113. | <M> | Training and associated documentation shall be in Dutch or English. | M C | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
9 Documentation Requirements
114. | <M> | Tenderer shall include a comprehensive and complete set of manuals and technical documentation written in Dutch or English with instructions for operation, installation, maintenance and verification of the PWS. | M C | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
This set shall at least include information on: • Technical data/specifications; • Data output description; • Status information output description; • Data communication options; • Installation instructions; • Mechanical and electronic connections; • Electronic wiring schematic diagrams; • Electrical specifications of digital outputs, power supply, power consumption and temperature range of the instrument; • Measurement examples;
• Documentation on the internal algorithms and thresholds used in the PWS processing for the calculation of output quantities; • Documentation on the internal algorithm to determine the 1 minute (or less) precipitation type; • Description of internal data quality control algorithms; • Verification protocol; • Description of factory acceptance test; • Maintenance guidelines (on site and indoors); • Safety instructions; • Guidelines for transport; • Troubleshooting; • Training documentation.
115. | <M> | Tender shall supply new or updated documentation when new firmware releases or upgrades of the PWS are delivered to KNMI, or when incorrect or incomplete information in the current documentation set is found. | M C | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
116. | <M> | All documentation files shall be formatted in Portable Document Format (PDF). | M C | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
117. | <D> | The documentation set should specify how to implement, use and maintain the PWS. | D O | 0.80214 | Provide the documentation package in a separate file. The documentation will be assessed and a score is given between 0 to 5 maximum. The assessment is based on expert judgement with special attention given to the following assessment criteria: - Level of detail of the documentation - Completeness of the documentation - Ease with which information can be found. | Select your answer: | The score will be assigned by the KNMI assessment team. | #VALUE!

## Summary of Quality criteria
Summary of Quality criteria points
Chapter | # of desirables | Sub-total points achieved | Total points achieved
|  | (without KNMI-assessed desirables) | (with KNMI-assessed desirables)
1 - Functional | 12 | #VALUE! | #VALUE!
2 - Maintainability | 4 | n/a | #VALUE!
3 - Technical | 3 | #VALUE! | #VALUE!
4 - Hardware | 6 | #VALUE! | #VALUE!
5 - Software | 5 | #VALUE! | #VALUE!
6 - Reliability | 1 | #VALUE! | #VALUE!
7 - Support and Maintenance | 1 | n/a | #VALUE!
8 - Training | 1 | n/a | #VALUE!
9 - Documentation | 1 | n/a | #VALUE!
Totals | 34 | #VALUE! | #VALUE!
(maximum) |  | (76,2) | (750)
Please note: this column will only be filled in after the assessment by KNMI is completed.

## Technical Drawings Mandatory-75
Technical drawings of the current KNMI PWS mast - part of Mandatory #75

## Technical Area (hidden)
This sheet contains technical structures to assist with Excel Data Validation. Tenderers do not need to take any actions with respect to this sheet.
For all Mandatory Closed & Desirable Closed
Select your answer:
Yes
No
Desirable 6:
Please provide the actual range of the PWS:
50 kilometers
51 to 63 kilometers
63 to 75 kilometers
75 to 88 kilometers
88 to 99 kilometers
≥100 kilometers
Mandatory 18, 19, 75, 105 (Y/N+case)
Select your answer:
Yes, case 1 is applicable
Yes, case 2 is applicable
No
Mandatory 44, 66 (y/n + case where both can apply)
Select your answer:
Yes, case 1 is applicable
Yes, case 2 is applicable
Yes, both cases apply
No
Desirable 67
Please provide the maximum power consumption of the PWS:
≥150 Watt (at 24 VDC ± 20%)
138 to 149 Watt (at 24 VDC ± 20%)
125 to 138 Watt (at 24 VDC ± 20%)
113 to 125 Watt (at 24 VDC ± 20%)
100 to 113 Watt (at 24 VDC ± 20%)
<100 Watt (at 24 VDC ± 20%)
Unable to operate at 24 VDC ± 20%.
Desirable 68
Please provide the average power consumption of the PWS:
>100 Watt
78 to 100 Watt
55 to 78 Watt
33 to 55 Watt
10 to 33 Watt
<10 Watt
Desirable 97
Please specify the MTTR:
>16 hours
13 to 16 hours
10 to 13 hours
7 to 10 hours
4 to 7 hours
<4 hours
Scaling Factor for weights
0.16043
Substantiations Yes/No
Select applicable option:
Yes, Substantiation is provided in Annex 12.
No, Substantiation is not submitted.
Mandatory 36
Select your answer:
Yes, file is included.
No, file is not included.