## Appendix A05
Appendix A05, Conformity list 'Programme of Requirements' Guidelines for the European Public Tender for a single-photon detector, ref. TNO 2023 FPL/INK 029, dated 03-07-2023.
By the following Programme of Requirements TNO will assess if Tenderer meets the relevant requirements as laid down in Section 6 and 8. The Tenderer must indicate for each item, the conformity for the requirements in the table below.
If to any of the requirements, as stated in section 8, are not met by Tenderer, he will be excluded for any further assessment. By answering a question with “Yes, complies with requirement, see” will mean that in the event of awarding the contract, non-conformity to the requirement shall count as failure to meet an obligation from the Agreement to be concluded.
Requirement | Decription | Tenderer complies with requirement | Reference in Tenderer's documentation where compliances with the requirement is described | Remark
Requirement 8.1.1 | The single-photon detector consists of a 19-inch rack that houses three main systems: the cryostat system; the electronics system; and the computer. | Click the drop-down box to select
Requirement 8.1.2 | The cryostat system consists of a closed-cycle cryogenic system including the coldhead, compressor, water chiller, and temperature controller | Click the drop-down box to select
Requirement 8.1.3 | The coldhead contains the single-photon detector devices and is capable of interfacing those with the optical fiber and electronic cables (for driving and detection signals) | Click the drop-down box to select
Requirement 8.1.4 | The electronics system consists of driving electronics that bias the detectors and time-tagging electronics that interpret the detection signals. | Click the drop-down box to select
Requirement 8.1.5 | The computer shall carry all the softwares required to communicate with the individual systems and extract current system status: temperature of coldhead; pressure of the cold head; status of SNSPD (latched/unlatched); photon counts. | Click the drop-down box to select
Requirement 8.1.6 | The cryostat shall be capable of supporting 16 single-photon detectors as well as the fibers and electrical cables required to interface with them | Click the drop-down box to select
Requirement 8.1.7 | 8 of the detectors shall be optimized for 1550nm with a detection efficiency higher than 90% | Click the drop-down box to select
Requirement 8.1.8 | The maximum acceptable dark count rate for the 1550nm detectors shall be 100 counts per second at a condition of 90% efficiency | Click the drop-down box to select
Requirement 8.1.9 | The maximum acceptable jitter for the 1550nm detectors shall be 50 ps at a condition of 90% efficiency | Click the drop-down box to select
Requirement 8.1.10 | The maximum recovery time for the 1550nm detectors shall be 80ns | Click the drop-down box to select
Requirement 8.1.11 | The optical fibers connected to the 1550nm detectors shall be SMF-28 | Click the drop-down box to select
Requirement 8.1.12 | 8 of the detectors shall be optimized to a wavelength region between 600 and 850nm with an averaged detection efficiency higher than 70% over this bandwidth | Click the drop-down box to select
Requirement 8.1.13 | The maximum acceptable dark count rate for the 600-850nm detectors shall be 10 counts per second at a condition of 70% efficiency | Click the drop-down box to select
Requirement 8.1.14 | The maximum acceptable jitter for the 600-850nm detectors shall be 50 ps at a condition of 70% efficiency | Click the drop-down box to select
Requirement 8.1.15 | The maximum recovery time for the 600-850nm detectors shall be 60ns | Click the drop-down box to select
Requirement 8.1.16 | The optical fibers connected to the 600-859nm detectors shall have a core diameter of 4.5um | Click the drop-down box to select
Requirement 8.1.17 | The efficiency and dark count curves as a function of bias current for each detector shall be provided together with the single-photon detector. | Click the drop-down box to select
Requirement 8.1.18 | The single-photon detector (including the compressor and cryogenic cables) shall fit in a 19'' rack | Click the drop-down box to select
Requirement 8.1.19 | The single-photon detector shall include driving electronics that directly interface with all the 16 detectors and allow characterizing and operating each detector individually | Click the drop-down box to select
Requirement 8.1.20 | The single-photon detector shall include time-tagging electronics capable of processing the detection signals from all 16 detectors | Click the drop-down box to select
Requirement 8.1.21 | The software that interfaces with the driving electronics shall include a user interface and pre-defined characterization measurements, such as an I/V curve for each detector | Click the drop-down box to select
Requirement 8.1.22 | The software that interfaces with the driving electronics shall include a user interface and basic measurement functions, such as correlations between any combination of detector channels | Click the drop-down box to select
Requirement 8.1.23 | The single-photon detector shall include a single computer that connects to both driving and time-tagging electronics and provides an interface between the user and the system | Click the drop-down box to select
Requirement 8.1.24 | All interfaces between computer and hardware of the single-photon detector shall include a python API | Click the drop-down box to select
Requirement 8.1.25 | The single-photon detector shall be installed by the manufacturer at the TNO site designated for it | Click the drop-down box to select
Requirement 8.1.26 | The compressor shall be water cooled | Click the drop-down box to select
Requirement 8.1.27 | A closed-cycle water chiller shall be part of the single-photon detector and shall be connected to the water-cooled compressor | Click the drop-down box to select
Requirement 8.1.28 | The rack structure that supports the single-photon detector shall provide compressor sound noise damping | Click the drop-down box to select
Requirement 8.1.29 | The rack shall include wheels with breaks in case the single-photon detector needs to be moved | Click the drop-down box to select
Requirement 8.1.30 | The single-photon detector shall have a reachable patch panel to allow the user to freely connect the fibers and electrical cables to their setup | Click the drop-down box to select
Requirement 8.1.31 | The 19-inch rack shall have a maximum height of 2.2 meters | Click the drop-down box to select
Requirement 8.2.1 | Terms and conditions of payment • 20% at order, after receipt of a bank guarantee covering 80% of the total amount; • 20% after Critical Design Review (CDR); • 40% after delivery and visual inspection at location TNO Delft. • 20% after installation and commissioning On-site and Site Acceptance Test (SAT) including test periods, approved by TNO, Management Quantum Technology and Procurement. At this moment TNO will return the bank guarantee. The Tenderer to whom the Contract will be awarded must provide a bank guarantee to TNO, covering 80% of the total price. The Bank guarantee and the bank who issues the Bank guarantee have to be approved by TNO (minimal A rating). TNO must have the right to draw the bank guarantee in any and all cases where Tenderer falls short in fulfilling its obligations under the Contract. The validity of the bank guarantee may not expire until approval of the SAT by TNO Management and Procurement. Remark: please note, that approving the CDR this does not mean that TNO will take final responsibility if during SAT it is found that the system does not meet the minimum requirements. TNO will at all times hold Tenderer ultimately responsible for delivering an single-photon detector in accordance with the requirements as stated in this Tender. | Click the drop-down box to select
Requirement 8.2.2 | Delivery shall be made in accordance with the Incoterms 2020, Delivered Duty Paid (DDP). | Click the drop-down box to select
Requirement 8.2.3 | A minimum of 3 TNO employees will be trained to operate the system safely and properly, as well as on basic preventive and corrective maintenance. Rationale: to solve small issues independently, with supplier approval, and minimise downtime. | Click the drop-down box to select
Requirement 8.2.4 | The single-photon detector shall be CE certified. | Click the drop-down box to select
Requirement 8.2.5 | The offered single-photon detector assembly and associated (necessary) accessories comply at all times with the applicable Dutch and European legal requirements, such as (but not limited to): NEN, EN and CE standards and their use including the working conditions - legislation. (93/42 / EEC); | Click the drop-down box to select
Requirement 8.2.6 | Documentation: • Instruction manual; • Maintenance manual including all required technical documentation for Preventive maintenance and Corrective Maintenance; • Quick- start guide (if available); • List of spare parts/consumables, including price for TNO; • Long lead times; • CE certification; • Safety measures and their implementation using the setup. All documentation must be available in English. | Click the drop-down box to select
Requirement 8.2.7 | Indicate if the instrument/ system is export controlled, and if so, please indicate if applicable the parts that are export controlled. | Click the drop-down box to select
Question 8.2.8 | Are there any additional parts that are or might become export controlled which we need to take into account for future maintenance and repair?
Requirement 8.2.9 | Certificate of Origin shall be provided with the shipment. | Click the drop-down box to select
Requirement 8.2.10 | Site Acceptance Test (SAT) - The SAT contains at least the following checkpoints: • The single-photon detector will be visually inspected for damages; • The proper functioning of the equipment will be checked: o Mechanical functions, o Electrical safety- and switching functions, o User interface(s), including any software. • All requirements (such as straightness etc.) and possible additional requirements will be checked; • The presence and quality of the documentation will be checked The details of the site acceptance test shall be agreed between TNO and the Tenderer to whom the Contract will be awarded, but Tenderer must submit a draft test protocol for the site acceptance test with its Tender. | Click the drop-down box to select
Requirement 8.2.11 | Warranty, as per Art. 6 of TNO’s Purchasing Conditions for Goods, and starts after a successful SAT | Click the drop-down box to select
Requirement 8.2.12 | Basic maintenance contract proposal for periods of one year, up to and including 2028 (see para. 1.3.2) | Click the drop-down box to select
Question 8.3.1 | Corporate Social Responsibility is becoming increasingly important in our own organization and in the future TNO will also ask more from our suppliers to play a part in this. What possibilities have you already initiated and where do you see opportunities for yourself and for TNO?
Question 8.3.2 | As you're aware, ISO 20400 is an international standard that provides guidelines on sustainable procurement. The standard can be used by any organization that intends to improve their social, economic, and environmental sustainability. Considering the growing awareness with regard to sustainable business practices, nationally and internationally, whether required by laws and regulations, promoted by the UN Sustainable Development Goals, or even by the changing societal expectations – there are a myriad of reasons in today’s world that drive organizations to be more sustainable, including TNO. Do you embrace the ISO 20400 guideline, and if so, how do you implement it in your own organisation?
Question 8.3.3 | Laboratory instruments, like the single-photon detector, that TNO uses for its scientific research often consume a lot of energy and also often consist of very high-quality raw materials. For the purpose of corporate social responsibility (CSR), do you, or will you in the future, make technological investments to reduce power consumption during instrument use, without negatively affecting the technical application of the instrument?
Question 8.3.4 | The design of products with recycling in mind should be an important place to start for all of us. The inclusion of recycled or reused materials is a good way to increase demand for recycled and reused materials as it can significantly reduce the greenhouse gas emissions impact of the materials used in such laboratory instruments. Can you give TNO some insight on how and in which way you are using recycled or reuse materials now, or are working on solutions to use more recycled or reuse materials in the future, but still safeguard the basic use of such laboratory instruments?
Question 8.3.5 | TNO focuses within its organization on reducing energy, water use and emissions. What greening initiatives have you taken or can you take in your product portfolio to reduce energy and water consumption and emissions?
Question 8.3.6 | In the case of transport packaging, one important opportunity is that of substituting reusable packaging for single trip packaging. Can you give TNO some insight on how and in which way you are using recycled or reuse materials now, or are working on solutions to use more recycled or reuse materials in the future?
Question 8.3.7 | In the case of shipping, from the sold products to the customer, we all have an important possibility to reduce emissions. Can you give TNO some insight into how you are trying to reduce the greenhouse gas impact in this perspective?

## validatie (hidden)
Click the drop-down box to select
Yes, complies with requirement, see:
No, does not comply with the requirement.
note: dit blad 'verbergen' indien je dit gaat delen intern/extern (behoudens aan inkoop collega's)