| # | Requirement | Required proof | Evaluation criteria | Reason for exclusion |
| --- | --- | --- | --- | --- |
| 4 | Three similar depositions of a metal-oxide or hetero-structures should result in three samples where the sample-to-sample variance in average thickness is preferably not more than 1% (for both small and large area PLD, consecutive depositions). | Experimental data of a Tenderer defined experiment in which metal-oxide films or hetero-structures of at least 50 nm in thickness are sequentially deposited with a deposition temperature of at least 700 degrees Celsius that shows that sample-to-sample variance in average thickness is preferably not more than 1% of the average film thickness. The data should include all relevant process conditions, such as partial-pressures, substrate temperature, and laser fluence, etc. For 4-inch substrates a minimum of 10 thickness measurement points are required scattered uniformly across the substrate surface. Proving homogenous epitaxial growth with x-ray lines across the film or hetero-structure surface is considered sufficient as well. Equal or better results shall be shown during a site acceptance test. | Points are assigned according to the following rules: - 9 points if sample-to-sample thickness variance is <= 1% - 3 points if 1% < sample-to-sample thickness variance <= 5% - 1 points if 5% < sample-to-sample thickness variance <=10% - 0 points if 10% < sample-to-sample thickness variance Tenderers may receive zero points if the provided proof is considered to be insufficient for proper evaluation. | The tenderer will be excluded if no data is provide that proofs that the sample-to-sample thickness variance is less than 20% or no equivalent proof of sufficient control of film homogeneity is provided. |
| 5 | The small and large area PLD systems shall ensure thickness uniformity in minimum of state-of-the-art research tools, while other material properties are not affected by the scaling to larger substrate areas. | Experimental data of a Tenderer defined experiment in which metal-oxide films or hetero-structures of at least 50 nm in thickness are deposited with a deposition temperature of at least 700 degrees Celsius that shows that the thickness variance across the film is preferably at most 1% of the average film thickness. The data should include all relevant process conditions such as partial-pressures, substrate temperature, and laser fluence. For 4-inch substrates a minimum of 10 thickness measurement points are required scattered uniformly across the substrate surface. Proving homogenous epitaxial growth with x-ray lines across the film or hetero-structure surface is considered sufficient as well. Equal or better results shall be shown during a site acceptance test. | Points are assigned according to the following rules: - 9 points if thickness variance is <= 1% - 3 points if 1% < thickness variance <= 5% - 1 points if 5% < thickness variance <=10% - 0 points if 10% < thickness variance Tenderers may receive zero points if the provided proof is considered to be insufficient for proper evaluation. | The tenderer will be excluded if no data is proved that proof that the thickness variance is less than 20% or no equivalent proof of sufficient control of film homogeneity is provided. |