## Appendix A05
Appendix A05, Conformity list 'Programme of Requirements' Guidelines for the European Public Tender for a scanning electron microscope (SEM), ref. TNO 2023 FPL-INK 029, date 10-07-2023.
By the following Programme of Requirements TNO will assess if Tenderer meets the relevant requirements as laid down in Section 6 and 8. The Tenderer must indicate for each item, the conformity for the requirements in the table below.
If to any of the requirements, as stated in section 8, are not met by Tenderer, he will be excluded for any further assessment. By answering a question with “compliant” will mean that in the event of awarding the contract, non-conformity to the requirement shall count as failure to meet an obligation from the Agreement to be concluded.
Requirement | Decription | Tenderer is compliant to the requirement | Reference in Tenderer's documentation where compliances with the requirement is described | Further clarification or remark
Requirement R-0000-005 | Meet all requirements The SEM supplier shall deliver a system that is capable of meeting all the stated requirements, and accepted preferences, in this document at the same time, unless indicated otherwise in a specific requirement. | Click the drop-down box to select
Requirement R-0000-010 | Human, machine and sample safety The SEM shall ensure, human, machine and sample safety. Remark: close-out of this requirement is done by review of the provided FMEA and demonstration that the implemented mitigation actions function, therefor VM is demonstration. | Click the drop-down box to select
Requirement R-0000-015 | Operation lifetime The SEM system shall have an operation lifetime of at least 10 years. Remark: Close out shall be done through review of design, proof of heritage, used materials and equipment. | Click the drop-down box to select
Requirement R-1000-005 | Source for imaging and material analysis The SEM shall have a source that matches best the stated requirements on both imaging (Requirement R-1000-025) and material analysis (Requirement R-2000-005). | Click the drop-down box to select
Requirement R-1000-010 | Electron optical column The Electron optical column shall enable the SEM to meet the required imaging (Requirement R-1000-025) and material analysis (Requirement R-2000-005) requirement stated in this document. | Click the drop-down box to select
Requirement R-1000-015 | Source lifetime demand The SEM source shall have a minimum lifetime of 2 years when operated 1000 hours per year. Remark 1: close out of requirement by proof from similar systems in use, or free repair when the source fails before stated requirement. Remark 2: lifetime is stated a time before full replacement of source, not tip maintenance or exchange. The supplier shall state the expected time before tip maintenance of tip exchange is needed. | Click the drop-down box to select
Requirement R-1000-025 | Imaging quality demand The SEM shall meet the stated requirements on imaging as stated in Appendix D001. Remark: this consists for example, but not limited to: - resolution, - brightness, - tuneability of parameters like voltage, current. | Click the drop-down box to select
Requirement R-1000-035 | Imaging detectors TNO concluded that both, different names and, slightly different implementations are used for similar detectors. To ensure equal opportunities for every supplier, the SEM shall be equipped with the detectors needed to generate the required measurements as stated in Appendix D001, and the minimum required detectors stated in Appendix D002. Remark 1: In the tender document the used/ offered detectors shall be stated, and the detectors that are not offered. Remark 2: VM is stated as analysis, this because the detectors will be tested, and the results analysed on the stated performance/ function. First check will be review of the offered detectors. | Click the drop-down box to select
Requirement R-1000-040 | Charging samples The SEM shall have one or more mitigations in place to measure a sample as described in Appendix D003. Remark 1: Several mitigation options are present in the market, TNO does not have a specific preference for the proposed solution, and several mitigation solutions may be used for the specific best solutions for the sample, or the measurement type. Remark 2: close-out of this requirement is verification that stated samples can be measured according to requirement R-1000-025. | Click the drop-down box to select
Requirement R-1000-045 | Beam acceleration voltage on sample between 0.1 keV and 30 keV The beam energy on the sample shall be able to be controlled between 0.1 keV and 30 keV. | Click the drop-down box to select
Requirement R-1000-055 | Horizontal field width at eucentric height The SEM shall have a maximum horizontal field width (HFW), at eucentric height, of at least 2 mm, both for SE and BSE imaging. | Click the drop-down box to select
Requirement R-1000-057 | Beam current setting The SEM shall have the possibility of tuning the beam current stepless. | Click the drop-down box to select
Requirement R-1000-060 | Imaging multiple channels simultaneously demand It is expected that the SEM has several detectors, the SEM shall be able to image at least 2 detectors simultaneously. Remark: imaging simultaneously is not equal to showing all the images simultaneously, TNO accepts that not all images will be displayed at the same time. | Click the drop-down box to select
Requirement R-1000-070 | Segmented detector The SEM shall be equipped with at least one segmented detector, which can be read out per segment to be able to obtain detailed topographical information. Remark: in the tender document the supplier shall state the amount of segments. | Click the drop-down box to select
Requirement R-1000-085 | Imaging integration settings In the SEM software it shall be possible to set or tune, at least the following imaging settings: • Pixel-, line-, frame-integration options. • Drift correction option. • Dwell time setting. Remark: It is expected that this can all be done from within the SEM software GUI. | Click the drop-down box to select
Requirement R-2000-005 | Material analysis (EDX) The SEM shall be delivered with an integrated, and functional EDX, with the requirements stated within Appendix D005. Remark 1: TNO opted to state the requirements in one Appendix document, this because TNO is of the opinion that all suppliers can deliver an EDX to the stated requirement, this by either installing a COTS EDX, or install an in-house developed EDX. Remark 2: the supplier shall state in the tender document the selected detector, with all the specifications. | Click the drop-down box to select
Requirement R-2000-010 | Material analysis samples The EDX of the SEM shall be able to measure the stated samples in Appendix D003 Remark: including in this sample list, charging samples are present. It is expected that charge mitigation will be required. This may be done by so called low vacuum SEM. | Click the drop-down box to select
Requirement R-2000-015 | Minimum detector surface EDX The minimum detector surface area shall be 60 mm^2. | Click the drop-down box to select
Requirement R-3000-005 | Sample size loading demand The SEM shall be able to load the samples that are stated in Appendix D003. | Click the drop-down box to select
Requirement R-3000-015 | Sample measurement area demand The SEM shall be able to measure the samples in the stated measurement area in Appendix D003 without making use of stage rotations. Remark: it is expected that an XYZ translation of the sample stage will be used to position the sample to measure at the required measurement spot | Click the drop-down box to select
Requirement R-3000-030 | Sample measurement at tilt demand The SEM shall be able to tilt, and measure, the specified samples, in the specified measurement area, and to the specified angle as stated in Appendix D003. Remark: The supplier shall indicate how the design of the sample stage will prevent the samples from movement during tilting. | Click the drop-down box to select
Requirement R-3000-045 | Sample weight demand The SEM shall be perform all stated SEM requirements with a sample weight up to 0,5 kg, with the exemption of the tilt motion. Remark: this stated sample weight is excluding sample holders weight that are required for normal SEM usage, with the stated samples in Appendix D003, and if applicable Appendix D004, and therefor are in the delivery of the SEM. | Click the drop-down box to select
Requirement R-3000-055 | Sample weight including tilt motion demand The SEM shall be perform all stated SEM requirements with a sample weight up to 0,5 kg, including the tilt motion. Remark: this stated sample weight is excluding sample holders weight that are required for normal SEM usage, with the stated sample in Appendix D003, and if applicable Appendix D004, and therefor are in the delivery of the SEM. | Click the drop-down box to select
Requirement R-3000-065 | Mounting of stated samples The SEM shall enable mounting of samples as stated in Appendix D003, and if applicable for the samples stated in Appendix D004. Remark: The supplier shall state in the tender documents how the supplier intends to mount the samples. | Click the drop-down box to select
Requirement R-3000-070 | Mounting of sample time needed demand The SEM shall enable the mounting of the sample on the sample stage within 10 minutes for all the samples indicated in Appendix D003 and if applicable Appendix D004. Remark: this is the time without pump down of the system. Sample holders that enable easy mounting of the sample will reduce the time required for mounting the samples, and will reduce the risk of sample damage and will safe time during loading. | Click the drop-down box to select
Requirement R-3000-076 | Mounting interface The SEM supplier shall share the interface description to the sample stage in the tender document. Remark: detailed 2D and STEP files are expected to be part of this delivery. | Click the drop-down box to select
Requirement R-3000-080 | Mounting of non-stated samples The SEM shall ensure that non-stated samples can be mounted safely. Remark 1: This may be achieved by ensuring a (large) flat surface that enables the sample to be secured with carbon tape. Remark 2: The supplier shall state in the tender documents how the supplier intends to mount non-stated samples. Remark 3: it is expected that this mounting can be done within 10 minutes, if this is not expected the supplier shall put forward this in the tender document. | Click the drop-down box to select
Requirement R-3000-085 | Sample stage rotation The SEM shall be able to rotate the stated samples in Appendix D003 with a range of +/- 180°. | Click the drop-down box to select
Requirement R-3000-086 | Compucentric sample stage rotation The SEM shall be able to perform compucentric rotations of the stage, when requested by the operator. | Click the drop-down box to select
Requirement R-3000-090 | Sample position repeatability demand With the stated sample motions of requirement R-3000-015, and the stated sample weight in requirement R-3000-045, the repeatability of the sample stage position shall be better than or equal to 3 um. Remark 1: this requirement directed to the demanded measurement areas and sample weight. Unless indicated differently TNO expects that when a demanded requirement are upgraded to a desired requirement, the stated value of this requirement still applies. If this is not the case, the supplier shall make this clear in the tender document. Remark 2: repeatability is defined as; how well does the sample stage does return back to previous position. Remark 3: this requirement is without unloading the sample Remark 4: this requirement is without tilt of the sample. | Click the drop-down box to select
Requirement R-3000-100 | Stage position accuracy demand With the stated sample motions of requirement R-3000-015, and the stated sample weight in requirement R-3000-045. The accuracy of the sample stage shall be better than or equal to 3 um over a travel range of 100 mm in either X or Y direction. Remark 1: this requirement direct to the demanded measurement areas and sample weight. Unless indicated differently TNO expects that when a demanded requirement is upgraded to a desired requirement, the stated value of this requirement still applies. If this is not the case, the supplier shall make this clear in the tender document. Remark 2: accuracy is defined as; how well does the sample stage know it absolute position. Remark 3: this requirement is without unloading the sample Remark 4: this requirement is without tilt of the sample. | Click the drop-down box to select
Requirement R-3000-110 | Resolution of stage movements The control of the stage shall have the following resolution: - XYZ= 0.0001 mm - R = 0.1° - Tilt =0.5° Remark: We expect that both can use absolute coordinates and relative steps can used to move the stage via the GUI. | Click the drop-down box to select
Requirement R-3000-115 | Analysis position readout The SEM shall display a readout of all axes of the analysis position (i.e. at least x, y, z, R and θ) in the SEM software. | Click the drop-down box to select
Requirement R-4000-005 | Contamination growth demand The SEM shall meet the stated values, with the stated method in appendix D006 with respect to contamination growth. Remark: if the stated method is not possible perform, this shall be stated during the question period of the tender phase, with a counter proposal how to obtain the values TNO seeks, TNO will investigate to proposal. | Click the drop-down box to select
Requirement R-4000-015 | Cleaning of SEM with plasma cleaner Remark 1: Because TNO expects that this a common component on a SEM, TNO expect that either COTS plasma cleaner will be offered, or an “inhouse” developed cleaner. Remark 2: Because of the expected common component, no specific performance requirements are set. TNO however expects, in example, an Evactron model 25 or 50, or a plasma cleaner similar performance. Remark 3: this plasma cleaner does not necessary be directly connected to the measurement chamber of the SEM Remark 4: Ideally the plasma cleaner is permanently installed on the SEM Remark 5: the tender document will state the selected plasma cleaner, the location, and used method of plasma cleaning. | Click the drop-down box to select
Requirement R-4000-025 | Forbidden material list The SEM shall not use the stated materials in appendix D007. Remark: if a material is listed in appendix D007 that is needed for functional or performance, this shall be ideally stated in the tender document, but at latest in the CDR documentation. TNO expects that mitigation options are present to resolve the used material conflict. | Click the drop-down box to select
Requirement R-4000-035 | Pump down time with sample demand The SEM shall reach the required high vacuum level for SEM operation within 10 minutes after loading a sample as stated in Appendix D003 and D004 from atmosphere with a sample that has been exposed to atmospheric conditions for at least 24 hours. | Click the drop-down box to select
Requirement R-4000-045 | Venting time demand The venting time, including all internal SEM actions that are required, shall be within 10 minutes after giving the venting signal. Remark: Internal SEM actions could be stop and cool down of SEM source. | Click the drop-down box to select
Requirement R-4000-065 | Automated pump down and venting The SEM shall, after receiving the signal to pump-down or vent, execute all steps to needed to for fill this command without further instructions from the operator. Remark: it is expected that giving the instruction to pump down or vent can be done with a very simple instruction, like a button pump down or vent in the HMI. | Click the drop-down box to select
Requirement R-4000-070 | Prevention against damage due to over- pressure The SEM shall prevent the system against damage due to over-pressure. I.e., source on at a too high pressure will impacting lifetime. A second example would be that during venting the entire vacuum vessel is set on an over-pressure resulting in, in example windows will burst Remark: as stated previously, TNO expects a COTS SEM, and assumes that the suppliers have adamant experience in manufacturing a save SEM. The proposed mitigations against the risks shall be stated in the tender document. | Click the drop-down box to select
RequirementR-4000-075 | Dry roughing pump The SEM shall use dry roughing pump(s) Remark: Dry rough pumps are considered roughing pumps that prevent oil contamination reaching the SEM, this could be by not having oil in the vacuum part of the roughing pump, or shielding this i.e. with N2 flow from the vacuum side of the pump. | Click the drop-down box to select
Requirement R-5000-005 | External resources specifications The SEM system shall meet its requirements while complying to the following external resource specifications: - GN2 supply pressure = up to 6 bar(a) - CDA supply pressure = up to 7.5 bar(a) - Cooling water supply temperature <18 °C - Maximum cooling water return temperature = 30 °C - Maximum cooling water flow available is 500 l/hour - An electric power supplied of; 240 volt, 50 hrz, and/or 400 volt, 3 phase, 50hrz. | Click the drop-down box to select
Requirement R-5000-010 | Electric power connection fusing The electric power connections of the SEM system shall have double poled fusing or circuit breakers (fuses placed in both the phase and neutral conductors). | Click the drop-down box to select
Requirement R-5000-015 | Protected earth grounding The SEM system shall be provided with at least one protected earth groundings to enable connection to protected ground connection in the VLL. Remark 1: It is expected that each electric cabinet has a protected ground connection. Remark 2: Within the VLL dedicated grounding points are available. | Click the drop-down box to select
Requirement R-5000-020 | Power consumption during normal operation The power consumption of the SEM during normal operation shall not be more than 5 kW. Remark: The power consumption and all electrical requirements shall be listed in the quotation of the supplier. | Click the drop-down box to select
Requirement R-5000-025 | Provided gas cleanliness The SEM shall be able to operate, to the stated requirements (for instance cleanliness) with a GN2 and XCDA quality provided by the VLL facility of 99.999% purity. Remark: If a higher quality is required this may be solved by using end point filters located on the gas distribution boards. These end point filters are then part of the delivery of the SEM system. The tenderer shall put forward during PDR the expected lifetime of these end point filters. | Click the drop-down box to select
Requirement R-5000-030 | Location of the roughing pump(s) The roughing pump(s) shall be located in the VLL corridor. Remark 1: the supplier shall state the maximum distance that the roughing pump can be away from the SEM in the tender document. Remark 2: the lay-out of the SEM in the VLL, will be agreed on in an interface control document during the PDR meeting. | Click the drop-down box to select
Requirement R-5000-035 | Roughing pumps - connection to exhaust The roughing pumps shall allow for connection to a central exhaust in the VLL building. Remark: The SEM supplier shall state the flange type and dimensions of the pre-pump(s) in the ICD documentation. | Click the drop-down box to select
Requirement R-5000-040 | Maximum weight The floor load per 1 m2 of both the SEM and the support equipment shall not exceed 1000 kg. | Click the drop-down box to select
Requirement R-5000-045 | Required volume claim At the PDR, or preferably in the tender document, the supplier shall specify the required volume claim in the cleanroom and in the VLL corridor of the entire SEM system. | Click the drop-down box to select
Requirement R-5000-050 | Required volume claim for maintenance At the PDR the supplier shall specify the required volume claim for maintenance. | Click the drop-down box to select
Requirement R-5000-055 | SEM electronics cabinet All electronics shall be installed in an electronics cabinet. | Click the drop-down box to select
Requirement R-5000-060 | Clean room conditions The SEM system and the support equipment located in the cleanroom shall be compatible with operating under cleanroom conditions (room temperature 21 °C +/- 1 °C, humidity between 30% and 70%). | Click the drop-down box to select
Requirement R-5000-065 | VLL corridor conditions The support equipment of the SEM system that is located in the VLL corridor shall be compatible with operating under the VLL corridor conditions (room temperature 15 °C to 30 °C, humidity between 30% and 70%). | Click the drop-down box to select
Requirement R-5000-070 | Cleanroom compatibility All external surfaces of the SEM shall be easily cleanable, i.e. smooth surfaces and no hidden areas. | Click the drop-down box to select
Requirement R-5000-075 | External cleanliness The external surface of the SEM and support equipment shall be free of grease, oils, and other manufacturing residues. | Click the drop-down box to select
Requirement R-5000-080 | Particle contamination on external surfaces The SEM and the support equipment shall be visibly clean, i.e. they shall have a Particle Area Coverage (PAC) of less than 250 ppm. Remark 1: The SEM and support equipment may be cleaned by the supplier before entering the clean room. Remark 2: TNO believes that this requirement will be met if the supplier takes care to use smooth and easy to clean materials. | Click the drop-down box to select
Requirement R-5000-085 | Restrict amount of air-cooled equipment The amount of active/forced air-cooled equipment needed for the SEM that is located inside the ISO 14644 class 6 cleanroom shall be kept to a minimum. Remark: In case active/forced air-cooled equipment inside the cleanroom is needed to operate the system, this shall be listed and put forward by the supplier during PDR. | Click the drop-down box to select
Requirement R-5000-090 | Noise output cleanroom The noise output of the SEM system and support equipment located in the cleanroom shall be less than 75 dB(A) in the cleanroom Remark: The preferred solution is to select equipment that produces less than the required 75 dB(A). Isolation or locating equipment in the VLL corridor is allowed. | Click the drop-down box to select
Requirement R-5000-095 | Grouped cables The cables both in vacuum and outside vacuum shall be grouped together as much as possible. | Click the drop-down box to select
Requirement R-5000-100 | Labelled cables The cables outside vacuum shall be labelled. | Click the drop-down box to select
Requirement R-6000-005 | Computer controlled SEM operation All operations of the SEM system shall be computer-controlled. Remark: After awarding the tender, TNO and supplier shall agree on the best way to configure and install one or multiple computers for the SEM and EDX software and how it should be integrated into the TNO network. This will be documented in the ICD. | Click the drop-down box to select
Requirement R-6000-010 | Computer username and password The computer used for controlling the SEM shall use username and password to protect the SEM against unauthorized use. | Click the drop-down box to select
Requirement R-6000-015 | Multiple user profiles The control system of the SEM system shall have at least four user profiles. Remark: It is expected that each operator shall be able to log in on their own user account, or profile, with their own user preferences in the SEM and EDX software. | Click the drop-down box to select
Requirement R-6000-025 | Navigation on sample The SEM shall be equipped with a method to navigate across the loaded sample using an image of the loaded sample. Remark: it is expected that this at least can be done using a camera, see requirement R-6000-030. | Click the drop-down box to select
Requirement R-6000-030 | Camera on SEM The SEM shall be equipped with a permanently installed camera that is able to obtain an image that can be used for navigation across the loaded sample. | Click the drop-down box to select
Requirement R-6000-035 | Navigation camera field of view The camera stated in R-6000-035 shall be able to make an image with field of view dimensions of 10 cm x 10 cm. Remark 1: the supplier shall state the field of view in the tender document. Remark 2: Ideally imaging all the stated samples in D003 is preferred. | Click the drop-down box to select
Requirement R-6000-040 | Navigation camera resolution The navigation camera image shall be able to resolve details on a sample of 0.1 mm or smaller for the defined wafer samples as stated in Appendix D003. | Click the drop-down box to select
Requirement R-6000-050 | Navigation image alignment The SEM software shall be able to perform an alignment procedure between the navigation image (Requirement R-6000-025) and the SEM position on the sample. Remark: This may be achieved by performing a 3 point alignment procedure, linking the images with the navigation image. | Click the drop-down box to select
Requirement R-6000-055 | Scan rotation The SEM shall be able to perform scan rotation, to align to features in the obtained image without having to physically rotate the sample stage. | Click the drop-down box to select
Requirement R-6000-060 | Scan rotation range and resolution Scan rotation (Requirement R-6000-055) shall be possible over 360 degrees. Remark: It is expected that the scan rotation can be set with an accuracy of 0.1 degrees. | Click the drop-down box to select
Requirement R-6000-065 | Find back previous imaging locations The system shall be able to return to locations of interest, or locations where images were previously made, on a sample. Remark: This may be done, for example, by saving a location of interest in a positions list embedded in the GUI. | Click the drop-down box to select
Requirement R-6000-075 | Find back previous imaging locations on sample accuracy The SEM shall be able to find back image locations made in a previous imaging session, with an accuracy of 10 um or better. Remark: This may be solved by for example setting up a custom coordinate system based on three easy to find markers on the sample, after which image locations expressed in this coordinate system can be saved. | Click the drop-down box to select
Requirement R-6000-085 | Perform basic scripting It shall be possible to perform basic scripting, either with a simple scripting language, C++ or Python, or using visual scripting (e.g. using blocks with certain functions and parameters). Remark: A script or recipe should be able to be saved and loaded, such that it can be used multiple times. | Click the drop-down box to select
Requirement R-6000-095 | Stitched imaging The system shall be able to make a large area, high resolution image. Remark: This may be achieved by performing stitched imaging, in which an area is defined and a resolution is set, after which the SEM will acquire the necessary micrographs and combine these in one large image. | Click the drop-down box to select
Requirement R-6000-105 | Simple analysis tools The SEM software shall be able to perform various analyses on the acquired image data, such as: • Measuring length of features • Counting the number of particles • Determine size and surface area of particles Remark: In the tender document it shall be stated which analysis tools are available. | Click the drop-down box to select
Requirement R-6000-110 | Save analysis data Analysis data shall be able to be saved in a file format, such as PDF, DOCX, or similar. Remark: In the tender document it shall be stated what filetypes are available. | Click the drop-down box to select
Requirement R-6000-115 | Customizable data bar The SEM software shall allow for customization of the micrograph data bar, including the ability to select which parameters are displayed and the order in which they are presented. | Click the drop-down box to select
Requirement R-6000-125 | SEM remote operation The SEM shall support remote operation. Remark: Protocol to be determined after awarding tender. Minimum requirements are Microsoft Windows 10 with up to date anti-malware software and firewall. | Click the drop-down box to select
Requirement R-6000-130 | SEM operation single operator The SEM shall be able to be safely operated by a single person. | Click the drop-down box to select
Requirement R-6000-135 | SEM human machine interface The SEM shall be delivered including a specialised input unit, that will control the most basic parameters of the SEM, such as magnification, zoom, astigmatism controls. | Click the drop-down box to select
Requirement R-6000-140 | EDX colour element mapping The SEM EDX system shall be capable of producing element colour maps, which display the distribution of elemental composition within a sample by assigning each element a unique colour. Remark: expected is that TNO can assign a colour to an element and can change this freely. If this is not possible, this shall be stated in the Tender document. | Click the drop-down box to select
Requirement R-6000-145 | SEM software display operating parameters The SEM software shall display the key operating parameters on a dedicated part of the GUI. Remark 1: The SEM supplier shall put forward in the tender the proposed key parameters that are shown and logged. Remark 2: Key operating parameters are to be discussed and agreed on after the tender is awarded. | Click the drop-down box to select
Requirement R-6000-155 | Measurement data file system information When a micrograph is saved in 16-bit grayscale TIFF format, the file shall contain EXIF metadata, such as date and time created, accelerator voltage, magnification, stage coordinates, etc. Remark: It is expected that other file formats, such as 8-bit grayscale TIFF and various JPG formats, will also contain EXIF metadata. | Click the drop-down box to select
Requirement R-6000-160 | SEM computer data transport The SEM software shall enable transport of the data stored on the SEM computer to other storage locations. Remark: Network shares are the expected destination of the data. | Click the drop-down box to select
Requirement R-6000-165 | Abort button The SEM software shall feature an abort button which stops immediately all motion, and will stop immediately the SEM measurement. After the abort button is activated the SEM will await new commands. Remark: This button shall not power the SEM down | Click the drop-down box to select
Requirement R-6000-175 | SEM control with gloves The control of the SEM shall function while operating it with non-conductive gloves. | Click the drop-down box to select
Requirement R-6000-180 | Software to be delivered The SEM shall be delivered with all the installation software required to (re-)install the control system of the SEM. Remark 1: This includes all the programs installed in the PC and the PLC (if applicable). Remark 2: The installation process shall be documented in a manual that is part of the data pack to be delivered. | Click the drop-down box to select
Requirement R-6000-185 | Control interface user friendly It shall be possible to perform basic operations with the system after three days of instructions. Remark 1: The supplier shall put forward in the tender document the number of days that are needed for initial instruction. Remark 2: The supplier shall put forward in the tender document a relevant example of the control interface screen, e.g. by using print-screen(s). | Click the drop-down box to select
Requirement R-0000-020 | Installation requirements The PDR documentation shall include all the documentation/information necessary for preparing the installation of the SEM. Remark: This includes for instance fitting diameters, electrical connections, floor load, required space for transportation, EMC fields, hoisting points etc. | Click the drop-down box to select
Requirement R-0000-025 | Specification of external resources At the PDR, or preferably in the tender document, the supplier shall specify the required external resources for the SEM system. Remark 1: TNO expects to be able to provide all required external resources. Remark 2: Information provided in the tender document shall at least include ball-park figures of: - electrical power requirements; - power consumption; - cooling requirements; - required mass flow rates of gasses (such as nitrogen and compressed dry air); - liquids. | Click the drop-down box to select
Requirement R-0000-030 | Output to VLL At the PDR, or preferably in the tender document, the supplier shall specify the SEM generated output to the cleanroom or lab. Remark 1: It is expected that the SEM system will produce heat, magnetic fields, etc., the amount and source of which shall be stated. Remark 2: TNO will inspect the supplied information and expects that that none of the output will be unacceptable for the environment. If an output to the environment is unacceptable, a method of mitigation shall be discussed between both parties. | Click the drop-down box to select
Requirement R-0000-035 | Documentation of third party hardware The supplier shall deliver certificates, CE declarations of conformity, and manuals for all the third party hardware that is part of the SEM system. Remark: Delivery of documentation is expected during FAT. | Click the drop-down box to select
Requirement R-0000-040 | CE certification The SEM shall be CE certified. | Click the drop-down box to select
Requirement R-0000-045 | SEM manual The SEM shall be delivered with a manual. This manual shall include the necessary drawings and electric layouts to perform simple troubleshooting and shall include basic operation instructions. | Click the drop-down box to select
Requirement R-0000-050 | Certificates The supplier shall document all the certificates. Remark 1: TNO expects certificates on: - Calibration certificates (if applicable); - Specifications and manuals of COTS equipment; - Outcome of the review of the used materials against the forbidden material list Remark 2: Delivery of documentation is expected during FAT. | Click the drop-down box to select
Requirement R-0000-055 | Drawings to be delivered The SEM shall be delivered with all relevant STEP files and "as built" drawings. Remarks 1: It is expected that it will include, relevant mechanical drawings (interfaces, sample holder), electric, gas, water, etc. Remarks 2: Delivery of documentation is expected during FAT. | Click the drop-down box to select
Requirement R-0000-060 | Used standards The SEM system design and implementations shall implement the relevant and applicable legislation and regulations in the following order of preference if not already specified in this requirements document. - International standards; - European standards; - Commercial standards; - National standards; - Manufacturing standards, newly developed standards | Click the drop-down box to select
Requirement R-0000-065 | Radioactivity The SEM system shall comply with all radioactive codes and regulations in order to be allowed to be installed and operated in the Netherlands. Remark: If any actions are needed on the facility side, like a permit, the supplier shall inform TNO that this is needed before PDR. | Click the drop-down box to select
Requirement R-0000-070 | FMEA The supplier shall perform an FMEA and shall share the outcome with TNO. Remark 1: It is expected that TNO is asked for input to the FMEA in an early stage of the process. Remark 2: TNO expects that identified critical failures are safe-guarded with at least two independent mitigation methods. | Click the drop-down box to select
Requirement R-0000-075 | SEM emergency stop button The SEM shall have at least one physical emergency stop button. Remark 1: This emergency stop will initiate the EMO state. Remark 2: The location of the emergency stop button shall be agreed on with TNO. | Click the drop-down box to select
Requirement R-0000-080 | Functionality during SAT The SEM shall be fully functional during SAT. Remark 1: The supplier shall provide a validation protocol for SAT at least two months before SAT. Remark 2: The supplier shall put forward what is expected from TNO during SAT at least two months before SAT. | Click the drop-down box to select
Requirement R-0000-085 | Transport to TNO The SEM supplier shall be responsible for the transport of the entire SEM system from their facility to the VLL building, according Incoterm Delivered Duty Paid (DDP) 2020. | Click the drop-down box to select
Requirement R-0000-090 | Transport inside VLL TNO will arrange and bear the costs of the transport of the SEM system from outside the VLL building to its final location inside the VLL. The supplier shall be present during this transport, and shall act actively, in case the supplier deems this transport to be unsafe for the SEM system. Remarks 1: Since at time of transport the SEM is not yet the property of TNO, both the supplier and TNO share the responsibility that this transport will not damage the SEM system. Remarks 2: The installation/assembly of the SEM system remains the responsibility of the supplier. Remarks 3: On request, TNO can supply information about companies that have permission and experience to work in the VLL building concerning installation work and transportation. | Click the drop-down box to select
Requirement R-0000-095 | SEM external volume claim during transport During transport the SEM shall fit in the volume claim of 1.8 by 1.8 by 2.0 meters. Remark: During transport the SEM may be delivered in several parts, as long as the available space in the clean room is sufficient for assembly. | Click the drop-down box to select
Requirement R-0000-100 | Windows update The control software shall support, free of charge, all future versions of windows during the expected lifetime of the SEM. Remark: the free of charge is as long as no hardware upgrade is needed. | Click the drop-down box to select
Requirement R-0000-105 | Control software updates The supplier shall update the software free of charge up to the minimum expected lifetime of the SEM. Remark: if a hardware update is needed for the software update, this option will be presented to TNO, and if TNO wishes to perform this hardware update, the no additional license fee remains in place. The hardware cost shall be marked conform prizes. | Click the drop-down box to select
Requirement R-0000-110 | License costs The offer includes all possible software licence fees, for all required applications, including all options and updates, without any limitation, for the minimum expected lifetime of the SEM. | Click the drop-down box to select
Requirement R-0000-115 | Advanced training Advanced application training for the SEM and selected accessories shall be given to TNO personnel on site, within 4 months after successful SAT. This training shall be given by an expert in the field of the supplier. Remark: it is expected that this training will be a 2- or 3-days training. | Click the drop-down box to select
Requirement R-0000-118 | Terms and conditions of payment • 20% at order, after receipt of a bank guarantee covering 80% of the total amount; • 20% after (Preliminary/Critical) Design Review (CDR); • 40% after delivery and visual inspection at location TNO Delft. • 20% after installation and commissioning On-site and Site Acceptance Test (SAT) including test periods, approved by TNO, Management Nano Instrumentation and Procurement. At this moment TNO will return the bank guarantee. Remark 1: The Tenderer to whom the Contract will be awarded must provide a bank guarantee to TNO, covering 80% of the total price. The Bank guarantee and the bank who issues the Bank guarantee have to be approved by TNO (minimal A rating). TNO must have the right to draw the bank guarantee in any and all cases where Tenderer falls short in fulfilling its obligations under the Contract. The validity of the bank guarantee may not expire until approval of the SAT by TNO Management and Procurement. Remark 2: please note, that approving the CDR this does not mean that TNO will take final responsibility if during SAT it is found that the system does not meet the minimum requirements. TNO will at all times hold Tenderer ultimately responsible for delivering an single-photon detector in accordance with the requirements as stated in this Tender. | Click the drop-down box to select

## validatie (hidden)
Click the drop-down box to select
Yes, compliant to requirement
No, not compliant to requirement
note: dit blad 'verbergen' indien je dit gaat delen intern/extern (behoudens aan inkoop collega's)