## Appendix A06
Appendix A06, Prescribed format for answering questions_notification of preferences Guidelines for the European Public Tender for a scanning electron microscope, ref. TNO 2023 FPL-INK 029, date 10-07-2023
The Tenderer demonstrates the ability to meet the preferences by means of his answers to the various questions/preferences, which must be structured in accordance with the instructions given in Chapter 6 and 8.
Preferences | Decription | Max. Score | Tenderer declares that he is able to meet the Preference | Reference in Tenderer's documentation where Tenderer indicates in which way he can fulfil the preference | Further clarification or remark
Preference R-1000-020 | Source lifetime desire The SEM should have a minimum lifetime of 10 years when operated 1000 hours per year. Remark 1: close out of requirement by proof from similar systems and use, or free repair when the source fails before stated requirement. Remark 2: the supplier shall state in the tender document the promised lifetime. Remark 3: lifetime is stated a time before full replacement of source, not tip maintenance or exchange. The supplier shall state the expected time before tip maintenance of tip exchange is needed. | 400 | Click the drop-down box to select
Preference R-1000-050 | Beam acceleration voltage on sample between 0.1 keV and 30 keV stepless desire The beam energy on the sample should be able to be controlled stepless, with a resolution of not bigger than 50 V steps, between 0.1 keV and 30 keV. Remark: the supplier shall state the minimum step size in the tender document. | 30 | Click the drop-down box to select
Preference R-1000-056 | Horizontal field width at low vacuum, at eucentric height desire The SEM should ideally have a large maximum HFW of 2 mm or more when imaging non-conducting samples, at eucentric height. | 75 | Click the drop-down box to select
Preference R-1000-065 | Imaging multiple channels simultaneously desire It is expected that the SEM has several detectors, the SEM should be able to image all detectors simultaneously. Remark 1: the supplier shall state (if not all detectors can be imaged simultaneously) the detectors that can be imaged simultaneously Remark 2: imaging simultaneously is not equal to showing all the images simultaneously, TNO accepts that not all images will be displayed at the same time. | 100 | Click the drop-down box to select
Preference R-1000-075 | Use of all detectors without manual hardware changes desire The SEM should be able to use all offered detectors without any manual hardware actions needed. Remark: If not possible for all the detectors, the supplier shall state what is possible. | 100 | Click the drop-down box to select
Preference R-1000-080 | Electron energy filtering option desire The SEM should be equipped with a BSE detector where the detected electron energy can be filtered. | 200 | Click the drop-down box to select
Preference R-1000-081A | Stage bias desire The SEM should be equipped with a negative stage bias with a tuneable range between at least -2000 V to 0 V. Remark: The performance of the stage bias will be stated in the tender document, if applicable. | 100 | Click the drop-down box to select
Preference R-1000-081B | Stage bias desire The SEM should be equipped with a positive stage bias with a tuneable range between at least 0 V to +250 V. Remark: The performance of the stage bias will be stated in the tender document, if applicable. | 100 | Click the drop-down box to select
Preference R-2000-020 | Level of integration desire The by the supplier selected EDX should have an as high as possible integration of the EDX control into the SEM control system. Remark 1: The supplier shall provide documentation (up to 10 slides) how the level of integration of the EDX is achieved in the overall SEM control system. Remark 2: the stated level of integration will be validated with demonstration. | 800 | Click the drop-down box to select
Preference R-3000-010 | Sample size loading desire The SEM should be able to load samples which are larger than stated in D003, which is explained in Appendix D004. | 500 | Click the drop-down box to select
Preference R-3000-020 | Sample measurement area desire The SEM should be able to measure a surface area which is larger than stated in Appendix D003 baseline. This is explained in Appendix D004. Remark 1: it is expected that an XYZ translation of the sample stage will be used to position the sample to measure at the desired measurement spot. It is allowed to also use a rotation to increase the measurement area. If this rotation is used, this shall be fully automated in the software of the SEM. Remark 2: if rotation is used, this shall be stated in the tender documentation. | 400 | Click the drop-down box to select
Preference R-3000-035 | Sample measurement at tilt desired The SEM should be able to tilt more than 60 degrees, as stated in Appendix D003. Remark: The supplier shall indicated how the design of the sample stage will prevent the samples from movement during tilting. | 50 | Click the drop-down box to select
Preference R-3000-050 | Sample weight desired The SEM should be perform all stated SEM requirements with a sample weight up to 2,5 kg, with the exemption of the tilt motion. Remark 1: this stated sample weight is excluding sample holders weight that are required for normal SEM usage, with the stated sample in Appendix D003, and if applicable Appendix D004, and therefor are in the delivery of the SEM. Remark 2: the supplier shall state the maximum sample weight | 30 | Click the drop-down box to select
Preference R-3000-060 | Sample weight including tilt motion desired The SEM should be perform all stated SEM requirements with a sample weight up to 2,5 kg, including the tilt motion. Remark: this stated sample weight is excluding sample holders weight that are required for normal SEM usage, with the stated sample in Appendix D003, and if applicable Appendix D004, and therefor are in the delivery of the SEM. | 30 | Click the drop-down box to select
Preference R-3000-066 | Limit of use of carbon tape desire The use of carbon tape for mounting the defined standard samples, see Appendix D003 and D004, should be kept as minimal as possible. | 100 | Click the drop-down box to select
Preference R-3000-075 | Mounting of sample time needed desired The SEM should enable the mounting of the sample on the sample stage within 2 minutes for all the samples indicated in Appendix D003 and if applicable Appendix D004. Remark: this is the time without pump down of the system. | 30 | Click the drop-down box to select
Preference R-3000-095 | Sample position repeatability desired With the stated sample motions of requirement R-3000-015, and the stated sample weight in requirement R-3000-045, the repeatability of the sample stage position should be better than or equal to 1 um. Remark 1: this requirement direct to the demanded measurement areas and sample weight. Unless indicated differently TNO expects that when a demanded requirement is upgraded to a desired requirement, the stated value of this requirement still applies. If this is not the case, the supplier shall make this clear in the tender document. Remark 2: repeatability is defined as; how well does the sample stage does return back to previous position. Remark 3: this requirement is without unloading the sample Remark 4: this requirement is without tilt of the sample. | 400 | Click the drop-down box to select
Preference R-3000-105 | Stage position accuracy desire With the stated sample motions of requirement R-3000-015, and the stated sample weight in requirement R-3000-045. The accuracy of the sample stage should be better than or equal to 1 um over a travel range of 100 mm in either X or Y direction. Remark 1: this requirement direct to the demanded measurement areas and sample weight. Unless indicated differently TNO expects that when a demanded requirement is upgraded to a desired requirement, the stated value of this requirement still applies. If this is not the case, the supplier shall make this clear in the tender document. Remark 2: accuracy is defined as; how well does the sample stage know its absolute position. Remark 3: this requirement is without unloading the sample Remark 4: this requirement is without tilt of the sample. | 400 | Click the drop-down box to select
Preference R-4000-010 | Contamination growth desired The SEM should meet the stated values, with the stated method in Appendix D006 with respect to contamination growth. | 250 | Click the drop-down box to select
Preference R-4000-030 | Materials in contact with sample desire Unless needed for SEM performance the material(s) that are in contact with the sample should be aluminium. Remark: If there are items in contact with the sample that are made from a different material, the supplier shall state this in the tender document. | 30 | Click the drop-down box to select
Preference R-4000-040 | Pump down time with sample desired The SEM should reach the required high vacuum level for SEM operation within 2 minutes after loading a sample as stated in Appendix D003 from atmosphere with a sample that has been exposed to atmospheric conditions for at least 24 hours. | 30 | Click the drop-down box to select
Preference R-4000-050 | Venting time desired The venting time, including all internal SEM actions that are required, should be within 2 minutes after giving the venting signal. Remark: Internal SEM actions could be stop and cool down of SEM source. | 30 | Click the drop-down box to select
Preference R-4000-055 | Switching time between vacuum modes desire The switching time between high vacuum operation mode and low vacuum operation mode, and vice versa should be within 5 minutes. | 30 | Click the drop-down box to select
Preference R-4000-060 | Switching time between vacuum modes manual actions desire The switching time between high vacuum and low vacuum mode, and vice versa should not require manual actions. | 30 | Click the drop-down box to select
Preference R-4000-080 | Turbo Molecular Pump (TMP) without oil, desire The SEM should use a TMP without oil. Remark: this may be achieved by using a magnetically levitated TMP. | 75 | Click the drop-down box to select
Preference R-6000-020 | Customizable GUI desire The SEM software GUI should be able to be customized and saved per user profile. | 150 | Click the drop-down box to select
Preference R-6000-045 | Use external images as navigation image desire The SEM software should be able to load an external image, to be used as navigation image (Requirement R-6000-025). Remark: expected formats are: .JPG, .PNG, .TIFF | 200 | Click the drop-down box to select
Preference R-6000-070 | Export and import stage coordinate list desire A list of saved coordinates should be able to be exported and imported, for example in a text format file. Remark: The coordinate list should at least consist of X, Y, Z coordinates. Preferable also rotation and tilt. | 200 | Click the drop-down box to select
Preference R-6000-080 | Custom coordinate system desire The SEM system should be able to set-up a custom coordinate system, which can be set-up based on one, two or three positions. Remark: see Appendix D008 for the detailed description on the expected implementation. | 500 | Click the drop-down box to select
Preference R-6000-090 | External control of SEM desire The SEM should be able to be controlled by an external programming script using a programming language, such as Python or C++. | 50 | Click the drop-down box to select
Preference -6000-100 | 3D image acquisition desire The system should be able to process micrographs made with a segmented detector, to acquire a 3D reconstruction. Remark: This may be done using external software, which will be supplied with the system. | 200 | Click the drop-down box to select
Preference R-6000-120 | Save and load data bar settings desire Users should have the ability to save and recall custom data bar configurations for future use. | 30 | Click the drop-down box to select
Preference R-6000-150 | SEM software system logging desire The SEM should ideally log all operations in a system logfile. | 50 | Click the drop-down box to select
Preference R-6000-170 | Imaging file name desire The SEM software should enable an easy to program file name, for saving micrograph images, with options like automatically adding relevant imaging information, e.g.: - Magnification - Accelerator voltage - Stage coordinates Remark: The Tender document shall provide the level of easy naming of files, including all options to add to the filename. | 300 | Click the drop-down box to select
Options | Decription |  | Tenderer declares that he is able to meet the option | Reference in Tenderer's documentation where Tenderer indicates in which way he can fulfil the option | Further clarification or remark
Option R-0000-120 | List available options The supplier shall list the available options including the cost of each item individually. |  | Click the drop-down box to select
Options R-0000-125 | Not affecting the performance of any other requirement The stated option(s) shall not affect the required performance / the compliance, with any other requirement stated in this document. |  | Click the drop-down box to select
Option R-0000-130 | Clear description of performance of the options The supplier shall put forward the relevant performance parameters of each option. If the option is selected, this stated performance will become a requirement that shall be met. |  | Click the drop-down box to select
Options R-0000-135 | Option for load lock The SEM supplier shall state if the SEM can be equipped with a load lock, and if so, which of the samples can be loaded from Appendix D003 and D004. |  | Click the drop-down box to select
Option R-0000-140 | Option dedicated sample holders Appendix D003 samples Requirements R-3000-065, R-3000-070 and Preference R-3000-075 aim for easy, and speedy loading of the standard samples, stated in Appendix D003. The supplier shall state all standard sample holder available and mark the sample holders that can be used to load the samples stated in Appendix D003, including prizes. Next to standard sample holders, and if needed, the supplier shall state the cost for dedicated sample holders to load the stated samples, stated in Appendix D003 Remark: TNO expect that wafer, reticle and a multipurpose sample holder is in the default sample holder catalogue are available. |  | Click the drop-down box to select
Option R-0000-145 | Option cold trap The SEM supplier shall state if a cold trap is of benefit for TNO’s anticipated usage of the SEM. |  | Click the drop-down box to select
Questions | Decription |  | Reference in the tenderer's documentation where the tenderer replies to the questions asked. |  | Further clarification or remark
Question 8.10.1 | Corporate Social Responsibility is becoming increasingly important in our own organization and in the future TNO will also ask more from our suppliers to play a part in this. What possibilities have you already initiated and where do you see opportunities for yourself and for TNO?
Question 8.10.2 | As you're aware, ISO 20400 is an international standard that provides guidelines on sustainable procurement. The standard can be used by any organization that intends to improve their social, economic, and environmental sustainability. Considering the growing awareness with regard to sustainable business practices, nationally and internationally, whether required by laws and regulations, promoted by the UN Sustainable Development Goals, or even by the changing societal expectations – there are a myriad of reasons in today’s world that drive organizations to be more sustainable, including TNO. Do you embrace the ISO 20400 guideline? - if so, how do you implement it in your own organisation? - If not, are you still planning to so, or are there specific reasons why your organisation is not embracing it.
Question 8.10.3 | Laboratory instruments, like the single-photon detector, that TNO uses for its scientific research often consume a lot of energy and also often consist of very high-quality raw materials. For the purpose of corporate social responsibility (CSR), do you, or will you in the future, make technological investments to reduce power consumption during instrument use, without negatively affecting the technical application of the instrument? The design of products with recycling in mind should be an important place to start for all of us. The inclusion of recycled or reused materials is a good way to increase demand for recycled and reused materials as it can significantly reduce the greenhouse gas emissions impact of the materials used in such laboratory instruments. Can you give TNO some insight on how and in which way you are using recycled or reuse materials now, or are working on solutions to use more recycled or reuse materials in the future, but still safeguard the basic use of such laboratory instruments?
Question 8.10.4 | The design of products with recycling in mind should be an important place to start for all of us. The inclusion of recycled or reused materials is a good way to increase demand for recycled and reused materials as it can significantly reduce the greenhouse gas emissions impact of the materials used in such laboratory instruments. Can you give TNO some insight on how and in which way you are using recycled or reuse materials now, or are working on solutions to use more recycled or reuse materials in the future, but still safeguard the basic use of such laboratory instruments?
Questioin 8.10.5 | TNO focuses within its organization on reducing energy, water use and emissions. What greening initiatives have you taken or can you take in your product portfolio to reduce energy and water consumption and emissions?
Question 8.10.6 | In the case of transport packaging, one important opportunity is that of substituting reusable packaging for single trip packaging. Can you give TNO some insight on how and in which way you are using recycled or reuse materials now, or are working on solutions to use more recycled or reuse materials in the future?
Question 8.10.7 | In the case of shipping, from the sold products to the customer, we all have an important possibility to reduce emissions. Can you give TNO some insight into how you are trying to reduce the greenhouse gas impact in this perspective?
Request | Decription |  | Reference in the tenderer's documentation where the tenderer replies to the request asked. |  | Further clarification or remark
Request 8.10.8 | TNO expects suppliers to take responsibility in sustainable product development and product use. Therefor we expect suppliers to offer a service for buying back equipment at the end of lifetime or at the end of service at TNO. Please elaborate on the service you will offer to TNO, and make your offer explicit in terms of actions and impact.

## validatie (hidden)
Click the drop-down box to select
Yes, Tenderer is able to meet this Preference
No, Tenderer is not able to meet this Preference
Click the drop-down box to select
Yes, Tenderer is able to meet this option
No, Tenderer is not able to meet this option
note: dit blad 'verbergen' indien je dit gaat delen intern/extern (behoudens aan inkoop collega's)