## Program of Requirements
1. GENERAL
| The Tenderer its offer must unconditionally meet the requirements set by the LUMC for the performance of the contract. This means that all requirements must be answered with “yes”. When the LUMC concludes that the Tender does not comply with the requirements, the offer will be rejected and the Tenderer will be excluded from any further participation (knock-out). | To be filled in by Tenderer
|  | YES | NO | Comments/Proof
Program of Requirements
1 | The microscope is state-of-the-art and highly suitable for automated transmission electron microscopy.
2 | The microscope can operate at 80 and 120 kV.
3 | The microscope has an LaB6 electron source. In total 6 electron sources and Wehnelt cups (or similar) are included in the offer.
4 | The complete system is safe, i.e. not harmful for users, safety points include but are not limited to X-ray, electronic and chemical safety.
5 | The offer includes a water-water chiller.
6 | The stage should have excellent movement (fast), drift (low), accuracy (high) and stability (high) characteristics. The relocation accuracy is better than 0.5 micron.
7 | System has a remote control to operate the microscope, i.e. system can be controlled from on a separate desk with hand panels, computer screens and controls. This desk can be positioned freely in the microscope room with respect to the microscope, cabling is included which allows this setup. For this remote control setup 4 identical high-quality screens of at least WUXGA 1920x1200 resolution are included.
8 | The microscope is equipped with a bottom mounted high-end electron detection camera, with square frame dimensions of at least 4k x 4k pixels and a frame rate of at least 40 fps at full resolution and a pixel size of at least 14 micron. The camera should allow to work from a remote computer without additional lag. The camera has integrated live drift correction for frames. Specify brand, type etc.
To acquire a fullframe, unbinned and drift corrected image with an expsosure time of 1.0 second takes less than 2.5 seconds. The time is measured from the moment the acquisition command is given through the API (requirement 9) to the moment the image is in computer memory and available for further processing in python.
Characteristics of the camera (stability, read-out, dead pixels, distortion, etc.) are of highest quality. Please specify characteristics and what efforts will be made to ensure the highest quality.
9 | Generic API’s are included to control camera, e.g. Thermo Fisher Advanced Temscripting or Gatan GMS3 with python plugin or similar. API’s are compatible with Python3 and are preferably COM interfaces. API’s offered by other parties than camera vendor are not accepted, e.g. serial EM. Deployment and distribution on a non-commercial basis of client developed software using supplied API’s is permitted without restriction. The API’s should unlock the full potential of the camera.
10 | COM interfaces are included for microscope control (excluding camera control), e.g. Thermo Scientific Advanced TemScripting , Jeol TemExt or similar. All user controllable parts of the microscope should be accessible e.g. but not limited to stage, apertures, lenses, deflectors, vacuum system, electron source and high tension settings. API’s offered by other parties than microscope vendor are not accepted, e.g. serial EM. Documentation is to be submitted at tender submission. Contact details of a technical expert are provided before microscope acceptance. This engineer should be available for support and assistance on software related questions. Deployment and distribution on a non-commercial basis of client developed software using supplied API’s is permitted without restriction. The API’s should unlock the full potential of the microscope.
11 | Pre-installation manuals should be made available for all equipment at tender submission.
12 | A small anti-contamination Dewar, max 5 liter. The Dewar should not be located inside an enclosure and should be readily available for a refilling system (requirement 13). Documentation illustrating the microscope Dewar location and construction is to be included in tender documents.
13 | A liquid nitrogen refilling system that is able to fill the small anti contamination Dewar on the microscope (requirement 12). The solution should have similar functionality as the “Simple Origin cryogen refilling system Next-gen” (https://simpleorigin.us/) system or can be based on a Norhof dispensing system (https://www.norhof.com/product/900-automatic-advance-ln2-microdosing-pump/) . The filling procedure should be both safe for environment, user as well as for the microscope system. An API is included to allow client software control of the filling. Documentation illustrating the microscope Dewar location and construction is to be included in tender documents.
14 | In case the anti-contamination Dewar on the microscope is fixed to the microscope, it is required to have a anti contamination/de-icing cycle, i.e. by means of a heating system. Also, then the decontamination cycle should be automatic and require no user intervention beyond starting the procedure in the software.
15 | The microscope vacuum system has a decontamination / cleaning cycle.
16 | The system has software controllable deflectors, lenses, vacuum valves, stage, electron source, apertures (condenser, objective and SA apertures) and camera. Also, all present apertures should at least be possible to move in and out of the beam. All this functionality should be available from the offered software API, see requirement 10.
17 | System has an oil free vacuum system. System does not contain any ODP.
18 | Client can install third party software packages on user accessible computers including e.g. python3, Matlab, up-to-date web browser, data storage, debugging software, login scripts and third party packages. This to allow integration of the system in the local environment, scientific developments and use of third party packages. This all without affecting warranty or maintenance and support. Conditions will be defined and agreed upon with mutual consent.
19 | Standalone simulators for the camera and microscope API’s are made available to develop and test in-house developed software, including relevant licenses. Please describe simulator, e.g. but not limited to OS version, requirements, camera support, specify limitations if any as well. On request, updates are made available without additional cost when software versions deviate significantly from the microscope versions, i.e. when functionality in the development system does not match (anymore) with that on the microscope computers.
20 | All accessory third party software like e.g. but not limited to virus scanners, backup software, cloning software, disk mapping software is included with proper licenses and offline activation.
21 | For external communication the following dedicated network cards/ports are present which are exclusively to be used for communication with the clients ICT environment, i.e. not for internal system communications: Computers with detectors and or significant data storage: at least 1 x 10Gbit/s network port for data transfer. All computers: at least 1x 10/100/1000 Mbit/s network port, for communication with inhouse services and remote service (This can mean that for e.g. a camera controlling computer two interfaces should be available to connect to the clients network).
22 | The internal microscope network layout is specified by the manufacturer at least a month before delivery. A concise document describing the network diagram is made available electronically for all components with an exposed network interface including all IP addresses used. The external network will be setup by the client to not conflict with the given information.
23 | All upgrades to hardware and software are included and will be installed to bring the versions of the operating systems to a version which has at least 5 years full support by the operating system supplier this from the date of system acceptance. For windows a pro version is required.
24 | The LUMC requires an warranty of 2 years. This includes service as offered in the service, maintenance and support contract.
25 | Image-beam shift allows imaging of a larger area without moving the stage. Specify the maximum imaging area at the specimen level for the magnification range between 5k and 15k times. A minimum of 3x3 non overlapping images should be possible at 5k magnification.
26 | On site CWAT is part of the final acceptance.
27 | Supplier has option for site inspection before tender submission. The inspection is limited to a day. Supplier shares findings with LUMC.
28 | Supplier is responsible for transport to installation location and installation of complete system including unforeseen additional costs.
29 | All electronic equipment not fitted in an enclosure should be fitted in a computer rack, if needed a 42 U computer rack should be included in the offer. In general wires, tubes and cables should be tidy and protected by covers outside enclosures.
30 | Water, pneumatic and nitrogen distribution manifolds/systems should be housed in an enclosure.
31 | Specify which parts are not covered by the service and maintenance contract, e.g. filaments and other consumables.
32 | A basic user training for a maximum of 6 participants and an advanced training for 3 persons.
33 | All needed to connect to the LUMC infrastructure is included in the offer like e.g. but not limited to: tubing, connectors, and cables.
34 | A Maintenance offer for 15 years is included, each part of the offered system is specified, i.e. microscope, camera, chiller,….
35 | Service and maintenance requires a yearly written report, covering the:
- Status of the apparatus, i.e. describing the condition of the system as a whole
- 4-year forecast of expected maintenance
- Expected remaining life-time of whole system and individual parts of the system
- General conclusions and advice to keep system safe, operable and up-to-date
36 | Each service and maintenance work requires a service report with e.g. but not limited to:
- Service report, job number according to clients specification
- General description of the work
- Replaced parts
- Changes made to hardware
- Changes made to software
- Describes follow up if needed
- Transfer to and acceptance by client of completed work. This includes a functionality check on functioning apparatus before handover to client
37 | Software updates and changes require a system backup to be in place before the action is performed. A new backup will be generated after such action.
38 | There will be no compensation for service after premature decommissioning of the system by client.