## Vragenformulier
Document | Paragraph/Rq.Number | Question
Programma van eisen | 3/1 | How is precisely the sensitivity defined here? How precisely was the desired improvement (4-8x) calculated? Would a technology that exceeds this improvement level be evaluated positively on the tender?
By sensitivity here we mean achievable SNR at a given light dose (which in turn depends on both light-collection efficiency and noise levels). Our imaging applications require maximizing use of limited photon budget while also minimizing phototoxicity. For example in organoids, we need to image at 63x and 100x magnification at the highest possible sensitivity, speed, resolution, and imaging depth (sample thickness). We need the capability to fluorescently image cell nuclei, complete membranes, and other expressed proteins for organoids that are 200-400 micron thick. We also have other applications where resolution is most important, and here we need to image at 120 nm resolution while maintaining high sensitivity. Sensitivity of existing GaAsP detectors fall short of our requirements by 4-8x fold (i.e. we need to average 4-8 images to achieve desired SNR at permissible light doses). Any improvements beyond 4-8x sensitivity will of course be evalulated positively.
Programma van eisen | 3/3 | In our opinion there are several points to consider for quantitative imaging, such as: 1) linear dwell time; 2) laser power stability; 3) detector range; 4) detector linearity; 5) detector photon counting ability; 6) detector stability over time. Why is there an emphasis only on linear dwell time ? If a vendor can offer a technical solution exceeding point (1), would this be evaluated positively?
Linear dwell time (1) is important for quantitative intensity imaging. We already have a resonant scanning microscope (with non-linear dwell time) and we wish to complement with a linear dwell-time system. We appreciate points (2)-(6) are also important, but improvements there would not substitute for a linear dwell-time system.
Programma van eisen | 3/4 | Minimum 4x GaAsP detectors are specified.We assume that this does not exclude latest detector technologies exceeding GaAsP performance. Would such latest technology detectors be evaluated positively in the tender? Is fast camera-based overview image acquisition with subsequent navigation within the overview image a mandatory requirement? Are alternative solutions offering similar functionality accepted? Is the water immersion a mandatory requirement? For biological samples (especially organoids) there are objectives with alternative immersion media available that exceed the performance of water immersion objectives in terms of refractive index matching, resolution and light efficiency.
Regarding 4xGaAsP detectors, we of course would not exclude other technologies that provide superior performance but having high sensitivity across 4x channels is important for our multicolor imaging needs. Regarding fast camera-based overview image acquisition with subsequent navigation, we are in principle open to other technologies, but speed and large field of view are crucial for our applications. Regarding immersion media, those other than water could in principle be usable, but we need it to function with an autoimmersion module (see system requirement No. 6), and for our FCS applications, our samples will be in acqueous media and hence we favor water immersion, which (because of refractive index matching) provides the highest accuracy of absolute quantifications of physical parameters (of concentration, diffusion coefficient, etc.).
Programma van eisen | 3/6 | Can you please specify the term "large range"?
We mean by "large ranges in X-Y" the full scale of multi-well plates.
Programma van eisen | 3/11 | Here 48 fps are specified. Can you please specify at what scan field resolution/scanning conditions?
We require here 512x512 scanning at 2x Nyquist resolution.
Programma van eisen | 3/12 | What is the desired application for the 5 megapixel camera? Is the 5 megapixel camera a mandatory requirement?
The desired application is overview imaging prior to subsequent high-resolution confocal scanning. We specify 5 megapixel to to ensure at least Nyquist sampling at low (10x) magnification over a large field of view.
Programma van eisen | 3/15 | There are five different fps values specified in the document: >25 fps at point 3/1; >20 at point 3/4; 25 or 48 fps at point 3/11 and 13 fps here (point 3/15). Can you please specify minimum scan speed and scan conditions?
For single-wavelength super-resolution functionality, we require minimally 25 fps at 512x512 scanning, but for some applications we require faster imaging (at ~48fps) at the expense of some resolution. For multicolor (standard, non-SR) confocal we require a minimum of 13 fps.
Programma van eisen | 3/19 | Is multi-immersion objective a mandatory requirement? Is it acceptable to quote dedicated objectives with different immersion media with similar or better performance?
We will consider replacing multi-immersion objective with multiple separate objectives, but we prefer multi-immersion objective to enable fitting a larger repertoire of objectives within the objective turret.
Programma van eisen | 3/22 | Can you please explain this point (3/22) with an application example and/or experiment workflow?
For imaging transport processes in fungal networks (which span centimeters), we wish to survey at low resolution over a large area spanning multiple fields of view, analyze the stitched image overview to identify regions of interest, and automatically switch to high-magnification super resolution imaging. And repeat the process over many hours. We have similar applications also in organoid work.
Information Notice
Market Consultation Confocal Microscope

## Legenda
Beschrijvend document
Concept overeenkomst
Kerncompetentieformulier
Programma van eisen
Social return
Social return
UEA
Verklaring aansprakelijheid concern
Verklaring conformiteit
Verklaring Derde