## 1. Front page
Appendix A04: Schedule of Prices and Costs
Belongs to the Public Procurement Guidelines for the European public procurement for the delivery of a Quantum Scanning Probe Microscope (QSPM)
2024 FPL-INK 006
Date:
Tenderer company name:
Name:
Signature:

## 2. Instructions
Instructions | sittingduck
Secur | The white boxes are secur and Tenderer can not change these fields.
Entry box | The light bleu boxes within Tab "price page" Tenderer is requested to enter, and are part of Price (TP) (also the fields on the front page: Date, Tenderer Company Name and Name)
Entry box | The Yellow boxes within Tab "price page" Tenderer is requested to enter, and are NOT part of Price (TP)
Price (TP) | The Light Green box is the price (TP) offered by the Tenderer

## 3. Price page
Appendix A04, Schedule of Prices and Costs
Guidelines for the European Public procurement tender for the delivery of a Quantum Scanning Probe Microscope (QSPM) Ref. Nr. 2024 FPL-INK 006, date 08-02-2024
Items quoted for price (TP), Tenderer is permitted to add any additional necessary items and price. Priced items quoted in this part of the price sheet will be part of the tender price and purchased by TNO.
Line Number | Reference ID | Description | Price each | Quantity | Total amount
1 | 6.1.1 | Price - "Quantum Scanning Probe Microscope (QSPM)” This includes: 1 - two Year Warrenty 1 - (critical) design review 1 - Site Acceptance Test (SAT) 1 - Installation on site 1 - Operator & basic maintenance training 1 - Delivery according Incoterms "Delivered Duty Paid" (DDP), 2020 (location TNO Delft, the Netherlands | 0 | 1 | 0
2 | 6.1.1 | Price - "Maintenance/services" A one-year basic maintenance contract multiplied by five (5) years. Providing insight into at least the following elements: a) All activities covered by the contract; b) All activities which are not part of the contract; c) Giving insight into response times (both inside and outside business hours), where we distinguish the following service levels; o Service level 1: A complete loss of service in which the single-photon detector is down. A major issue prevent access to the system or the use of significant portions of the system; o Service level 2: Users are unable to perform a non-mission critical function, such as change a portion of the system’s configuration; o Service level 3: The single-photon detector issue is cosmetic in nature and/or displays minor flaws which are easily circumvented with little or no impact on the customer’s normal business operations. d) Service-desk reachability (both inside and outside business hours), either by telephone or via an online portal; e) Number of planned maintenance visits; Replacement parts prices or discounted prices on Tenderer enlisted selling price. | 0 | 5 | 0
Price (TP) |  |  |  |  | 0
Price quotation preference items. Tenderer is requested to quote line numbers 3 till 38 IMPORTANT: if the preference doesn't lead to an additional cost for TNO, please enter '0' (zero)
Line Number | Preference ID | Description | Price each | Quantity | Total amount
3 | 8.2.1 | NV magnetometry Demonstrated NV magnetometry with increased sensitivity better than 2µT/√Hz. | 0 | 1 | 0
4 | 8.2.2 | High speed magnetic field measurement Demonstrated high speed magnetic field measurement with a scan speed of more than 100 pixels/s while having a sensitivity (or equivalent minimum detectable field at such scan speed) better than 5µT/√Hz. | 0 | 1 | 0
5 | 8.2.3 | Static magnetic field sensitivity Demonstrated increase in the static magnetic field sensitivity down to 100nT/√Hz using advanced techniques (such as gradiometry) with integrated software interface. | 0 | 1 | 0
6 | 8.2.4 | AFM probe height control AFM probe height control including scanning at a constant height as well as scanning at an offset height from a previously measured line and a measured two-dimensional topography map. | 0 | 1 | 0
7 | 8.2.5 | Sample scanner noise Sample scanner noise better than 0.1nm RMS for XY and better than 50pm RMS for Z, for a bandwidth of at least 100Hz in XY and 1kHz in Z. | 0 | 1 | 0
8 | 8.2.6 | Stitch coarse movements Software implemented possibility to stitch the coarse movements to achieve larger scan areas. | 0 | 1 | 0
9 | 8.2.7 | Sample size Possibility to mount samples with larger sizes: XY size of 50mm (or above if possible) and Z size up to 30mm. | 0 | 1 | 0
10 | 8.2.8 | Sample weight Increased sample weight capacity of up to 250g. | 0 | 1 | 0
11 | 8.2.9 | Sample mount Offering ready-made or custom-made solutions (e.g. PCB plates), for mounting and accessing electrical connections to the sample. | 0 | 1 | 0
12 | 8.2.10 | Regulated environment Temperature regulated environment with temperature control of ±0.1 °C at a target temperature of 20 °C and a lab temperature range of 18 to 24 °C. | 0 | 1 | 0
13 | 8.2.11 | Vibration isolation stage Active vibration isolation stage with active isolation bandwidth of at least 1-100Hz or better and with a transmissibility of >20dB at 5Hz and >=40dB above 10Hz. | 0 | 1 | 0
14 | 8.2.12 | Tip-sample drift rate Tip-sample drift rate of less than 2nm/h. | 0 | 1 | 0
15 | 8.2.13 | Microwave source Microwave source with increased frequency range of up to 6GHz. | 0 | 1 | 0
16 | 8.2.14 | Green excitation laser The green excitation laser with a tuneable power between 0 to >10mW at the focal point, with rise/fall time of <2ns, and modulation/pulsing bandwidth of at least 200MHz. | 0 | 1 | 0
17 | 8.2.15 | Confocal scanning and signal optimization Confocal scanning over the probe as well as automatic optimization of the NV signal, by for example (re-)tracking the NV fluorescence signal. Possibility to re-track the signal during the measurement if needed. | 0 | 1 | 0
18 | 8.2.16 | Stable optics Stable optics resulting in low drop in the detected NV fluorescence signal over several hours of measurement when no signal re-tracking is used. | 0 | 1 | 0
19 | 8.2.17 | Automated sensitivity optimization routine Automated magnetic sensitivity optimization routine to set the best laser and microwave power. | 0 | 1 | 0
20 | 8.2.18 | Access port Access port on the (back-side of the) microscope enclosure to feed in and out extra coaxial cables (e.g. BNC, SMA) and optical fibers, to and from outside. | 0 | 1 | 0
21 | 8.2.19 | Microwave antenna positioning stage Microwave antenna positioning stage with positioning resolution of <1µm and positioning repeatability of better than ±2µm. | 0 | 1 | 0
22 | 8.2.20 | Extended pulsed measurement modalities Extended pulsed measurement modalities with integrated GUI to control and display the data at least for the following modes: Pulsed ODMR, Rabi, Relaxometry, Ramsey, Spin-Echo, CPMG. | 0 | 1 | 0
23 | 8.2.21 | Custom-defined pulse-sequences Possibility to apply custom-defined pulse-sequences and display the results. | 0 | 1 | 0
24 | 8.2.22 | Vector electromagnet Calibrated, low noise, and temperature stabilized vector electromagnet with strength tunability in arbitrary direction up to at least 50mT, including software interface to control and calibrate the strength and orientation of the field. Resolution and repeatability is expected to be better than 0.1mT and 0.1 degree. The magnetic noise due to the electromagnet in steady-state operation is expected to be below the noise floor of the rest of the system. | 0 | 1 | 0
25 | 8.2.23 | Measurement data Record, save and export the measurement data in a user-accessible data format. This should include (all) the data that are displayed in the software user interface, including but not limited to: NV measurement results data, AFM-related data (topography, amplitude, phase), raw as well as processed fluorescence data and photon counts per pixel, hardware status including the laser and microwave signals, etc. | 0 | 1 | 0
26 | 8.2.24 | Hardware access to APD signals Hardware access to the electronic signals from the APD. | 0 | 1 | 0
27 | 8.2.25 | Hardware access to electronic signals Hardware access to electronic signals or triggers (e.g. TTL) to monitor the state of the hardware, including at least AFM (pixel, line, frame), laser/AOM modulation, and MW sequence triggers. | 0 | 1 | 0
28 | 8.2.26 | Hardware access to MW source Hardware access to the analog I/O outputs used for the MW modulation. | 0 | 1 | 0
29 | 8.2.27 | Hardware access to clock signal Hardware access to a 10MHz clock signal. | 0 | 1 | 0
30 | 8.2.28 | Polarization control of laser Possibility to control the polarization of the laser. | 0 | 1 | 0
31 | 8.2.29 | Custom beam path Customized beam path by adding two extra beam lines with wavelengths of 594(±1)nm and 637(±1)nm that are confocal with the green laser. This customization includes the necessary hardware (lasers, optics, electronic interfaces) as well as software interface to control the beams (individually and in combination) from the user interface. The fluorescence loss from the green laser due to the additional optics should be less than 5%. | 0 | 1 | 0
32 | 8.2.30 | Magneto-optic Kerr effect Integrated magneto-optic Kerr effect (MOKE) microscopy for increased speed in finding the regions of interest. | 0 | 1 | 0
33 | 8.2.31 | Scripting Possibility to control the hardware via scripting language, e.g. Python, etc. | 0 | 1 | 0
34 | 8.2.32 | Accessibility software Access to the future upgrades in the software and the user interface, free of charge, for at least the during of the expected lifetime of the instrument. | 0 | 1 | 0
35 | 8.2.33 | User interface software The user interface software shall support, free of charge, all future versions of the operating system of the instrument's computer (e.g. MS Windows) during the expected lifetime of the instrument. | 0 | 1 | 0
36 | 8.2.34 | Compatible NV probes Availability of compatible NV probes up to the minimum expected lifetime of the system. | 0 | 1 | 0
37 | 8.2.35 | Compatible probes with integrated antenna Possibility of using the instrument with compatible probes including integrated MW antenna. | 0 | 1 | 0
38 | 8.2.36 | Inverted optical configuration Compatibility with inverted optical readout configuration for transparent samples used e.g. in biological or NMR applications. | 0 | 1 | 0
Price (TP) |  |  |  |  | 0
Price quotation optional items. Tenderer is requested to quote line numbers 39 till 51 Tenderer is permitted to quote any other opptional needed items which are not mentioned by TNO, line numbers 52 till 60.
Line Number | Optional ID | Description | Price each | Quantity | Total amount
39 | 1.4.1 | Basic maintenance/services - 1 Year | 0 | 1 | 0
40 | 1.4.1 | Basic maintenance/services - 2 Years | 0 | 1 | 0
41 | 1.4.1 | Basic maintenance/services - 3 Years | 0 | 1 | 0
42 | 1.4.1 | Basic maintenance/services - 4 Years | 0 | 1 | 0
43 | 1.4.1 | Basic maintenance/services - 5 Years | 0 | 1 | 0
44 | 1.4.1 | Variation maintenance/services contract (as an alternative to basic maintenance contract) - 1 Year | 0 | 1 | 0
45 | 1.4.1 | Variation maintenance/services contract (as an alternative to basic maintenance contract) - 2 Years | 0 | 1 | 0
46 | 1.4.1 | Variation maintenance/services contract (as an alternative to basic maintenance contract) - 3 Years | 0 | 1 | 0
47 | 1.4.1 | Variation maintenance/services contract (as an alternative to basic maintenance contract) - 4 Years | 0 | 1 | 0
48 | 1.4.1 | Variation maintenance/services contract (as an alternative to basic maintenance contract) - 5 Years | 0 | 1 | 0
49 | 1.4.2 | Option for TNO, to buy a second quantum scanning probe microscope with the same requirements/preferences, within a period of one (1) year | 0 | 1 | 0
50 | 1.4.2 | Option for TNO, to buy a second quantum scanning probe microscope with the same requirements/preferences, within a period of two (2) years | 0 | 1 | 0
51 | 1.4.2 | Option for TNO, to buy a second quantum scanning probe microscope with the same requirements/preferences, within a period of three (3) years | 0 | 1 | 0
52 |  | ** optional ** | 0 |  | 0
53 |  | ** optional ** | 0 |  | 0
54 |  | ** optional ** | 0 |  | 0
55 |  | ** optional ** | 0 |  | 0
56 |  | ** optional ** | 0 |  | 0
57 |  | ** optional ** | 0 |  | 0
58 |  | ** optional ** | 0 |  | 0
59 |  | ** optional ** | 0 |  | 0
60 |  | ** optional ** | 0 |  | 0