## Appendix A05
Appendix A04, Conformity list 'Programme of Requirements and Preferences' Guidelines for the European Public Tender for a Quantum Scanning Probe Microscope (QSPM), ref. TNO 2024 FPL-INK 006, dated 08-02-2024
By the following Programme of Requirements TNO will assess if Tenderer meets the relevant requirements as laid down in Section 6 and 8. The Tenderer must indicate for each item, the conformity for the requirements in the table below.
If to any of the requirements, as stated in section 8, are not met by Tenderer, he will be excluded for any further assessment. By answering a question with “comply” will mean that in the event of awarding the contract, non-compliance to the requirement shall count as failure to meet an obligation from the Agreement to be concluded. If Tenderer does not comply with any of the preferences, as mentioned in point 8, he will not be excluded from further assessment. Answering a preference with "comply with" means that, at the time of award, failure to comply with the preference counts as failure to fulfil an obligation under the Agreement to be concluded. The abbreviations of the validation method during SAT are described in section 8.
Requirements
Requirement | Decription | Method of Validation | Does Tenderer comply to the requirement? | Reference in Tenderer's documentation where compliances with the requirement is described | Remark
8.1.1 | AFM and diamand probe An atomic force microscope that employs a diamond probe containing a single nitrogen-vacancy (NV) centre in the tip. The diamond tip is attached to a tuning fork that works in shear mode. Rationale: Basic functionality of the system. | SAT, DV | Click the drop-down box to select
8.1.2 | Integrated confocal microscope An integrated confocal microscope for excitation of the NV centre in the probe and detection of the emitted fluorescence photons including an objective with an NA >= 0.7, magnification >=50x, working distance >=5mm. Rationale: Optics for addressing and reading the NV centre. | SAT, DV | Click the drop-down box to select
8.1.3 | Closed-loop sample scanner Closed-loop sample scanner with scanning range of at least 80x80x15 µm3. Rationale: Performing a scan of the sample with respect to the NV centre. | SAT, DT | Click the drop-down box to select
8.1.4 | Coarse sample positioning stage Coarse sample positioning stage with a range of at least 5x5x15 mm3 and a resolution of at least 1µm. Rationale: Moving the sample to the position of interest. | SAT, DT | Click the drop-down box to select
8.1.5 | Sample size capacity Sample size capacity of more than 20x20x15 mm3 and sample weight capacity of at least 50g. Rationale: Minimum requirements for the sample size and weight. | SAT, DT | Click the drop-down box to select
8.1.6 | Sample mount Possibility to mount and access samples with electrical connections. Rationale: The instrument should provide means to mount and address samples that require electrical connections. | SAT, DT | Click the drop-down box to select
8.1.7 | Green excitation laser A green excitation laser with the nominal wavelength in the range of 515nm to 532nm with an output power of at least 5mW at the focal point. Rationale: Laser to excite the NV centres. | SAT, DT | Click the drop-down box to select
8.1.8 | Microwave generator electronics A microwave generator electronics with an output power of >= 30dbm and frequency range up to at least 3.5 GHz. Rationale: Microwave source to address the NV spin. | SAT, DT | Click the drop-down box to select
8.1.9 | Microwave transmission circuitry A robust microwave transmission circuitry to address NV spins. Rationale: To guide the microwave radiation from the source to the position of the NV centre | SAT, DT | Click the drop-down box to select
8.1.10 | Single-photon detector A single-photon detector with high quantum efficiency in the emission band of the NV centres Rationale: To detect the fluorescence photons emitted from the NV centre. | SAT, DT | Click the drop-down box to select
8.1.11 | Magnetic bias Magnetic bias with adjustable magnetic field strength at the position of the NV with a tunability range of >=5mT. Rationale: Provide bias magnetic field for the initial split of the NV spin states. | SAT, DT | Click the drop-down box to select
8.1.12 | Sample scanning Simultaneous and real-time AFM topography and magnetic field maps during the sample scanning. Rationale: AFM topography and magnetic field maps are the basic required functionality of the instrument and should be visualized while the measurement is running. | SAT, VT | Click the drop-down box to select
8.1.13 | Wide-field imaging capability Top-view wide-field imaging capability for inspection and alignment of the sample, probe and antenna. Rationale: To provide visuals and assist with alignment. | SAT, VT | Click the drop-down box to select
8.1.14 | Wide-field imaging capability Side-view wide-field imaging capability to inspect the relative position of the probe and the MW antenna and the sample's surface. Rationale: To provide visuals and assist with alignment. | SAT, VT | Click the drop-down box to select
8.1.15 | NV magnetometry sensitivity Scanning NV magnetometry with a DC, CW-ODMR sensitivity better than 3µT/√Hz Rationale: The minimum required sensitivity.z | SAT, DT | Click the drop-down box to select
8.1.16 | NV imaging modes NV imaging modes including at least: - Photoluminescence or quenching mode - Iso-B or magnetic field contour mode - full CW-ODMR mode Rationale: Minimum NV imaging capabilities of the instrument. | SAT, VT | Click the drop-down box to select
8.1.17 | Software package Software package with graphical user interface to control the hardware’s and perform the measurements, and record the results. Rationale: Interface to communicate with and control the instrument. | SAT, VT | Click the drop-down box to select
8.1.18 | User manual User manual with detailed description of the hardware and software, including any custom made options. Rationale: To provide clear understanding of the system's functionalities, facilitate effective utilization and help with the troubleshooting of the instrument. | SAT, V | Click the drop-down box to select
8.1.19 | Installation On-site system installation. Remark: The installation documentation shall be provided at least two months before the agreed installation date. The document shall include all the information necessary for preparing the installation, for instance fitting diameters, electrical connections, floor load, required space for transportation, EMC fields, hoisting points. etc. Rationale: To ensure efficient setup, configuration, validation and immediate troubleshooting of the instrument. | SAT, VT | Click the drop-down box to select
8.1.20 | On-site basic training On-site training for the instrument and selected modalities for three users with the duration of at least two days within maximum one week after installation and testing of the system. Rationale: Increasing the efficient and successful operation of the instrument. | SAT, VT | Click the drop-down box to select
8.1.21 | Meet all requirements The supplier shall provide a system that simultaneously meets all defined minimum requirements and all accepted preferences as listed in this tender, unless otherwise specified in a specific requirement. Rationale: Required to meet main function of the instrument. | SAT, DVT | Click the drop-down box to select
8.1.22 | Operation lifetime The instrument shall have an operation lifetime of at least 10 years. Remark 1: Tenderer must describe in what way it guarantees a operational lifetime of at least 10 years. Remark 2: Close out shall be done through review of design, proof of heritage, used materials and equipment. Rationale: Minimum expected lifetime of such a system. | FAT | Click the drop-down box to select
8.1.23 | Grouped cables The cables shall be grouped together as much as possible. Rationale: Grouped cables decrease the risk of damage to a single cable and make routing easier. | SAT, VT | Click the drop-down box to select
8.1.24 | Labelled connectors and cables All accessible connectors and cables to and from the instrument as well as the peripheral devices and units (e.g. controllers, computers, etc.) shall be clearly labelled. Rationale: Labelled cables will increase safety and decrease installation and troubleshooting time and facilitates future system expansions. | SAT, VT | Click the drop-down box to select
8.1.25 | Username and password The computer used for controlling the instrument shall use username and password to protect the instrument against unauthorized use. Rationale: Untrained personnel must be prevented from operating the instrument. | SAT, VT | Click the drop-down box to select
8.1.26 | Remote operation The instrument shall support remote operation. Remark: Protocol to be determined after awarding tender. Rationale: Efficient use of the instrument will be increased if the operator can control the instrument remotely. | SAT, VT | Click the drop-down box to select
8.1.27 | Abort button in controller software The controller software shall feature an abort button which stops immediately all motion, and will stop immediately the measurement. After the abort button is activated the instrument will await new commands. This button shall not power down the instrument. Rationale: Necessary for safety and convenience of the operation of the instrument | SAT, VT | Click the drop-down box to select
8.1.28 | Documentation The supplier shall deliver certificates, CE declarations of conformity, and manuals for all the third party hardware that is part of the system.. Rationale: TNO quality requirement. | SAT, V | Click the drop-down box to select
8.1.29 | Emergency stop button The instrument shall have at least one physical emergency stop button. Remark 1: This emergency stop will initiate the EMO state. Remark 2: The location of the emergency stop button shall be agreed on with TNO. Rationale: Machine and operator safety. | SAT, VT | Click the drop-down box to select
8.1.30 | Safely retraction probe from sample The instrument shall provide a means to safely retract the probe from the sample after an accident such as software crash, emergency stop, power outage, etc. Rationale: Damage prevention to the sample, probe or the instrument. | SAT, DT | Click the drop-down box to select
8.1.31 | Applicable delivery terms and Incoterm The supplier shall be responsible for the transport of the entire system from their facility to TNO facility where the instrument will be located, according Incoterm Delivered Duty Paid (DDP) 2020 (further specified in the General Purchasing Conditions TNO - February 2022). Rationale: Part of the delivery. | FAT | Click the drop-down box to select
Preferences
Preference | Decription | Method of Validation | Does Tenderer comply to the preference? | Reference in Tenderer's documentation where compliances with the preference is described | Remark
8.2.1 | NV magnetometry Demonstrated NV magnetometry with increased sensitivity better than 2µT/√Hz. Rationale: Higher sensitivity is desirable to reduce the measurement time or to measure on samples with smaller magnetic field variations. | SAT, DT | Click the drop-down box to select
8.2.2 | High speed magnetic field measurement Demonstrated high speed magnetic field measurement with a scan speed of more than 100 pixels/s while having a sensitivity (or equivalent minimum detectable field at such scan speed) better than 5µT/√Hz. Remark: Please specify the details. Rationale: Reduce in the measurement time with an acceptable sensitivity. | SAT, DT | Click the drop-down box to select
8.2.3 | Static magnetic field sensitivity Demonstrated increase in the static magnetic field sensitivity down to 100nT/√Hz using advanced techniques (such as gradiometry) with integrated software interface. Remark: Please specify the details. Rationale: Higher sensitivity is desirable to reduce the measurement time or to measure on samples with smaller magnetic field variations. | SAT, DT | Click the drop-down box to select
8.2.4 | AFM probe height control AFM probe height control including scanning at a constant height as well as scanning at an offset height from a previously measured line and a measured two-dimensional topography map. Remark: Please specify if the height control can be done only over the range of the z-piezo actuator or if it can be extended to be used also together with the coarse z-stage. Rationale: Measure the sample at a desired tip sample separation and while the probe is not engaged with the sample. | SAT, DT | Click the drop-down box to select
8.2.5 | Sample scanner noise Sample scanner noise better than 0.1nm RMS for XY and better than 50pm RMS for Z, for a bandwidth of at least 100Hz in XY and 1kHz in Z. Rationale: Enhanced precision, stability and dynamic performance of the scanner. | SAT, DT | Click the drop-down box to select
8.2.6 | Stitch coarse movements Software implemented possibility to stitch the coarse movements to achieve larger scan areas. Rationale: Increasing the effective scanning range. | SAT, DT | Click the drop-down box to select
8.2.7 | Sample size Possibility to mount samples with larger sizes: XY size of 50mm (or above if possible) and Z size up to 30mm. Remark: Larger sample volume capacity receives higher score, relatively. Rationale: Increased flexibility on the sample size is desirable | SAT, DT | Click the drop-down box to select
8.2.8 | Sample weight Increased sample weight capacity of up to 250g. Remark: Lower maximum weight capacity receives part of the score depending on the effect on the sample scanner performance. Please specify such effect. Rationale: Increased flexibility on the sample weight without degrading the scanner performance. | SAT, DT | Click the drop-down box to select
8.2.9 | Sample mount Offering ready-made or custom-made solutions (e.g. PCB plates), for mounting and accessing electrical connections to the sample. Rationale: Flexibility and ease of mounting and addressing samples with electrical connection is desired. | SAT, DT | Click the drop-down box to select
8.2.10 | Regulated environment Temperature regulated environment with temperature control of ±0.1 °C at a target temperature of 20 °C and a lab temperature range of 18 to 24 °C. Rationale: Temperature variation has negative influence on the stability of the system | SAT, DT | Click the drop-down box to select
8.2.11 | Vibration isolation stage Active vibration isolation stage with active isolation bandwidth of at least 1-100Hz or better and with a transmissibility of >20dB at 5Hz and >=40dB above 10Hz. Rationale: Increased stability of the system. | SAT, DT | Click the drop-down box to select
8.2.12 | Tip-sample drift rate Tip-sample drift rate of less than 2nm/h. Remark: Please indicate if a temperature controlled environment is needed to achieve such stability. Rationale: Lower drift rate is desirable for increased stability. | SAT, DT | Click the drop-down box to select
8.2.13 | Microwave source Microwave source with increased frequency range of up to 6GHz. Rationale: Increase flexibility on microwave generation. | SAT, DT | Click the drop-down box to select
8.2.14 | Green excitation laser The green excitation laser with a tuneable power between 0 to >10mW at the focal point, with rise/fall time of <2ns, and modulation/pulsing bandwidth of at least 200MHz. Rationale: Increase flexibility with the excitation laser that can be used in pulsed measurements. | SAT, DT | Click the drop-down box to select
8.2.15 | Confocal scanning and signal optimization Confocal scanning over the probe as well as automatic optimization of the NV signal, by for example (re-)tracking the NV fluorescence signal. Possibility to re-track the signal during the measurement if needed. Remark: Please specify the options available for NV signal re-tracking. Rationale: Please specify the options available for NV signal re-tracking. | SAT, DT | Click the drop-down box to select
8.2.16 | Stable optics Stable optics resulting in low drop in the detected NV fluorescence signal over several hours of measurement when no signal re-tracking is used. Remark: Please quantify the signal stability. Please indicate if a temperature controlled environment is needed to achieve such stability. Rationale: This indicates the robustness of the system without the need for extra signal re-tracking. The higher robustness with no signal re-tracking is desirable when long measurements are carried out. | SAT, DT | Click the drop-down box to select
8.2.17 | Automated sensitivity optimization routine Automated magnetic sensitivity optimization routine to set the best laser and microwave power. Rationale: Ensuring the optimal setting of the laser and microwave source in an easy-to-use manner. | SAT, DT | Click the drop-down box to select
8.2.18 | Access port Access port on the (back-side of the) microscope enclosure to feed in and out extra coaxial cables (e.g. BNC, SMA) and optical fibers, to and from outside. Rationale: Flexibility on feeding extra/custom cables and fibers. | SAT, DT | Click the drop-down box to select
8.2.19 | Microwave antenna positioning stage Microwave antenna positioning stage with positioning resolution of <1µm and positioning repeatability of better than ±2µm. Rationale: Repeatable and accurate positioning of the microwave antenna. | SAT, DT | Click the drop-down box to select
8.2.20 | Extended pulsed measurement modalities Extended pulsed measurement modalities with integrated GUI to control and display the data at least for the following modes: Pulsed ODMR, Rabi, Relaxometry, Ramsey, Spin-Echo, CPMG. Rationale: Expand the functionality and sensitivity of the instrument. | SAT, DT | Click the drop-down box to select
8.2.21 | Custom-defined pulse-sequences Possibility to apply custom-defined pulse-sequences and display the results. Rationale: Provide the possibility to implement complex pulse sequences that are not integrated in the software. | SAT, DT | Click the drop-down box to select
8.2.22 | Vector electromagnet Calibrated, low noise, and temperature stabilized vector electromagnet with strength tunability in arbitrary direction up to at least 50mT, including software interface to control and calibrate the strength and orientation of the field. Resolution and repeatability is expected to be better than 0.1mT and 0.1 degree. The magnetic noise due to the electromagnet in steady-state operation is expected to be below the noise floor of the rest of the system. Rationale: Increased flexibility in generating custom magnetic fields for magnetization of the sample and/or biasing the NV centre. | SAT, DT | Click the drop-down box to select
8.2.23 | Measurement data Record, save and export the measurement data in a user-accessible data format. This should include (all) the data that are displayed in the software user interface, including but not limited to: NV measurement results data, AFM-related data (topography, amplitude, phase), raw as well as processed fluorescence data and photon counts per pixel, hardware status including the laser and microwave signals, etc. Remark: Please specify if any specific data (e.g. raw data, optical images, confocal maps, etc.) that are accessible within the GUI can not be stored. Rationale: High flexibility in data recording and data accessibility is desired. | SAT, DT | Click the drop-down box to select
8.2.24 | Hardware access to APD signals Hardware access to the electronic signals from the APD. Rationale: Flexibility in accessing the relevant signals for custom hardware expansion and signal processing. | SAT, DT | Click the drop-down box to select
8.2.25 | Hardware access to electronic signals Hardware access to electronic signals or triggers (e.g. TTL) to monitor the state of the hardware, including at least AFM (pixel, line, frame), laser/AOM modulation, and MW sequence triggers. Rationale: Flexibility in accessing the relevant signals for custom hardware expansion and signal processing. | SAT, DT | Click the drop-down box to select
8.2.26 | Hardware access to MW source Hardware access to the analog I/O outputs used for the MW modulation. Rationale: Flexibility in accessing the relevant signals for custom hardware expansion and signal processing. | SAT, DT | Click the drop-down box to select
8.2.27 | Hardware access to clock signal Hardware access to a 10MHz clock signal. Rationale: Clock signal for external synchronization. | SAT, DT | Click the drop-down box to select
8.2.28 | Polarization control of laser Possibility to control the polarization of the laser. Rationale: Increased flexibility with laser excitation. | SAT, DT | Click the drop-down box to select
8.2.29 | Custom beam path Customized beam path by adding two extra beam lines with wavelengths of 594(±1)nm and 637(±1)nm that are confocal with the green laser. This customization includes the necessary hardware (lasers, optics, electronic interfaces) as well as software interface to control the beams (individually and in combination) from the user interface. The fluorescence loss from the green laser due to the additional optics should be less than 5%. Rationale: Advanced readout of the NV spin. | SAT, DT | Click the drop-down box to select
8.2.30 | Magneto-optic Kerr effect Integrated magneto-optic Kerr effect (MOKE) microscopy for increased speed in finding the regions of interest Rationale: Increased speed in finding the regions of interest. | SAT, DT | Click the drop-down box to select
8.2.31 | Scripting Possibility to control the hardware via scripting language, e.g. Python, etc. Remark: Please elaborate on the possibilities. Rationale: Flexibility in hardware control. | SAT, DT | Click the drop-down box to select
8.2.32 | Accessibility software Access to the future upgrades in the software and the user interface, free of charge, for at least the during of the expected lifetime of the instrument. Remark: Please specify the applicable terms. Rationale: TNO expects that the instrument is delivered with a software that in time will be updated with new features or bugfixes. TNO demands to receive these without any charge. | SAT | Click the drop-down box to select
8.2.33 | User interface software The user interface software shall support, free of charge, all future versions of the operating system of the instrument's computer (e.g. MS Windows) during the expected lifetime of the instrument. Rationale: With the expected lifetime of the instrument it can be expected that several new windows versions can be released. To ensure safety of the IT infrastructure the control PC of the instrument might be required to update to the latest version. When this is required, it is vital for the IT infrastructure that this upgrade is performed, not doing this will result in the instrument being disconnected from the IT infrastructure, therefore hindering the operation of the instrument. | SAT | Click the drop-down box to select
8.2.34 | Compatible NV probes Availability of compatible NV probes up to the minimum expected lifetime of the system. Rationale: The instrument needs compatible NV probe (consumable). | SAT | Click the drop-down box to select
8.2.35 | Compatible probes with integrated antenna Possibility of using the instrument with compatible probes including integrated MW antenna. Rationale: Expected increase in the flexibility, reproducibility and ease of use. | SAT, DT | Click the drop-down box to select
8.2.36 | Inverted optical configuration Compatibility with inverted optical readout configuration for transparent samples used e.g. in biological or NMR applications. Rationale: Provide possibility to use the system for relevant biological or NMR applications where transparent samples are used without the need to acquire a fully separate instrument. | SAT, D | Click the drop-down box to select

## validatie (hidden)
Click the drop-down box to select
Tenderer declares to comply with the requirement
Tenderer declares non-compliance with requirement
Click the drop-down box to select
Yes, Tenderer can comply with preference
No, Tenderer cannot comply with preference
note: dit blad 'verbergen' indien je dit gaat delen intern/extern (behoudens aan inkoop collega's)