## Appendix A04
Appendix A04, Conformity list 'Programme of Requirements and Preferences' Guidelines for the European Public Tender for a Closed-Cycle refrigerator Cryostat system, ref. TNO 2024 FPL-INK 081, dated 30-08-2024
By the following Programme of Requirements and Preferences TNO will assess if Tenderer meets the relevant requirements or preferenes as laid down in Section 6 and 8. The Tenderer must indicate for each item, the conformity for the requirements and preferences in the table below.
If to any of the requirements, as stated in section 8, are not met by Tenderer, he will be excluded for any further assessment. By answering a question with “comply” will mean that in the event of awarding the contract, non-compliance to the requirement shall count as failure to meet an obligation from the Agreement to be concluded. If Tenderer does not comply with any of the preferences, as mentioned in point 8, he will not be excluded from further assessment. Answering a preference with "comply with" means that, at the time of award, failure to comply with the preference counts as failure to fulfil an obligation under the Agreement to be concluded. The abbreviations of the validation method during SAT are described in section 8.
Requirements
Requirement | Decription | Method of Validation | Does Tenderer comply to the requirement? | Reference in Tenderer's documentation where compliances with the requirement is described | Remark
Requirement R-0000-005 | The supplier shall deliver a system that is capable of meeting all the stated requirements in this document at the same conditions and time, unless indicated otherwise in a specific requirement. | R | Click the drop-down box to select
Requirement R-0000-010 | The system shall be an existing product, which has been produced, tested and delivered at least once before. | R | Click the drop-down box to select
Requirement R-0000-015 | The system shall be able to cool down the device under test (DUT). | I | Click the drop-down box to select
Requirement R-0000-020 | The cryostat shall be electrically isolated from the gas handling system, compressor and control unit. | T | Click the drop-down box to select
Requirement R-0000-025 | The cryostat pumping system shall be able to keep a pressure below 10e-5 mbar at the room temperature platform on the inside of the cryostat when the cryostat’s temperature is at 1 Kelvin. | D | Click the drop-down box to select
Requirement R-0000-030 | The cryostat shall be able to go from ambient to cryogenic temperatures and vice-versa without requiring any physical action from the operator besides performing the command via the remote software. | D | Click the drop-down box to select
Requirement R-0000-035 | The cryostat shall be able to pump-down and vent without requiring any physical action from the operator besides performing the command via the remote software. | D | Click the drop-down box to select
Requirement R-0000-040 | The system shall be of the type closed-cycle cryogen-free cryostat. | I | Click the drop-down box to select
Requirement R-0000-045 | CE certification – the cryostat shall be CE certified. | C | Click the drop-down box to select
Requirement R-1000-005 | The Cryostat shall have a base temperature of less than 1 kelvin, as measured on the cold plate. | D | Click the drop-down box to select
Requirement R-1000-010 | The cooling power of the Cryostat at 1 Kelvin shall be at least 30 mW with all components in place and thermalized (e.g. the vector magnet, optical windows and optical fibers, RF lines and DC wiring) as measured on the cold plate. | D | Click the drop-down box to select
Requirement R-1000-015 | The Cryostat shall be able operate continuously at the 1 Kelvin or below with a heat load of 30 mW, for at least 1 week. | D | Click the drop-down box to select
Requirement R-1000-020 | The Cryostat shall be monitored by at least one thermometer per stage with a resolution of 0.1 Kelvin or better. | D | Click the drop-down box to select
Requirement R-1000-025 | The Cryostat shall be monitored by at least one thermometer per stage calibrated with an accuracy of 0.2 Kelvin or better. | C | Click the drop-down box to select
Requirement R-1000-030 | The Cryostat shall be upgradable to a working dilution refrigerator capable of base temperatures below 100 mK. | A | Click the drop-down box to select
Requirement R-1000-035 | The Cryostat shall ensure that the temperature of the cold plate has a maximum peak to peak variation of 0.05 K with a varying heat load. | D | Click the drop-down box to select
Requirement R-1000-040 | The vacuum housing of the cryostat shall be removable in separate parts, each part weighing equal or less than 22.5 kg. | I | Click the drop-down box to select
Requirement R-1000-045 | The vacuum housing of the cryostat shall be removable in the direction of the floor, and the housing design shall take this into account. | I | Click the drop-down box to select
Requirement R-1000-050 | The required tooling for the removing of the vacuum housing of the cryostat shall be included in the delivery. | I | Click the drop-down box to select
Requirement R-1100-005 | The Cryostat shall contain a free-space optical interface connecting the optics at atmospheric pressure to the DUT using an optical window on the top platform of the cryostat and removable optical windows at the 50 K, 4 and 1 K stages of the cryostat. | I | Click the drop-down box to select
Requirement R-1100-010 | The optical window on the housing shall be located centered on the top of the cryostat. | I | Click the drop-down box to select
Requirement R-1100-015 | The optical windows on the stages within the cryostat shall be located centered on the stage and allow for line-of-sight between the window of the housing and the center of the sample space. | I | Click the drop-down box to select
Requirement R-1100-016 | At the Room Temperature, 50K, 4K and 1K stages there shall be a clear aperture of at least 40 mm when the optical windows are removed. | I | Click the drop-down box to select
Requirement R-1100-017 | At the 50K stage there shall 6 holes surrounding the optical window such that a lens mount can be mounted there. | I | Click the drop-down box to select
Requirement R-1100-025 | The optical window on the housing (room temperature platform) shall have a diameter of 40 mm. | D | Click the drop-down box to select
Requirement R-1100-030 | The mount for the optical window on the housing (room temperature platform) shall be of type KF40. | D | Click the drop-down box to select
Requirement R-1100-040 | Where used, the optical windows shall have a transmission of 95% or more for light of any wavelength within the range between 532 nm to 800 nm. | T | Click the drop-down box to select
Requirement R-1100-045 | Where used, the optical windows shall have a transmission of 97% or more for light of any wavelength within the range between 619 nm to 637 nm. | T | Click the drop-down box to select
Requirement R-1100-055 | Where used, the optical windows defined in requirements R-3000-010 and R-3000-015 shall have a transmitted wavefront error of λ/4 or better (less). | I | Click the drop-down box to select
Requirement R-1100-060 | The optical windows at every stage (including the housing on top of the cryostat) shall be centered with respect with the optical axis of the cryostat with a maximum deviation of 1 mm from the axis. | I | Click the drop-down box to select
Requirement R-1210-005 | The Cryostat shall contain at least 10 RF feedthroughs and RF lines from room temperature to the sample space. | I | Click the drop-down box to select
Requirement R-1210-010 | The RF Feedthroughs in the Cryostat shall be of the type SMA on the room and vacuum side. | I | Click the drop-down box to select
Requirement R-1210-015 | The RF lines shall connect the feedthroughs to the sample space. | I | Click the drop-down box to select
Requirement R-1210-020 | The RF lines shall be thermalized at each stage. | I | Click the drop-down box to select
Requirement R-1210-025 | The RF lines shall consist of segments between: - The room temperature and 50 k stage - The 50k and the 4K stage - 4K stage and final (1K or below) stage | I | Click the drop-down box to select
Requirement R-1210-035 | The RF lines shall have a bandwidth of 15 GHz or more, defined as 0-15 GHz. | T | Click the drop-down box to select
Requirement R-1210-040 | The RF lines shall have SMA connectors at the sample space side. | I | Click the drop-down box to select
Requirement R-1210-045 | The RF lines shall allow the inclusion of a patch cable to convert the output connector to SMP at the sample space. | I | Click the drop-down box to select
Requirement R-1210-050 | The RF lines shall have a characteristic impedance of 50 Ohm. | T | Click the drop-down box to select
Requirement R-1210-055 | The RF lines shall be able to handle an input RF power on the room temperature side of at least 33 dBm (2 W). | T | Click the drop-down box to select
Requirement R-1210-060 | The RF lines shall have an attenuation below 5 dB between the input of the feedthrough to the output of the lines at the sample space, in the range of 0-15 GHz. | T | Click the drop-down box to select
Requirement R-1210-065 | At least 5 of the RF lines shall have an attenuation below 3 dB between room temperature and the sample space, in the range of 0-15 GHz. | T | Click the drop-down box to select
Requirement R-1220-005 | The Cryostat shall contain DC feedthrough and DC lines separated over 3 types: - Sample Scanner signal lines - Sample Scanner readout lines - Experimental lines. | R | Click the drop-down box to select
Requirement R-1220-010 | The DC lines shall be thermalized at each stage. | I | Click the drop-down box to select
Requirement R-1220-015 | The DC lines shall consist of segments between: - The room temperature and 50 k stage - The 50k and the 4K stage - 4K stage and final (1K or below) stage | I | Click the drop-down box to select
Requirement R-1220-020 | The Sample Scanner signal lines shall consist of at least 6 twisted pairs. | I | Click the drop-down box to select
Requirement R-1220-025 | The Sample Scanner signal lines shall have an electrical resistance of 5 ohm or less. | T | Click the drop-down box to select
Requirement R-1220-030 | The Sample Scanner signal lines shall have a bandwidth of 100 kHz or more. | T | Click the drop-down box to select
Requirement R-1220-035 | The Sample scanner readout lines shall consist of at least 6 twisted pairs. | I | Click the drop-down box to select
Requirement R-1220-040 | The Sample scanner readout lines shall have an electrical resistance of 20 ohm or less. | T | Click the drop-down box to select
Requirement R-1220-045 | The Sample scanner signal and readout feedthrough connectors on the room temperature side shall be of the type Fischer connector. | I | Click the drop-down box to select
Requirement R-1220-050 | The Sample scanner signal and readout lines connectors on the sample space side shall be of the type micro-D connector. | D | Click the drop-down box to select
Requirement R-1220-055 | The experimental lines shall consist of at least 12 twisted pairs. that allow for voltages up to 300 V with a bandwidth of 100 kHz or more. | I | Click the drop-down box to select
Requirement R-1220-060 | The experimental feedthrough connectors shall be of the type fischer connector, sub-D or micro-D. | I | Click the drop-down box to select
Requirement R-1220-065 | The experimental lines connectors at the sample space shall be of the type micro-D. | I | Click the drop-down box to select
Requirement R-1230-005 | The Cryostat shall contain at least 16 optical fibers and feedthrough's. | I | Click the drop-down box to select
Requirement R-1230-010 | The optical fiber feedthroughs shall be suited for FC-APC fiber connectors. | I | Click the drop-down box to select
Requirement R-1230-015 | There shall be optical fiber lines to connect the input feedthrough to the sample space. | I | Click the drop-down box to select
Requirement R-1230-020 | The optical fiber lines shall be thermalized at each plate. | I | Click the drop-down box to select
Requirement R-1230-025 | 8 out of the 16 fiber lines shall be of the type PM460-HP fiber with FC-APC connectors. Remark: Polarization maintaining is required too | I | Click the drop-down box to select
Requirement R-1230-030 | 8 out of the 16 fiber lines shall be of the type PM630-HP fiber with FC-APC connectors. | I | Click the drop-down box to select
Requirement R-1230-035 | The optical fiber lines shall provide strain relief at the sample space end of each fiber. | I | Click the drop-down box to select
Requirement R-1300-005 | The Vector Magnet shall have a bore of 93 mm or larger. | R | Click the drop-down box to select
Requirement R-1300-010 | The sample space (empty for the user to place components) shall have a height of at least 160 mm, as defined from the center of the magnetic field within the vector magnet to the bottom of the experimental space, and a radius of at least 93 mm. | I | Click the drop-down box to select
Requirement R-1300-015 | The Vector magnet shall be able to create a magnetic field in the range of 0 to 1 Tesla in steps of 0.1 milliTesla or smaller in the X, Y and Z directions. | T | Click the drop-down box to select
Requirement R-1300-020 | The Vector magnet shall be equipped with the persistent switches for all three magnetic coils. | D | Click the drop-down box to select
Requirement R-1300-025 | The Vector magnet shall be monitored by at least one thermometer. | I | Click the drop-down box to select
Requirement R-1300-030 | The vector magnet shall have a homogeneity of 1% or better over 10 mm diameter spherical volume in the Z direction. | D | Click the drop-down box to select
Requirement R-1300-035 | The vector magnet shall have a homogeneity of 1% or better over 10 mm diameter spherical volume in the X and Y direction. | D | Click the drop-down box to select
Requirement R-1300-040 | The system shall include a suitable power supply for the vector magnet. | I | Click the drop-down box to select
Requirement R-2000-005 | The supporting frame of the cryostat shall have a height equal to or below 2,60 m. | I | Click the drop-down box to select
Requirement R-2000-010 | At a facility with 320 cm height, the supporting frame of the cryostat shall allow for the placement of an optical breadboard of 90 cm by 90 cm, with 12,7 mm height, adjacent to the window on top of the cryostat mentioned in requirement R-1100-010. | I | Click the drop-down box to select
Requirement R-2000-015 | The system shall not contain components above a height of 3,00 m. | I | Click the drop-down box to select
Requirement R-2000-020 | The supporting frame of the cryostat shall have a maximum width of 1,30 m and a maximum length of 1,80 m. | I | Click the drop-down box to select
Requirement R-2000-025 | The supporting frame shall ensure vibrations in the horizontal plane are below 10 um RMS between the cold plate and support frame. | D | Click the drop-down box to select
Requirement R-3000-005 | The system shall be controllable and monitorable via a control unit. | I | Click the drop-down box to select
Requirement R-3000-010 | The system shall be controllable and monitorable remotely via a software package. | I | Click the drop-down box to select
Requirement R-3000-015 | The software package shall be able to run on windows and/or linux computers. | I | Click the drop-down box to select
Requirement R-3000-020 | The software package shall allow to turn off internet-enabled functionality, such as: - (Automatic) updates - Forwarding of usage information Remark: If the software package doesn’t contain internet-enabled functionality, this requirement is not applicable. | I | Click the drop-down box to select
Requirement R-3000-025 | The software package (requirement R-3000-010) shall be able to control the following functionality: - Pumping down and venting the system - Cooling down the sample to a temperature specified by the user down to the base temperature - Warm-up the system to ambient temperature - Power of the heaters at the 1K plate | I | Click the drop-down box to select
Requirement R-3000-030 | The software (requirement R-3000-010) shall be able to show the following system parameters: - Vacuum pressure - Temperature of each stage - Temperature of the DUT - Temperature of the Vector magnet - Magnetic field magnitude (based on a recent (<5 second) measurement of the electrical current that is computed to the magnetic field magnitude using a known & measured calibrated magnetic field magnitude vs electrical current dataset) | I | Click the drop-down box to select
Requirement R-3000-035 | The software package and python interface shall show the current values for the parameters defined in requirement R-3000-030 with a maximum delay of 5 seconds. | D | Click the drop-down box to select
Requirement R-3000-040 | The software package shall display a graph with the evolution of the parameters defined in requirements R-3000-030 for the previous 3 months. | I | Click the drop-down box to select
Requirement R-3000-045 | The software package and python interface shall allow the user to monitor the value of the parameters defined in requirement R-3000-030 with the minimum resolutions of (or better): - Temperature - 0.1 Kelvin - Pressure – 1 x 10-7 mbar | I | Click the drop-down box to select
Requirement R-3000-050 | The software package and python interface shall allow the user to set the temperature value with a minimum resolution of: - Temperature - 0.1 Kelvin | D | Click the drop-down box to select
Requirement R-3000-055 | The control unit shall allow be able to control the following functions: - Pumping down and venting the system - Cooling down the sample to a temperature specified by the user down to the base temperature - Warm-up the system to ambient temperature - Power of the heaters at the 1K plate | I | Click the drop-down box to select
Requirement R-4000-005 | The Gas handling system shall be able report the pressure of the gasses to the user. | I | Click the drop-down box to select
Requirement R-4000-010 | The system shall warn the user when the pressure of the gasses is too low. | I | Click the drop-down box to select
Requirement R-4000-015 | The system shall include the required gases for the operation of the system. | I | Click the drop-down box to select
Requirement R-4000-020 | The gas handling system shall be possible to be placed at a distance of 6 meter or more from the cryostat. | I | Click the drop-down box to select
Requirement R-4000-025 | The gas handling system shall be able to be filled by the operator. | I | Click the drop-down box to select
Requirement R-4000-030 | The gas handling system shall include all tooling required to be filled by the operator. | I | Click the drop-down box to select
Requirement R-4000-035 | The Gas Handling System components shall all be contained in a single rack/cabinet. | I | Click the drop-down box to select
Requirement R-5000-005 | The compressor shall be controlled remotely via a software package. | I | Click the drop-down box to select
Requirement R-5000-010 | The compressor shall have a strategy to prevent the existence of contaminant gases in the helium. | I | Click the drop-down box to select
Requirement R-5000-015 | The compressor shall be water cooled. | I | Click the drop-down box to select
Requirement R-5000-020 | The compressor shall be possible to be placed at a distance of 6 meter or more from the cryostat. | I | Click the drop-down box to select
Requirement R-5000-025 | The compressor shall have a minimal lifetime of 2 years of continuous operation. | I | Click the drop-down box to select
Requirement R-6000-005 | The electric power connections of the System shall have double poled fusing or circuit breakers (fuses placed in both the phase (s) and neutral conductors). | I | Click the drop-down box to select
Requirement R-6000-010 | The electric power supplies for the system shall be of the types: 240 Volt, 1 phase, 50 Hz and / or 400 volt, 3 phase, 50 Hz. | I | Click the drop-down box to select
Requirement R-6000-015 | The gas handling system and compressor shall able to be located in a separate room, separated by a wall from the cryostat. | I | Click the drop-down box to select
Requirement R-6000-020 | The gas handling system and compressor together shall be able to fit on a floorspace of 2 meter wide by 1 meter long. | I | Click the drop-down box to select
Requirement R-6000-025 | The shipping boxes including the equipment shall fit through a door of 250 cm wide and 250 cm height. | I | Click the drop-down box to select
Requirement R-6000-030 | The shipping boxes including the equipment shall not weight more than 1500 kg per shipping box. | I | Click the drop-down box to select
Requirement R-6000-035 | The unboxed equipment shall have a maximum size of 130cm wide and 202 cm height. | I | Click the drop-down box to select
Requirement 8.1.13.1 | Critical design review (CDR) The objective of the CDR is to confirm that the equipment and all deliverables satisfies the requirements. It establishes the compatibility of external & internal interfaces and the compliance of the design with contractual requirements (predicted and/or measured). The successful completion of the CDR permits the Supplier to start the manufacturing of the system. The CDR will contain the following details: - Physical dimensions of the system and system components. - External and internal interface location and dimensions of the system. - Showing the design meets the functional and performance requirements. The acceptance of the CDR shall by achieved using a written document, to be signed by both parties. In the event that the CDR fails to meet the requirements, the supplier or TNO must set out such failure in a written statement, to be signed by both parties. The supplier must then provide TNO with corrective action plan within fourteen (14) calendar days from the date of signing. The supplier must remedy these deficiencies at no additional cost within the timeframe set out in the corrective action plan, but not exceeding 1 month. | R | Click the drop-down box to select
Requirement 8.1.13.2 | Site Acceptance Test (SAT) The SAT contains the following checkpoints: - The cryostat will be checked for damages; - The proper functioning of the equipment will be checked; o Mechanical functions, o Electrical safety- and switching functions, o Gas safety and functions, o Vacuum functionality, o Cryogenic functionality, o User interface(s), including software. - All requirements (such as requirements R-0000-025, R-1000-005, R-1000-10, R-1300-10, etc.) and possible additional requirements will be checked; - The presence and quality of the documentation will be checked A full user acceptance test on-site of the contracted tenderer must be performed by the Tenderer in cooperation with TNO. Following approved user acceptance test on-site, the cryostat shall be at the unrestricted disposal of TNO for one month. If the cryostat works according to TNO’s functional and technical specifications, without faults and/or malfunctions during operational mode, the test period is considered successfully concluded by TNO. If the cryostat is not operating in this period according TNO’s functional and technical specifications, the faults and/or malfunctions will be remedied and another test period of one month must successfully be concluded and approved by TNO, up to a maximum of three test periods of one month each. The details of the site acceptance test shall be agreed between the user and the supplier. Tenderer must submit a draft test protocol for the site acceptance test with its Tender. In the event that the deliverables fail to meet the user Acceptance Criteria for the SAT, the Supplier must set out such failures in a written statement, to be signed by both parties. The supplier must then provide TNO with a corrective action plan for such failures within fourteen (14) calendar days from the date of signing of such statement. The supplier must remedy these deficiencies at no additional cost within the timeframe set out in the corrective action plan, but not exceeding 3 months. A new acceptance test will then take place. | I | Click the drop-down box to select
Requirement 8.1.13.3 | Terms and conditions of payment - 20% at order, after receipt of a bank guarantee covering 40% of the total amount. - 20% after the critical design review (CDR) acceptance by TNO (Requirement 8.1.13.1). - 40% after delivery at location TNO Delft, STW At this moment TNO will return the bank guarantee. - 20% after installation and commissioning On-site and Site Acceptance Test (SAT) including test periods, approved by TNO, Management Quantum Technology and Procurement. The Tenderer to whom the Contract will be awarded must provide a bank guarantee to TNO, covering at least 40% of the total price. The Bank guarantee and the bank who issues the Bank guarantee have to be approved by TNO (minimal A rating). TNO must have the right to draw the bank guarantee in any and all cases where Tenderer falls short in fulfilling its obligations under the Contract. The validity of the bank guarantee may not expire until approval of the SAT by TNO Management and Procurement. | R | Click the drop-down box to select
Requirement 8.1.13.4 | When supplying the cryostat, the Tenderer shall provide a Mill Test Certificate (MTC), or similar certificate such that it meets the requirements and obligations to establish the country of origin (further specified in paragraph 5.2.4). | C | Click the drop-down box to select
Requirement 8.1.13.5 | Maintenance/services Contract, further specified in paragraph 1.5. | R | Click the drop-down box to select
Requirement 8.1.13.6 | The Tenderer must have comprehensive maintenance facilities within Europe. These facilities should be equipped to handle all aspects of maintenance, including diagnostics, repairs, and parts replacement. The facilities must be strategically located to ensure quick and efficient service delivery | R | Click the drop-down box to select
Requirement 8.1.13.7 | User manual with detailed description of the hardware and software, including any custom made options. | R | Click the drop-down box to select
Requirement 8.1.13.8 | The Contract will be executed through a purchase order, incorporating the Purchasing Conditions for Goods TNO 2022 and Incoterms 2020 Delivered Duty Paid (DDP), see paragraph 2.2.18. | R | Click the drop-down box to select
Preferences & Questions
Preference | Decription | Method of Validation | Does Tenderer comply to the preference? | Reference in Tenderer's documentation where compliances with the preference is described | Remark
Preference QY 1 | As our current research line has uncertain factors, it is important that we have the flexibility to upgrade the Sub-Kelvin fridge to a dilution refrigerator, allowing to change the purpose of the fridge. Describe the process of upgrading the Sub-Kelvin fridge to a full dilution refrigerator and the specifications of the resulting dilution refrigerator. The following items must be elaborated by the Tenderer in this regard: - How long will the fridge be unusable for experiments during the upgrade? - What parts need to be added or replaced in the upgrade? - What is the cooling power for 100mK? - How many litres of He-3 and He-4 are necessary for operation? The complete answer for Preference QY 1 must not exceed a maximum of three (3) A4 pages. If the answer is longer than 3 pages, the excess pages will not be included in the assessment. Also, the Tenderer should submit its answer to each bullet point (!) and in the same order ("from top to bottom") of the bullet points under this preference. | R | Click the drop-down box to select
Preference QY 2 | As stated in requirement R-1000-010 the minimum cooling power at 1K stands at 30 mW. For expanding our experiments in the future, more cooling power at 1K will allow for more complicated experiments. Specify the cooling power at 1K with all components in place and thermalized (e.g. the vector magnet, optical window(s), and optical fibers, RF lines and DC wiring) as measured on the cold plate. The complete answer for Preference QY 2 must not exceed a maximum of one (1) A4 page. If the answer is longer than 1 page, the excess pages will not be included in the assessment. | D | Click the drop-down box to select
Preference QY 3 | The system should allow easy integration into our experimental setup and software framework, including remote control and readout of relevant system operations (such as temperature, pressure, etc.). Describe the control API, the features of the work panel and other features related to the control of the system, including (continuous) availability of drivers for different operating systems and automated scripts for operation of the system. Some of our preferences: - A properly documented protocol for monitoring and control is strongly preferred over access exclusively via programming-language specific API support that ultimately depends on binary components, such as DLLs for Windows or shared object (.so) files for Linux. Rationale: Such binary drivers are often tied to specific versions of the OS, and introduce undesirable dependencies in the software, and may render the device useless in the future if the OS makes an ABI-breaking change and the device vendor does not update their drivers. - A hardware interface based on Ethernet using TCP/IP is preferred over alternatives (e.g. Ethernet/UDP, Ethernet/Custom, USB-based interfaces using USBTMC or serial-over-USB). Rationale: Ethernet provides more flexibility with respect to the physical placement of the device relative to the controlling PC, and is generally a more robust interface compared to USB. TCP provides flexibility to monitor and control device traffic. An automated and robust operation of the system is important for our experiments. Describe how the system ensures automated and robust operation. Include in your answer at least the following: - Does the fridge rely on a local PC for operation? - Is the fridge unaffected by a crash of any PC connected to it? - Is the valve of the bypass of the compressor closed automatically after condensation? - Does the fridge store a user-accessible log of the received commands and errors it encountered? The complete answer for Preference QY 3 must not exceed a maximum of two (2) A4 pages. If the answer is longer than 2 pages, the excess pages will not be included in the assessment. Also, the Tenderer should submit its answer to each bullet point (!) and in the same order ("from top to bottom") of the bullet points under this preference. | R | Click the drop-down box to select
Preference QY 4 | As broken parts can cause downtime of the system, it is important that a good support infrastructure is present. Include in your answer the following points: - How is preventive maintenance of the fridge organized? - How is the repair or replacement of broken components organized? - What is the normal interval between maintenances? Indicate the typical timeframe (from failure to restart of the experiments) for replacing a critical component in the system including: - A broken part in general - Compressor parts - Cryostat parts (e.g. the pulse tube) - A broken pump - A broken turbo - A broken magnet Describe the required preventive maintenance and schedule of the cold parts, the vacuum system, and the electronics. What is the lifetime of the compressor? This should meet minimally requirement R-5000-025. The complete answer for Preference 8.2.1 must not exceed a maximum of three (3) A4 pages. If the answer is longer than 3 pages, the excess pages will not be included in the assessment. Also, the Tenderer should submit its answer to each bullet point (!) and in the same order ("from top to bottom") of the bullet points under this preference. | R | Click the drop-down box to select
Preference QY 5 | As our current research line has uncertain factors, it is important that we have the flexibility to upgrade the Sub-Kelvin fridge to a dilution refrigerator, allowing to change the purpose of the fridge. With this said we value previous experience on delivering systems that fulfill our needs as presented in this document. Demonstrate your previous experience in upgrading cryostats to full dilution fridges. Include in your answer the following points: - How many 1 K systems has the tenderer upgraded to a sub-Kelvin fridge? - TNO prefers a supplier with more experience The complete answer for Preference QY 5 must not exceed a maximum of two (2) A4 pages. If the answer is longer than 2 pages, the excess pages will not be included in the assessment. | R | Click the drop-down box to select
Preference QY 6 | In our experiments, we may have multiple emitters on a single DUT and it may be beneficial to minimize the differences in the magnetic field they are subjected to. The Tender which offers the highest Field Homogeneity (FH), corresponding to the lowest variation of magnetic field over 10 mm diameter spherical volume (DSV) from the center of the vector magnet bore, will be awarded the maximum of 25 points. Other Tenders will be scored pro rata in inverse proportion to Field Homogeneity. As stated in Requirement R-1300-025 and R-1300-030 a variation equal to or below 1% over a 10 mm diameter spherical volume minimum values is required, thus, tenders showing a higher variation than this will not be considered. The complete answer for Preference QY 6 must not exceed a maximum of one (1) A4 page. If the answer is longer than 1 page, the excess pages will not be included in the assessment. |  | Click the drop-down box to select
Preference QY 7 | In our experiments we are looking for low drifts in the magnetic field when it is in persistent mode. To ensure similar conditions for experiments performed at different points in time. Describe the expected drift of the magnetic field in ppm/hour or percentual points and the measures that are taken to minimize the drift. This should take into account the situation with the persistent switches turned on. The complete answer for Preference QY 7 must not exceed a maximum of one (1) A4 page. If the answer is longer than 1 page, the excess pages will not be included in the assessment. |  | Click the drop-down box to select
# | Questions |  |  | Answer to question
Question 1 | Corporate Social Responsibility is becoming increasingly important in our own organization and in the future TNO will also ask more from our suppliers to play a part in this. What possibilities have you already initiated and where do you see opportunities for yourself and for TNO in future?
Question 2 | As you're aware, ISO 20400 is an international standard that provides guidelines on sustainable procurement. The standard can be used by any organization that intends to improve their social, economic, and environmental sustainability. Considering the growing awareness with regard to sustainable business practices, nationally and internationally, whether required by laws and regulations, promoted by the UN Sustainable Development Goals, or even by the changing societal expectations – there are a myriad of reasons in today’s world that drive organizations to be more sustainable, including TNO. Do you embrace the ISO 20400 guideline and, if so, are you currently implementing it in your own organisation or will you be implementing it in the coming period?
Question 3 | Laboratory instruments that TNO uses for its scientific research consume a lot of power and also consist of very high-quality raw materials. For the benefit of the corporate social responsibility (CSR), are you making or willing to make technological investments to reduce power consumption during the use of the instrument, without negatively affecting the technical application of the instrument?
Question 4 | Biobased Goods are made of materials derived from fully renewable sources. TNO is considering to require in future Goods that devices be made (in part) from biobased materials. Can you give TNO some insight into whether this is feasible for the Goods in your portfolio, and if so, which biobased Goods do you think could be used as alternative Goods?
Question 5 | Are you experiencing negative impacts from global problems in your supply chain and, if so, what measures have you taken to ensure, for example, that you can meet agreed delivery deadlines?

## validatie (hidden)
Click the drop-down box to select
Tenderer declares to comply with the requirement
Tenderer declares non-compliance with requirement
Click the drop-down box to select
Yes, Tenderer can comply with preference
No, Tenderer cannot comply with preference
note: dit blad 'verbergen' indien je dit gaat delen intern/extern (behoudens aan inkoop collega's)