| Tenderer: | <fill in> |
| --- | --- |

Schedule of requirements

Plasma enhanced Atomic Layer

Deposition System

# 1	Introduction

The Tenderer’s offer must fulfil all the listed requirements in this schedule. Not fulfilling one (or more) of the requirements, means exclusion from this tendering procedure.

The Contracting authority wants to verify some of the requirements, as indicated by the letter ‘T’ (Text). In these cases, documentation must be submitted with your tender and the table in section 3 of this schedule must be filled in. The documentation is free format, with the exception of power point presentations (and likewise) or references to websites.

Verification takes place using the following criteria:

- Complete: all information needed to verify are present;

- Correct: the information is verifiable, specific, realistic, acceptable and (where possible) measurable and formulated in a time-bound manner;

• 	Consistent: the information provided is relevant, in line with the Assignment, and consistent with information given elsewhere;

• 	Convincing: the information given convinces the Client that the requirement is fulfilled.

# 2	Requirements

Requirement 1 General system description

| Req. | Description |
| --- | --- |
| 1.1 (T) | The total footprint, including all the apparatus for a complete working system, should fit in the restricted space indicated by the color green and number 1, 2 and 3 on the supplied ground plan (see Annex 5, which specifies dimensions). The ground plan shows other apparatus and items currently in the grey room. The numbers indicate: - Available space for ALD system in technical room. - Available space for ALD system in cleanroom. - Cabinet holding pumps. - Future plasma etching system. - Current Rapid Thermal Processor (AnnealSys ASOne). - Table with control screens for various equipment including the new ALD. - Electrical cabinets for the lab. Roughing pumps supplied with the system must be placed in the pump cabinet indicated as 3 on the floorplan. |
| 1.2 | System must be installed as a through-the-wall solution in a technical space and a cleanroom. |
| 1.3 | System must be able to process samples sizes down to 5×5 mm² and up to 100 mm wafers. |
| 1.4 | System must be prepared for a retrofit installment of an ellipsometer, either to be supplied by the Tenderer or by a third party at a later time. |
| 1.5 (T) | System should have a strategy in place to protect the vacuum pump from precursor material. |
| 1.6 | System must be able to deposit high quality oxide, nitride and metal layers (see requirement 2 Process specifications). |
| 1.7 | System must be equipped with the necessary vacuum pump units (see requirement 3 Vacuum system). |
| 1.8 | System must be equipped with enough precursor lines to allow for high material flexibility (see requirement 4 Precursor, reactant and gas delivery system). |
| 1.9 | System must be of the plasma-enhanced atomic layer deposition type and able to perform both thermal ALD and plasma-assisted ALD in the same chamber and within a single recipe without need for hardware re-configuration. (see requirement 5 Plasma system). |
| 1.10 | System must be equipped with a load-lock for substrate loading (see requirement 6 Load-lock). The load lock must be located on the cleanroom side of the wall. |
| 1.11 | System must be able to maintain both a process temperature and a different wall temperature (see requirement 7 Heating and temperature control). |
| 1.12 | System must be supplied with a comprehensive user manual (see requirement 9 Service, maintenance and warranty for further requested documentation). |
| 1.13 | System should be able to be connected to the Dutch electrical net. If not, proper connections and/or transformers must be included and included in the total price. |

Requirement 2 Process specifications

| Req. | Description |
| --- | --- |
| 2.1 (T) | System must be able to perform deposition of 10 nm Al2O3 with the following specifications: Dielectric constant; ≥8.0. Breakdown voltage; ≥6.8 MV/cm. Thickness variations; ≤2% over 100 mm wafer. Precursor; Trimethylaluminum. Reactant; Water. |
| 2.2 (T) | System must be able to perform deposition of 10 nm AlN with the following specifications: Thickness variations; ≤2% over 100 mm wafer. Precursor; Trimethylaluminum. Reactant; N2/H2. |
| 2.3 (T) | System must be able to perform deposition of 40 nm of Pt with the following specifications: Bulk resistivity; < 20 µΩ-cm. Thickness variations; ≤2% over 100 mm wafer. Precursor; Trimethyl(methylcyclopentadienyl)platinum(IV). Reactant; O2. |
| 2.4 (T) | Wafer thickness reproducibility must be ≤ 2% over 10 subsequent depositions measured on 100 mm wafers excluding a 1 mm zone at its edge. |

Requirement 3 Vacuum system

| Req. | Description |
| --- | --- |
| 3.1 (T) | Roughing pumps supplied with the system must be of the Edwards nXDS10i, nXDS35i or, if chemical resistance is necessary, appropriate (e.g. iGX100M, iXM100, iXH100, etc.) oil-free pump models to allow interchanging between already available pumps in the lab. |
| 3.2 (T) | System must be equipped with a turbomolecular pump to allow for base pressures below 1×10-6 mBar. |
| 3.3 (T) | System must be equipped with gaskets and other seals that can withstand the temperatures ranges used in the system and allow for base pressures below 1×10-6 mBar. |
| 3.4 | All the necessary vacuum valves and vacuum pressure control should be present for a fully automatic vacuum system. |
| 3.5 | The system should pass a pressure rise test of < 5×10-5 mbar/s in 10 minutes. |

Requirement 4 Precursor, reactant and gas delivery system

| Req. | Description |
| --- | --- |
| 4.1 (T) | System must be able to accommodate a minimum of 6 precursor delivery lines. 4 of these must be equipped with heater jackets and 1 of these with a N2 bubbler. |
| 4.2 | Precursor containers with heater jackets should be capable of being heated to 200°C. |
| 4.3 | System must be able to have a retrofit available for a future O3 generator. |
| 4.4 | System must allow for precursor switching when other processes are done in the future, including pump/purge cleaning of the delivery lines. |
| 4.5 | System must be able to accommodate 4 mass-flow controlled gas lines. Required process gasses are: - Oxygen (O2) - Hydrogen (H2) - Nitrogen (N2) - Argon (Ar) |
| 4.6 | System must have 2 additional spare gas line(s) for possible future use. |
| 4.7 | All gas lines fittings should be VCR. |
| 4.8 | All gas lines have stainless steel orbital welded joints. |

Requirement 5 Plasma system

| Req. | Description |
| --- | --- |
| 5.1 (T) | The reaction chamber should be equipped with a remote inductively coupled plasma source or remote hollow cathode source. |
| 5.2 | Plasma generator should operate at 13.56 MHz and have an automatic matching box. |
| 5.3 (T) | Plasma source must have a strategy in place to minimize deposition on the source during a deposition. |

Requirement 6 Load-lock

| Req. | Description |
| --- | --- |
| 6.1 | System must be equipped with a load-lock system to handle all substrate sizes outlined in requirement 1.3. |
| 6.2 | Load-lock must be equipped with a small turbomolecular pump for improving pumping times and base pressure. |
| 6.3 | A valve must be present separating the load-lock from the process chamber. |
| 6.4 (T) | Chamber should be fitted with full range vacuum gauge (1×10-7 mBar to 1000 mBar) of a reputed manufacturer. |

Requirement 7 Heating and temperature control

| Req. | Description |
| --- | --- |
| 7.1 | System must be able to operate at a maximum process temperature of 400 °C. |
| 7.2 (T) | Process temperature control must be accurate within 1 °C. Electrode temperature should be monitored and controlled during the process. |
| 7.3 | Temperature uniformity should be ±1 °C across a 100 mm Si wafer in the temperature range between 20 °C and 400 °C during a process. |
| 7.4 | Reactor must be equipped with hot walls of 150 °C or higher. |

Requirement 8 Software

| Req. | Description |
| --- | --- |
| 8.1 | System must come with a Windows operating system. |
| 8.2 (T) | Software must provide a form of manual operation to test functions and operations individually, with the aim of troubleshoot issues. This mode must be hidden or locked from normal users but available to our support staff. |
| 8.3 | A back-up of all hard drives and PLCs must be supplied on a data carrier. |
| 8.4 | Software interface must be duplicated to allow for access from the cleanroom as well as the adjacent grey-room (technical space). |
| 8.5 | PLC for controlling hardware, guaranteed timing and safety functions. |
| 8.6 (T) | System operation should be completely controlled through a PC interface, including but not restricted to operation and control of the vacuum pumps and valves, gas flows, power supplies, heater temperature, substrate handling and process control. |
| 8.7 (T) | Process with predefined parameters recipes should be implemented, enabling fully automated process cycles. |
| 8.8 | All relevant system and process parameters should be logged. |

Requirement 9 Service, maintenance and warranty

| Req. | Description |
| --- | --- |
| 9.1 (T) | All equipment should be guaranteed for trouble-free operation and against manufacturing defects for a minimum period of one year after site acceptance test. Sufficient amount of rapidly wearing and consumable parts are to be included with the system to cover the warranty period. A new equivalent warranty period shall apply to all the products delivered by way of replacement. If applicable, the warranty period shall be extended by the period during which the product could not be used due to damage. The Contracting Authority may demand from the supplier that he replaces the products at their expense within the set term, if they show defects which do not permit normal use in accordance with the terms of the contract, for a period of one year counting from the date of provisional acceptance at the place of delivery. |
| 9.2 | After the initial warranty period of at least one year, the service contract must cover four additional years. This contract should include 1 preventive maintenance visit per year to ensure smooth operation during this period, covering all necessary parts, personnel, and travel expenses. Minimum response time must be 72 hours. The yearly cost for the maintenance contract, excluding VAT, should be specified and included in the Total Price. Following the initial 1+4 year term, the Maintenance Service Contract may be extended. The scope and terms of the subsequent maintenance package may be subject to adjustment, including potential increases or decreases in content. |
| 9.3 | The following documentation should be provided with the system as electronic pdf files, in English language: installation Manual, operation Manual, maintenance Manual, electrical diagrams, pneumatic diagrams, vacuum diagrams, gas line diagrams. |
| 9.4 | Vendor should guarantee the availability of spare parts, software and service support for a period of at least 10 years. |
| 9.5 | Technical help desk support should be available during normal local office working hours of the service team. |
| 9.6 | The language of all support (visits and technical help desk) should be English. |
| 9.7 | System must be able to be cleaned by through the use of liners or ability to swap the process chamber. |

Requirement 10 General properties and conditions

| Req. | Description |
| --- | --- |
| 10.1 | The system should be ISO 6/7 clean room compatible. |
| 10.2 | The system should be new and not used or refurbished. |
| 10.3 | The system should comply with CE marking standards. |
| 10.4 (T) | The system should be interlocked to guarantee safe operation and to prevent damage to the system or parts of the system in normal user mode. |
| 10.5 | All cables, gas lines or vacuum lines should be long enough to make the connections between components indicated in the supplied ground plan |
| 10.6 | The system must be delivered within a period of 9 calendar months counting from the day following the signing of the contract |
| 10.7 | The vendor should provide all pre-installation requirements. |
| 10.8 | Cooling water required should be compatible with the existing system of 6 bar 21 °C using DOWCAL 100 diluted to 20-25% in DI water as cooling liquid. |
| 10.9 | The system should comply with all relevant safety regulations. |
| 10.10 | Manufacturer must have previous experience in building PE-ALD systems. |
| 10.11 | On-site installation and training must be included in the total system price. |

Requirement 11 System price

| Req. | Description |
| --- | --- |
| 11.1 | The Total Price of the system should not exceed € 550,000 excluding VAT. |

# 3	Documentation

| Requirement | Document title | Page | Paragraph |
| --- | --- | --- | --- |
| 1.1 |  |  |  |
| 1.5 |  |  |  |
| 2.1 |  |  |  |
| 2.2 |  |  |  |
| 2.3 |  |  |  |
| 2.4 |  |  |  |
| 3.1 |  |  |  |
| 3.2 |  |  |  |
| 3.3 |  |  |  |
| 4.1 |  |  |  |
| 5.1 |  |  |  |
| 5.3 |  |  |  |
| 6.4 |  |  |  |
| 7.2 |  |  |  |
| 8.2 |  |  |  |
| 8.6 |  |  |  |
| 8.7 |  |  |  |
| 9.1 |  |  |  |
| 10.4 |  |  |  |