## Form
Appendix Question Form For the public European tender procedure 'Precipitation gauges and 3rd line maintenance'
Tenderer is asked to submit questions with the use of this question form.
Instruction of use: Column A states the question number. Column B states the document that the question regards, and if possible, please also state the paragraph. Column C states the question.
Question number | Document | Question | Answers
1 | Descriptive document | Do I understand correctly that the supply of a windshield is not part of this tender? | It is true that a windscreen is not part of the standard offer: when the diameter of the offered precipitation gauge's collection orifice area is larger than the diameter that is suitable for the current windscreen, then it is required to offer suitable windscreens. The height restriction is due to the English pit set-up, not due to the windscreen set-up. So even then the height restriction still applies. The correct heigt is dependant on the set-up: with english pit the top of the gauge should be flush with the wall, with windscreen the top of the gauge should be flush with the top of the windscreen.
2 | Annex 4 Program of Requirements final version 2 | A number of specifications can only be met by the addition of a data processing unit. This could be a data logger, for example. How is the data received by the KNMI from the measurement locations? Is data transmission part of the tender and what requirements does this data transmission have to meet? Protocol, data format, etc. Or does KNMI itself provide this data transmission? | The precipitation gauge will be connected to a data logger provided by KNMI. The specifications for communication between precipitation gauge and data logger are described in the Program of Requirements.
3 | Annex 4 Program of Requirements final version 2 Question 7, 34, 37-45 | There is a number of mentions of a data telegram. Do we understand correctly that a telegram has been offered that provides the requested information? If that is the case, can you give us an exact description of what such a telegram should look like and how it should be delivered? | A datalogger polls the gauge and the gauge answers by sending a telegram. The datalogger processes this telegram into a standardized output which will be send to KNMI's database. Any output of the precipitation gauge is acceptable as long as the datalogger can be programmed to process it in any way we need.
4 | Annex 4 Program of Requirements final version 2 Question 7, 34, 37-45 | The SDI-12 protocol works more or less with telegrams. Will we meet these specifications when we use the SDI-12 interface protocol. | Yes you will meet these specifications.
5 | Annex 4 Program of Requirements final version 2 | Tenderer should provide the possibility to reprocess archived "raw" data offline by either Tenderer or KNMI in order to optimize the algorithms that derive the meteorological quantities. What is meant by this? Where should the raw data be stored? In the rain gauge, in a data logger or in the supplier's cloud? Or should it be offered with an extra message so that KNMI can store this data itself? | A rain gauge may not measure precipitation intensity, detection and accumulation directly, but may measure for example the weight of collected precipitation, the water level of collected precipitation or the time when tips of a tipping bucket occur. This “raw” data might be filtered, processed and converted to get an optimized intensity, detection and accumulation. Since this is about a desirable, the requirement asks you to indicate whether “raw” data and applied processing steps will be made available to KNMI. The ”raw” data can be made available via the sensor telegram that also contains the precipitation intensity, detection and accumulation and the sensor status or via a separate telegram. Data storage is not applicable for your offer.
6 | Annex 4 Program of Requirements final version 2 | We think the 118 questions is a lot. Why so extensive. KNMI has an extensive test field where many instruments have already been extensively tested, so many of these questions can already be answered by KNMI. In addition, we believe that the large number of questions will not provide the clarity that is expected. Many questions are unclear and leave a lot of room for free interpretation, so that the answers raise new questions. | You mention the 118 questions. We assume you are referring to the 118 requirements. We disagree with your opinion. It's correct that we have tested a few different types of precipitation gauges available on the market to see how they perform and to see if there are suitable replacements for our current gauge. We can not use that data for verification purposes, because the testing was done before the tender was issued. The KNMI is required to tender new gauges when the investment exceeds a certain value. To guarantee that the tender produces a gauge that fits what the KNMI needs, a stakeholder analysis was completed. The number of mandatory requirements and desirables stems from the complete package of what the stakeholders need.
7 | Time Schedule | The schedule coincides with the holiday period. Also taking into account the complexity of the number of questions, we would like to request that the period of submission and everything thereafter be extended by 1 month. So submitted before September 14. | Based on the information we currently have, we are not planning to move the deadline. Contracting Authority has tried to answer the questions from the first round of questions as soon as possible. The deadline for submitting questions for the second and final round of questions is set for July 18, 2024. Final publication of the Information Memorandum will take place before Friday, August 2, 2024 as you can see in the schedule.
8 | Annex 4 Program of Requirements final version 2 | is it correct that the delivery of a windshield is not part of the standard delivery unless the specifications cannot be met? Is the height still a "lock-out" spec. or is it allowed to offer | We refer to the answer on question 1
9 | Descriptive document Annex 1B | Is it correct that 2 systems are initially ordered and that they are tested for 12 months? What is the desired delivery time for these 2 systems and what is the desired delivery time after the test is completed? | We'd like you to offer one type of gauge. When this gauge cannot meet the duration requirements, include an auxiliary device for this purpose that can meet the duration requirements. Initially we'd like you to deliver us 9 pieces of that gauge or solution (gauge + auxiliary device) that we will use for testing. For more details on delivery times and testing, see section 2 and paragraph 4.6 of Annex 1B Verification phase and delivery.
10 | Draft agreement par 6 | MTBF of 60 months and a 99.98 % data availability is tough. Who will judge the cause of the failure and who is responsible for this failure. | The threshold for data availablility has been determined by experience with our current gauge and other sensors in our automated network. KNMI experts, and when needed in consultation with the Tenderer, will determine the cause of the failure and dertermine the party that is responsible.
11 | Annex 4 Program of Requirements final version 2 | I know a lot of waterboards send the data they collect with their rain gauges to the KNMI server. They use the KNMI protocol to send the data. Will the rain gauges prurchased by KNMI work with same protocol and is it allowed to transmit data from the rain gauges to our cloud server first and then transmit the data in the same way waterboards are doing now. | No it's not allowed to transfer the data in the same way that waterboards are doing now. For more information about data transmission we refer to section 2.3 of the Program of Requirements.
12 | notice from the contracting authority | Closing date for the second round of questions. | The deadline for submitting questions for the second and final round of questions is set for July 18, 2024. Final publication of the Information Memorandum will take place before Friday, August 2, 2024 as you can see in the updated schedule on Tenderned.
13
14
15
16
17
18
19
20
21
22
23
24
25
26
27
28
29
30
31
32
33
34
35
36
37
38
39
40
41
42
43
44
45
46
47
48