## Instructions and Explanations
Instructions regarding the Programme of Requirements and Desirables - New generation wind sensor and 3rd line maintenance – 3D sonic wind sensors
Introduction
This Program of Requirements and Desirables forms an integral part of the tender for new wind sensors, TN-463889. It details the technical Requirements and Desirables of the new wind sensors that the KNMI intends to buy. Every Requirement and Desirable is of a specific type, where the possible types are listed below in the Requirement and Desirable Types section. Every type has a specific answer format that must be adhered to. Some types are associated with a substantiation that will be given a score by the assessment team. Every Desirable can receive a score between 0 and 5. This score will be multiplied by the weight of the particular Desirable to yield the points awarded to the Tenderer for that specific Desirable. For more information on the award criteria, see the Descriptive Document. In general, the assessment is done using the guidelines described in the Descriptive Document. Additional focus points are mentioned in the Evaluation Criteria next to each Desirable, indicating the specific aspects that are of interest to the KNMI assessment team.
Fill-in instructions IMPORTANT INFORMATION
This document also functions as the formal declaration of the Tenderer to comply with the Requirements and to indicate in what extent the Tenderer complies with the Desirables. For this purpose, fill in the Requirements and Desirables sheet of this excel file for all yellow cells. These cells typically contain a drop-down menu that can be used to select the proper option. Tenderer may not change any other cells within this excel file even if technically possible. For the purpose of answering the open Desirables with a substantiation, Annex 08 is provided with a numbering of the Desirables that require substantiations. Tenderer may use this annex to structure the substantiations. Tenderer may also use an own template or file to structure the substantiations, as long as it is unambiguously clear which substantiation belongs to which Desirable. In all cases, the text in this Program of Requirements is leading for all Requirements. The amount of space available for the substantiations is given in terms of A4 pages. This is done deliberately to allow for the inclusion of figures & graphs which tenderer may use to support the substantiation. The text in these substantiations should be in "Calibri" font and must be of font size 11pt or larger. Any text within graphics must be of similar (or better) readability. Any substantiation text or graphics which are beyond the allowed size may be ignored during the assessment.
Requirement and Desirable Types
Every Requirement and Desirable in this Program of Requirements and Desirables is of a specific type. These types are listed here and includes an explanation of what is expected for each type.
Type | Sub-type | Abbreviation | Desired answer | Explanation | Scoring
Mandatory | Closed | M C | Yes or No | A singular Yes or No answer is requested, without any extra information. | n/a
Mandatory | Information | M I | Yes or No with extra information | A singular Yes or No answer is requested that is supplemented with particular objective information. | n/a, the supplementary information is not used in the scoring.
Desirable | Closed | D C | Yes or No | A singular Yes or No answer is requested, without any extra information. | Score 0 for No. Score 5 for Yes.
Desirable | Scaled | D S | Information | A specific piece of objective information is requested (such as a number or list). | Scoring according to the criteria given.
Desirable | Open | D O | Substantiation | A substantiation for the Desirable is requested, in accordance with the guidelines given. | Scoring by assessment team based on Assessment table (see below), and the specific criteria listed. Score of 0 to 5 is given.
Assessment levels
Level | Description |  |  |  | Score
Excellent | In the opinion of the assessment team, the Tenderer has addressed the requested Desirable using the assessment aspects and specific criteria in a highly complete, substantiative, appropriate, highly convincing and excellent way and/or complies to the full level or beyond with the wishes of the Contracting Authority as stated in the Desirable. Tenderer additionally adds new insights and/or additional value that the assessment team did not expect. |  |  |  | 5
Very good | In the opinion of the assessment team, Tenderer has addressed the requested Desirable using the assessment aspects and specific criteria in a complete, substantiated and very convincing way and/or complies to nearly the full level of the wishes of the Contracting Authority as stated in the Desirable. |  |  |  | 4
Good | In the opinion of the assessment team, Tenderer has addressed the requested Desirable using the assessment aspects and specific criteria in a relevant, appropriate, convincing and good way and/or complies mostly to the level of the wishes of the Contracting Authority as stated in the Desirable. |  |  |  | 3
Sufficient | In the opinion of the assessment team, Tenderer has addressed the requested Desirable using the assessment aspects and specific criteria in an only somewhat relevant, reasonably convincing and sufficient way, but without thorough elaboration and/or taking the assessment aspects and specific criteria into proper account. And/or the Tenderer only complies somewhat to the level of the wishes of the Contracting Authority as stated in the Desirable. |  |  |  | 2
Poor | In the opinion of the assessment team, the Tenderer has addressed the requested Desirable in a subpar way and to a very limited extend without sufficient elaboration and without taking into account the assessment aspects and specific criteria. And/or the Tenderer only complies to a very limited level with the wishes of the Contracting Authority as stated in the Desirable. |  |  |  | 1
Very poor | In the opinion of the assessment team, the Tenderer does not respond to the requested Desirable in a meaningful manner taking into account the assessment aspects and/or does not comply with the wishes of the Contracting Authority as stated in the Desirable. |  |  |  | 0

## Requirements and Desirables
3D Sonic Wind Sensors - Mandatory and Desirable Requirements | [6 empty] | Notice: Not all mandatories are completed.
# | M/D | Description | Score Type | Weight | Evaluation Criteria | Answer | Outcome | Points
1 Functional Requirements
1.1 General Requirements
1 | <M> | The wind sensor shall be capable of measuring and reporting the vertical wind speed (in SI units, e.g. m/s or cm/s) and the horizontal: • Wind direction (where the wind is coming from, so where 0 or 360° is true north, 90° east and so forth; mean shall be the average of unit vectors); • Wind speed (the magnitude of the instantaneous horizontal wind vector and scalar averages of the instantaneous magnitudes) in SI units. | M C | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
2 | <D> | Substantiation of any additional capabilities and any improved performance of the wind sensor above and beyond what is covered by the other Mandatory and Desirable Requirements in this Tender is desired. These should be described as specifically as possible, with special attention to the added value to the intended purpose of KNMI (see section 1.1 of the Descriptive Document). | D O | 7 | A score will be given between 0 and a maximum of 5. Please provide a substantiation of no more than 1 A4 page. The substantiation will be assessed in accordance with the general assessment criteria of the Descriptive Document. The assessment is based on expert judgement with special attention given to the following assessment criteria: - Higher score for better performing wind sensors in general, e.g. smaller bias, less failures. - Additional capabilities beyond those described in the rest of these Requirements will result in a higher score. - Note that no points will be given for any additional capabilities that do not have a proven positive impact with respect to the intended purpose. | Select applicable option: | The score will be assigned by the KNMI assessment team. | #VALUE!
3 | <M> | The internal sampling and output frequency of complete basic measuring sequences at 12 Hz shall be possible. | M C | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
4 | <D> | An internal sampling rate higher than 12 Hz is desired. | D S | 8 | Points are given according to the maximum internal sampling rate of the sensor: ≤ 12 Hz = score 0. 12-90 Hz = score 1. 90-160 Hz = score 2. 160-220 Hz = score 3. 220-270 Hz = score 4. > 270 Hz = score 5. | Select the actual sampling rate: | Score | #VALUE!
1.2 Wind Direction Requirements
5 | <M> | The maximum absolute bias (negative biases made positive) of the wind direction in a wind tunnel shall be < 3° for all wind directions and tunnel wind speeds in the range 1 to 65 m/s. To make the dependency of the bias on wind direction clear, the wind direction bins will be 5° or smaller. | M C | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
6 | <D> | The maximum absolute bias (negative biases made positive) of the wind direction in a wind tunnel is desired to be: • < 3° for all wind directions and tunnel wind speeds in the range 1 to 65 m/s. • < 5° for all wind directions and tunnel wind speeds in the range 65 to 90 m/s. In addition, good reproducibility of the bias from one sensor to the next is desired where small differences between the sensors is good. Substantiation in the form of wind tunnel test results is desired. The results should cover the entire wind speed range and all wind directions for at least three sensors, each from different production batches. To make the change in the bias with wind direction clear, the wind direction bins will be 5° or smaller. | D O | 8 | A score will be given between 0 and a maximum of 5. Please provide a substantiation of no more than 1 A4 page. The substantiation will be assessed in accordance with the general assessment criteria of the Descriptive Document. The assessment is based on expert judgement with special attention given to the following assessment criteria: - The following guidelines for the scoring are used, where low means ≤ 65 m/s and high > 65 m/s: • 5 if maximum bias < 1° (low) and < 3° (high) • 3 if maximum bias < 2° (low) and < 4° (high) • 0 if maximum bias < 3° (low) and < 5° (high) - In addition to this, a higher score can be achieved by the reproducibility (measured using the absolute difference between the bias of two sensors averaged over all combinations of the three sensors and all wind directions and tunnel speeds) as follows: • 0 more if the score for the maximum bias was 5 • at most 1 more if the score for the maximum bias was 3 • at most 2 more if the score for the maximum bias was 0 A larger portion of the maximum additional score will be awarded if the absolute difference is low. How much larger will depend on the absolute difference shown by the wind tunnel test results and can therefore not be decided before the results have been received. | Select applicable option: | The score will be assigned by the KNMI assessment team. | #VALUE!
7 | <M> | The absolute bias (negative bias made positive) of the wind direction in a wind tunnel, averaged over all wind directions, shall be ≤ 1.5° for each and every tunnel wind speed in the range 1 to 65 m/s. The wind direction bins will be 5° or smaller. | M C | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
8 | <M> | The maximum absolute bias (negative biases made positive) of the horizontal wind direction of the sensor, measured in a wind tunnel with the sensor inclined up to and including ±30° from the vertical position (both into the wind as well as away from the wind, in steps of 10° or less), shall be ≤ 2° for all wind directions and tunnel wind speeds in the range 1-16 m/s. To make the change in the bias with wind direction clear, the wind direction bins will be 5° or smaller. | M C | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
9 | <D> | The maximum absolute bias (negative biases made positive) of the horizontal wind direction of the sensor is desired to be < 2°, where the bias is measured in a wind tunnel with the wind sensor inclined (tilted) up to and including ±30° from the vertical position (both into the wind as well as away from the wind), for tunnel wind speeds 1-16 m/s and for all directions. Also, with the wind sensor inclined from +30° (wind approaching the sensor from "below") to -45° (wind from "above") an absolute bias < 15° is desired, for speeds 1-90 m/s and for all wind directions. In addition, as good (small differences between the sensors is good) a reproducibility of the bias, from one sensor to the next, is desired. Substantiation in the form of wind tunnel test results is desired, where the results cover all wind directions and all wind speeds in the range 1-90 m/s with the wind sensor inclined from +30° (wind approaching the sensor from "below") to -45° (wind from "above"). The test results should be for at least three sensors, each from different production batches. To make the change in the bias with wind direction clear, the wind direction bins will be 5° or smaller. | D O | 5 | A score will be given between 0 and a maximum of 5. Please provide a substantiation of no more than 1 A4 page. The substantiation will be assessed in accordance with the general assessment criteria of the Descriptive Document. The assessment is based on expert judgement with special attention given to the following assessment criteria: - The following guidelines for the scoring are used, where low means 1-16 m/s (and inclination up to and including ±30°) and high means 1-90 m/s (and inclinations +30 to -45°): • 5 if maximum bias < 1.0° (low) and < 9.0° (high) • 3 if maximum bias < 1.5° (low) and < 12.0° (high) • 0 if maximum bias < 2.0° (low) and < 15.0° (high) - In addition to this, a higher score can be achieved by the reproducibility (measured using the absolute difference between the bias of two sensors averaged over all combinations of the three sensors, all inclinations in the range between +30 and -45°, and all wind directions and tunnel speeds) as follows: • 0 more if the score for the maximum bias was 5 • at most 1 more if the score for the maximum bias was 3 • at most 2 more if the score for the maximum bias was 0 A larger portion of the maximum additional score will be awarded if the absolute difference is low. How much larger will depend on the absolute difference shown by the wind tunnel test results and can therefore not be decided before the results have been received. | Select applicable option: | The score will be assigned by the KNMI assessment team. | #VALUE!
10 | <M> | The absolute bias (negative biases made postive) of the horizontal wind direction of the sensor, measured in a wind tunnel and averaged over all wind directions and inclination angles, shall be < 1.0° for each and every tunnel wind speed in the range 1-16m/s and ≤ 4° in the range 16-65 m/s. The inclination or tilting of the sensor from the vertical position shall cover the range ±30° in steps of 10° or less, both into the wind as well as away from the wind. The wind direction bins will be 5° or smaller. | M C | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
11 | <M> | The wind direction measured by different individual sensors of the same model in a wind tunnel shall be reproducible to the extent that the absolute difference (wind direction measurement of the one sensor minus that of the other where negative differences are made positive) shall be < 1.5° for each and every combination of tunnel wind speed (in the range 1-65 m/s) and wind direction (bin width 5° or less). | M C | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
12 | <M> | The resolution of the wind direction (both internal and output) shall be ≤ 0.1°. | M C | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
13 | <M> | The threshold for reporting wind direction shall be ≤ 0.1 m/s for each and every wind direction. | M C | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
1.3 Wind Speed Requirements
14 | <M> | The range of the wind speed shall be at least 0-65 m/s. | M C | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
15 | <D> | The range of the wind speed is desired to be 0-90 m/s. | D S | 8 | Points are given according to the wind range of the sensor: 0-65 m/s = score 0. At least 0-70 m/s = score 1. At least 0-75 m/s = score 2. At least 0-80 m/s = score 3. At least 0-85 m/s = score 4. At least 0-90 m/s = score 5. | Select the applicable wind speed range: | Score | #VALUE!
16 | <M> | The wind direction dependent absolute bias (negative biases made positive) of the wind speed measured by the sensor in a wind tunnel for each and every wind direction shall be: • ≤ 0.4 m/s for each tunnel wind speed in the range 1-10 m/s • and ≤ 7% for each tunnel wind speed ≥ 10 m/s up to 65 m/s. To make the change in the bias with wind direction clear, the wind direction bins will be 5° or smaller. | M C | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
17 | <D> | The maximum absolute bias (negative biases made positive) of the wind speed in a wind tunnel is desired to be: • < 0.4 m/s for all wind directions and tunnel wind speeds in the range 1-10 m/s. • < 8% for all wind directions and tunnel wind speeds ≥ 10 m/s up to 90 m/s. In addition, reproducibility of the wind-direction dependency of the wind speed bias, from one sensor to the next, is desired for horizontal wind speeds in the range 1-90 m/s. The reproducibility of the bias from one sensor to the next should be as good (small differences between the sensors is good) as possible. This should be substantiated by providing wind tunnel test results that cover the entire wind speed range and all wind directions for at least three sensors, each from different production batches. To make the change in the bias with wind direction clear, the wind direction bins will be 5° or smaller. | D O | 8 | A score will be given between 0 and a maximum of 5. Please provide a substantiation of no more than 1 A4 page. The substantiation will be assessed in accordance with the general assessment criteria of the Descriptive Document. The assessment is based on expert judgement with special attention given to the following assessment criteria: - The following guidelines for the scoring are used, where low means 1-10 m/s and high 10-90 m/s: • 5 if the maximum bias < 0.2 m/s (low) and < 6% (high) • 3 if the maximum bias < 0.3 m/s (low) and < 7% (high) • 0 if the maximum bias < 0.4 m/s (low) and < 8% (high) - In addition to this, a higher score can be achieved by the reproducibility (measured using the absolute difference between the bias of 2 sensors averaged over all combinations of the 3 sensors and all wind directions and tunnel speeds) as follows: • 0 more if the score for the maximum bias was 5 • at most 1 more if the score for the maximum bias was 3 • at most 2 more if the score for the maximum bias was 0 A larger portion of the maximum additional score will be awarded if the absolute difference is low. How much larger will depend on the absolute difference shown by the wind tunnel test results and can therefore not be decided before the results have been received. | Select applicable option: | The score will be assigned by the KNMI assessment team. | #VALUE!
18 | <M> | The absolute bias (negative bias made positive) of the horizontal wind speed measured in a wind tunnel, averaged over all wind directions, shall be ≤ 2% for each and every tunnel wind speed in the range 5 to 65 m/s and < 0.1 m/s for tunnel speeds < 5 m/s. The wind direction bins will be 5° or smaller. | M C | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
19 | <D> | A Class D classification index < 6.0 is desired, with heating off and with all parameters influencing the result where the Class D parameters are defined in EN-IEC 61400-50-1:2022. | D C | 6 | Yes = Score of 5 No = Score of 0 | Select your answer: | No answer selected. | #VALUE!
20 | <M> | The wind speed measured by different individual sensors of the same model in a wind tunnel shall be reproducible to the extent that the absolute difference (wind speed measurement of the one sensor minus that of the other where negative differences are made positive) shall be: • < 0.1 m/s for tunnel wind speeds in the range 1-10 m/s • and < 3.5% for tunnel wind speeds ≥10 m/s and below 65 m/s. This applies for each and every combination of wind speed and wind direction. The difference shall be averaged over wind direction bins with a width of 5° or less. | M C | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
21 | <M> | The wind direction dependent absolute bias (negative biases made positive) of both the horizontal and the vertical wind speed measured by the sensor in a wind tunnel with the sensor inclined up to and including 30° from the vertical position shall be ≤ 8% for each and every combination of wind direction, tunnel wind speed in the range 4-16 m/s and inclination both into the wind as well as away from the wind. To make the change in the bias with wind direction clear, the wind direction bins will be 5° or smaller. | M C | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
22 | <D> | The wind direction dependent absolute bias (negative biases made positive) of both the horizontal and the vertical wind speed is desired to be: • < 8% for each and every combination of inclinations of the sensor (away from the vertical stand) up to and including ±30° with wind tunnel wind speeds 4-16 m/s and with all wind directions • and < 15% for each and every combination of inclinations between -30° and -45° (wind from "above") with tunnel speeds 5-16 m/s and with all wind directions. • and < 15% for each and every combination of inclinations between +30° (wind approaching the sensor from "below") and -45° (wind from "above") with tunnel speeds 16-90 m/s and with all wind directions. In addition, a good reproducibility of the bias from one sensor to the next is desirable, where small differences between the sensors is good. This should be substantiated by providing wind tunnel test results that cover all wind directions and all tunnel speeds in the range 1-90 m/s with the wind sensor inclined from +30° to -45°. The test results should be for at least three sensors, each from different production batches. To make the change in the bias with wind direction clear, the wind direction bins will be 5° or smaller. | D O | 16 | A score will be given between 0 and a maximum of 5. Please provide a substantiation of no more than 2 A4 pages. The substantiation will be assessed in accordance with the general assessment criteria of the Descriptive Document. The assessment is based on expert judgement with special attention given to the following assessment criteria: - The maximum absolute bias (considering both horizontal and vertical speed measurements) should be as low as possible. The following guidelines for the scoring are used, where low means 4-16 m/s (and inclination up to and including ±30°) and high means the other two categories of comprising both low speed (5-16 m/s) and high inclination (-30° to -45°) and high speed 16-90 m/s (and inclinations +30 to -45°): • 5 if maximum bias < 6% (low) and < 11% (high) • 3 if maximum bias < 7% (low) and < 15% (high) • 0 if maximum bias < 8% (low) and < 17% (high) - In addition to this, a higher score can be achieved by the reproducibility (measured using the absolute difference between the bias of two sensors averaged over the horizontal and vertical speed biases for all combinations of the three sensors, all inclinations in the range between +30 and -45°, and all wind directions and tunnel speeds) as follows: • 0 more if the score for the maximum bias was 5 • at most 1 more if the score for the maximum bias was 3 • at most 2 more if the score for the maximum bias was 0 A larger portion of the maximum additional score will be awarded if the absolute difference is low. How much larger will depend on the absolute difference shown by the wind tunnel test results and can therefore not be decided before the results have been received. | Select applicable option: | The score will be assigned by the KNMI assessment team. | #VALUE!
23 | <M> | The absolute bias (negative biases made positive) of the horizontal wind speed of the sensor, measured in a wind tunnel and averaged over all wind directions and inclination angles up to and including 20° (in steps of 10° or less; inclined both into the wind as well as away from the wind), shall be < 2% for each and every tunnel wind speed in the range 4-16 m/s. Also, for speeds 16-65 m/s the absolute bias, averaged over all wind directions and inclination angles up to and including 30°, shall be < 8%. The wind direction bins will be 5° or smaller. | M C | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
24 | <M> | The resolution of the wind speed (both internal and output) shall be ≤ 0.01 m/s. | M C | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
25 | <M> | The threshold of reporting wind speed shall be ≤ 0.01 m/s. | M C | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
26 | <D> | A small negative effect on measurements caused by precipitiation is desired, e.g. a small number of missing or incorrect measurements. Tenderer should substantiate this effect, ideally by including field and/or wind tunnel test results involving various types of precipitation (the complete range including drizzle, rain, snow, hail) and intensities (ranging from 0.1 to 70 mm/h). The sensor should continue to meet the mandatory requirements during the type and intensities of precipitation that can be expected at the KNMI measurement locations. | D O | 6 | A score will be given between 0 and a maximum of 5. Please provide a substantiation of no more than 1 A4 page. The substantiation will be assessed in accordance with the general assessment criteria of the Descriptive Document. The assessment is based on expert judgement with special attention given to the following assessment criteria: - The less the measurements are affected by precipitation, the higher the score. - The score will be adjusted by a factor between 0 and 1 to take account of the completeness of the test results, where 1 will be used for the most complete (many measurement points that are spread over the required ranges) set of results and 0 if tenderer choses not to provide any test results. | Select applicable option: | The score will be assigned by the KNMI assessment team. | #VALUE!
27 | <D> | A good reproducibility of the horizontal wind speed measured by different individual sensors of the same model is desired. For each combination of wind direction (for bin width 5° or less) and tunnel wind speeds in the given range, the absolute difference (negative differences made positive) of the wind speed measurement of the one sensor minus that of the other (at the same wind tunnel speed) is desired to be: • ≤ 0.05 m/s for the tunnel speed range 1-5 m/s • ≤ 1% for the tunnel speed range 5-25 m/s. | D C | 3 | Yes = Score of 5 No = Score of 0 | Select your answer: | No answer selected. | #VALUE!
28 | <M> | The range of the vertical wind speed shall be at least -50 m/s to +50 m/s. | M C | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
29 | <M> | The maximum absolute bias (negative biases made positive) of the vertical wind speed measured in a wind tunnel for all wind directions (bin width 5° or less) shall be: • ≤ 0.4 m/s for tunnel horizontal wind speeds 1-10 m/s • and ≤ 5% of tunnel horizontal wind speeds ≥ 10 m/s and ≤ 16 m/s. The vertical wind speed measured should ideally be 0 m/s as it is assumed that the sensor is mounted vertically in the wind tunnel. | M C | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
30 | <D> | A maximum absolute bias (negative biases made positive) of the vertical wind speed, when the sensor is mounted vertically in a wind tunnel (vertical speed is ideally 0 m/s), is desired to be: • < 0.4 m/s in the tunnel horizontal wind speed range 1-10 m/s • and < 8% of the tunnel horizontal wind speed in the range 10-90 m/s. In addition, a good reproducibility of the bias from one sensor to the next is desirable, where small differences between the sensors is good. This should be substantiated by providing wind tunnel test results that cover the wind speed range (1-90 m/s) and all wind directions for at least three sensors, each from different production batches. To make the change in the bias with wind direction clear, the wind direction bins will be 5° or smaller. | D O | 11 | A score will be given between 0 and a maximum of 5. Please provide a substantiation of no more than 2 A4 pages. The substantiation will be assessed in accordance with the general assessment criteria of the Descriptive Document. The assessment is based on expert judgement with special attention given to the following assessment criteria: - The following guidelines for the scoring are used, where low means 1-10 m/s and high 10-90 m/s: • 5 if the maximum bias is < 0.2 m/s (low) and < 6% (high) • 3 if the maximum bias is < 0.3 m/s (low) and < 7% (high) • 0 if the maximum bias is < 0.4 m/s (low) and < 8% (high) - In addition to this, a higher score can be achieved by the reproducibility (measured using the absolute difference between the bias of two sensors averaged over all combinations of the three sensors and all wind directions and tunnel speeds) as follows: • 0 more if the score for the maximum bias was 5 • at most 1 more if the score for the maximum bias was 3 • at most 2 more if the score for the maximum bias was 0 A larger portion of the maximum additional score will be awarded if the absolute difference is low. How much larger will depend on the absolute difference shown by the wind tunnel test results and can therefore not be decided before the results have been received. | Select applicable option: | The score will be assigned by the KNMI assessment team. | #VALUE!
31 | <M> | The resolution of the vertical wind speed (both internal and output) shall be better than or equal to 0.01 m/s. | M C | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
32 | <M> | The threshold of reporting vertical wind speed shall be better than or equal to 0.01 m/s. | M C | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
1.4 Sonic Temperature Requirements
33 | <M> | The sensor shall report the sonic temperature in SI units. | M C | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
34 | <M> | The range of the sonic temperature reportable by the Sensor shall be at least -25°C to +50°C. | M C | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
35 | <M> | The measurement uncertainty (root mean square error) of the sonic temperature shall be ≤ 2°C when taken over the entire range of wind directions. This should be the case for each and every combination of temperature and wind speed in the measurement range of the sensor. | M C | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
36 | <D> | It is desired to have a measurement uncertainty (root mean square error) of the sonic temperature of < 2°C, when taken over the entire range of wind directions. This should be the case for each and every combination of temperature and wind speed in the measurement range of the sensor. | D O | 6 | Please provide a substantiation of maximum 1 A4 page. The substantiation will be assessed and a score is given between 0 to 5 maximum. The assessment is based on expert judgement with special attention given to the following assessment criteria: - Matching of the sonic temperature to the reference temperature. - Presence of Calibration reports on this parameter is desirable, especially if the calibration is done by an EN-ISO 17025:2018-accredited laboratory. - Field trial comparisons with more traditional temperature sensors are desirable. - The level of detail with which the test results cover the ranges of temperature, wind speed and direction is important. Higher detail is desirable. - The highest attainable score depends on the measurement uncertainty: • 5 if the uncertainty is < 1.0°C • 3 if the uncertainty is < 1.5°C • 1 if the uncertainy is < 2.0°C. | Select applicable option: | The score will be assigned by the KNMI assessment team. | #VALUE!
37 | <D> | A measurement uncertainty (root mean square error) of the variance of the sonic temperature of ≤ 0.1°C is desired, when taken over the entire range of wind directions. This should be the case for each and every combination of temperature and wind speed in the measurement range of the sensor. Tenderer should substantiate this measurement uncertainty, and this substantiation should be representative of the climate conditions in the Netherlands, including North Sea platforms and the Carribean Netherlands. Notes: KNMI intends to use this variance to calculate the Obukhov length. [variance = average of (instantaneous measurement minus the average of many recent instantaneous measurements)^2] | D O | 8 | Please provide a substantiation of maximum 1 A4 page. The substantiation will be assessed and a score is given between 0 to 5 maximum. The assessment is based on expert judgement with special attention given to the following assessment criteria: - The higher the level of detail with which the test results cover the ranges of temperature, wind speed and direction, the better. - Matching of the sonic temperature to the reference temperature is desired to be as good as possible. - The highest attainable score depends on the measurement uncertainty: • 5 if the uncertainty is < 0.02°C • 3 if the uncertainty is < 0.06°C • 1 if the uncertainy is ≤ 0.10°C. | Select applicable option: | The score will be assigned by the KNMI assessment team. | #VALUE!
38 | <M> | The sonic temperature measured by different individual sensors of the same model shall be reproducible to the extent that the absolute difference (measurement of the one sensor minus that of the other where negative differences are made positive) shall be < 1.0°C for each and every combination of temperature, wind speed and direction that the sensor can measure. | M C | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
39 | <M> | The resolution of the sonic temperature (both internal and output) shall be ≤ 0.1°C. | M C | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
1.5 Environmental Requirements
40 | <M> | The offered Wind Sensors shall be suitable for use within the Netherlands, including North Sea platforms and the Caribbean Netherlands. | M C | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
-
2 Maintainability Requirements
2.1 Wind Speed Calibration Requirements
41 | <M> | The Tenderer shall include a calibration procedure that ensures the wind speed can be verified, adjusted and calibrated by KNMI staff in a wind tunnel. This calibration procedure must be clearly documented. Note: this calibration procedure (or page-specific reference to the manual chapter that contains this procedure) must be made available after initial award upon first request. | M C | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
42 | <D> | It is desired that Tenderer substantiates how the calibration and adjustment of the wind sensor can be used to optimize the wind speed measurement. In particular, access to correction tables for wind speed, including dependence on wind speed, wind direction and tilt angles, is highly desired. These correction tables should be adjustable by KNMI staff during calibration. It is highly desirable that the final measurement uncertainty after calibration and adjustment is at least a factor 3 better than the accuracy requirements listed in the mandatories in order to account for drift during the period in between two successive calibrations. Alternatively, Tenderer should substantiate why the design of the sensor is such that drift of the calibration is eliminated to a level where this factor 3 is no longer necessary. In addition, other factors such as traceability to a standard measurement reference are also very important. | D O | 8 | Please provide a substantiation of maximum 1 A4 page. The substantiation will be assessed and a score is given between 0 to 5 maximum. The assessment is based on expert judgement with special attention given to the following assessment criteria: - Correct traceability to a standard measurement reference. - Completeness of the verification and calibration procedure. - The final measurement uncertainty that can be achieved is important for the scoring. A lower measurement uncertainty after calibration will yield higher scores. - Ease of use to perform the calibration. | Select applicable option: | The score will be assigned by the KNMI assessment team. | #VALUE!
2.2 Wind Direction Calibration Requirements
43 | <M> | The Tenderer shall include a calibration procedure that ensures that the wind direction of the Wind Sensors can be verified, adjusted and calibrated by KNMI staff. This calibration procedure must be clearly documented. Note: this calibration procedure (or page-specific reference to the manual chapter that contains this procedure) must be made available after initial award upon first request. | M C | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
44 | <D> | It is desired that Tenderer substantiates how the calibration and adjustment of the wind sensor can be used to optimize the wind direction measurement. In particular, access to correction tables for wind direction, including dependence on wind speed, wind direction and tilt angles, is highly desired. These correction tables should be adjustable by KNMI staff during calibration. It is highly desirable that the final measurement uncertainty after calibration and adjustment is at least a factor 3 better than the accuracy requirements listed in the mandatories in order to account for drift during the period in between two successive calibrations. Alternatively, Tenderer should substantiate why the design of the sensor is such that drift of the calibration is eliminated to a level where this factor 3 is no longer necessary. In addition, other factors such as traceability to a standard measurement reference are also very important. | D O | 8 | Please provide a substantiation of maximum 1 A4 page. The substantiation will be assessed and a score is given between 0 to 5 maximum. The assessment is based on expert judgement with special attention given to the following assessment criteria: - Correct traceability to a standard measurement reference. - Completeness of the verification and calibration procedure. - The final measurement uncertainty that can be achieved is important for the scoring. A lower measurement uncertainty after calibration will yield higher scores. - Ease of use to perform the calibration. | Select applicable option: | The score will be assigned by the KNMI assessment team. | #VALUE!
2.3 Other Calibration Requirements
45 | <M> | The calibration interval of the Wind Sensor shall be 26 months or longer. | M C | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
46 | <M> | Tenderer shall commit to provide a clear verification protocol for the Wind Sensors and/or its modules directly after receiving an "Intention to Award" decision, and at the latest before the Final Award decision will be published. This protocol shall allow complete verification of the correct operation and the required quality of the Wind Sensor in the laboratory and on the test field by KNMI before operational deployment of the Wind Sensors. | M C | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
47 | <M> | The Tenderer shall supply a calibration certificate issued by an EN-ISO 17025:2018 accredited laboratory with each wind sensor. This calibration shall cover at least a representative number of samples of both wind speed and wind direction measurements. | M C | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
48 | <M> | The Tenderer shall commit to re-calibrate the wind sensors used in the verification phase after its conclusion, as detailed in Annex 07. This recalibration certificate must be issued by an EN-ISO 17025:2018 accredited laboratory. This recalibration shall cover at least the same set of samples for both wind speed and wind direction as is used in the initial calibration. | M C | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
49 | <D> | It is desired that Tenderer includes a calibration procedure to ensure that the sonic temperature as measured by the Wind Sensor can be verified, adjusted and calibrated by KNMI staff. This calibration procedure should be clearly documented. Note: KNMI has access to climate chambers where both temperature and relative humidity can be controlled. KNMI envisions that these chambers can be used for calibration of sonic temperature measurements. | D C | 3 | Note: the calibration procedure (or page-specific reference to the manual chapter that contains this procedure) must be made available after initial award upon first request. Yes = Score of 5 No = Score of 0 | Select your answer: | No answer selected. | #VALUE!
-
3 Technical Requirements
3.1 General Requirements
50 | <M> | A unique ID/serial number of the Wind Sensor, as well as the firmware version shall be available in the data message or is accessible through specific commands. | M C | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
51 | <M> | Commands shall be available to retrieve, change and store the parameter and configuration settings of the Wind Sensor. | M C | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
52 | <M> | The parameters/settings configured and stored in the Wind Sensor shall be retained after a re-start or shutdown due to power failure. | M C | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
53 | <M> | The time between switching on the Wind Sensor (power on) and the first output of a data message containing valid observational data shall be 1 minute or less. | M C | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
3.2 Communication Requirements
54 | <M> | The Wind Sensors shall be equipped with one serial interface for all data communications through either (case 1) RS422 full-duplex using 4 data wires (receive and transmit over separate wires); or (case 2) RS485 half-duplex using 2 data wires (receive and transmit over the same wires). | M I | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
55 | <M> | Handshake shall not be required for communication with the Wind Sensors, neither in hardware nor in software. | M C | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
56 | <M> | The Wind Sensors shall support at least one of the listed options for each of the following transmission characteristics: • Speed: 1200, 2400, 4800, 9600 or 19200 bits per second; • Parity: Even, Odd, or None; • Wordsize: 7 or 8 bits; • Stopbits: 1 or 2. | M C | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
57 | <M> | The Wind Sensors shall support operation in polling mode. | M C | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
58 | <M> | The Wind Sensors shall prevent blocking of further communication in case of transmission errors by ignoring any command which is not received completely within 10.000 milliseconds. | M C | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
59 | <M> | The response of the Wind Sensors to any command at any time shall start within 200 milliseconds. The response to any valid or invalid command shall be completely transmitted within 1500 milliseconds using the maximum requested and available transmission rate. | M C | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
3.3 Sensor Output Requirements
60 | <M> | The format of the Wind Sensor response to any command shall be specified. This specification shall include the length of any response. | M C | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
61 | <M> | The Wind Sensor response shall contain a data integrity check, for example a checksum. | M C | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
62 | <M> | The Wind Sensor response shall be enclosed by unique predefined delimiters, for example STX (start of text) and ETX (end of text). | M C | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
63 | <M> | The Wind Sensor response shall be fixed. Either all data is located at fixed positions in the data message, or a unique delimiter is unambiguously used to divide the data message into elements. | M C | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
64 | <M> | The Wind Sensor response shall be in readable ASCII. | M C | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
65 | <D> | It is desired to be possible within the Wind Sensor to create a user-defined output data message. | D C | 1 | Yes = Score of 5 No = Score of 0 | Select your answer: | No answer selected. | #VALUE!
3.4 Status Requirements
66 | <M> | The Wind Sensor shall report status information on the operational state of the Wind Sensor and clearly identify any problem, either within the regular data message or through documented diagnostic commands. | M C | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
67 | <M> | The status information shall unambiguously indicate whenever insufficient measurements have been made to produce valid measurement parameters, irrespective of the reason of insufficient measurements. | M C | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
68 | <M> | The Wind Sensor shall perform internal checks to ensure correct functioning of each module. Status information shall be reported for each module. | M C | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
-
4 Hardware Requirements
4.1 General Requirements
69 | <M> | The weight and dimensions of the Wind Sensor shall be such that it can be easily transported and installed by a single person, in accordance with Dutch ARBO legislation. Installation is done by KNMI on a tiltable wind mast. | M C | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
70 | <M> | The arms and ultrasonic transducers of the sensor shall be heated. | M C | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
71 | <M> | The Wind Sensor shall have a method to switch off the heating in order to conserve power. This method shall be useable during first-line maintenance activities. Note: KNMI intends to use this switch on specific sensors placed in off-grid locations where only a solar array may be available for power. | M C | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
72 | <M> | The mass of the sensor shall be less than 4 kg. | M C | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
73 | <M> | Each Wind Sensor shall be delivered in a reusable protective case that prevents damage to the instrument during transportation under normal conditions. A cardboard box with suitable protection can be used. | M C | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
74 | <D> | It is desired that Tenderer substantiates that their Wind Sensor is both environmentally and socially responsible. Factors to take into consideration might be the power usage of the sensor to achieve the desired functionality, the use of recycled raw materials in production of the sensor and its final recyclability, use of ethical labour practices during production, etc. Eco-friendly procurement is about preventing or minimising a negative impact on the environment. The reduction of environmental impact can be achieved in various ways, including the choice of materials, the production process through to placement, commissioning and maintenance. Eventually, a product will be disposed of, although placement and disposal is outside the scope of this agreement, the product offered may have a negative impact on the environment when taken out of service. Chosen materials or production techniques may have a positive or negative impact on the (degree of) recyclability of the offered product. Buyer aims for a product with limited environmental impact. Describe in what way and to what extent you include sustainability in your offer in the production, materials used, logistics, energy consumption, maintenance and recyclability of the product so that you meet the requested and substantiate its effectiveness. Consider the following preconditions: - All Requirements of the Programme of Requirements must be met; - Your solution must not involve separate costs from those specified in the price sheet; - Your proposal must be feasible and not entail an extension of the contract. | D O | 9 | Please provide a substantiation of no more than 2 A4 pages. The substantiation will be assessed and a score is given between 0 to 5 maximum. Evaluation criteria The assessment team will use your answer to assess whether and to what extent reservations have been made, your solution is effective, applicable and described in concrete terms, i.e.: - In what way and to what extent the environmental impact is limited from the production (including materials used), logistics, energy consumption, maintenance and recyclability of the product. The less environmentally harmful the solutions described are and the more this is quantified and made demonstrable, the higher the rating; the substantiation is an important element here. | Select applicable option: | The score will be assigned by the KNMI assessment team. | #VALUE!
4.2 Electrical Requirements
75 | <M> | The Wind Sensor shall have a valid and current European CE marking, or equivalent. Note: If compliance for any EN-norm is to any version other than the EU Harmonized version, a Declaration of Conformity including a Notified Body certificate must be made available to KNMI on first request. | M C | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
76 | <M> | The Tenderer shall guarantee electrically safe operation of the delivered Wind Sensor including its components according to Dutch and EU regulations. | M C | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
77 | <D> | Compliance by the Wind Sensor with electric safety standard EN-IEC 61010, or equivalent, is desired. | D C | 3 | Yes = Score of 5 No = Score of 0 | Select your answer: | No answer selected. | #VALUE!
78 | <M> | The Wind Sensor and related parts shall comply with all applicable EMC requirements, specifically European standards EN-IEC 61326-1 with compliance to EN-IEC 61000-6-2 for EMC immunity in industrial environments and EN-IEC 61000-6-3 for EMC emission for household environments, or equivalent. | M C | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
79 | <M> | The Wind Sensor shall be adequately protected against failures caused by lightning strikes. Note: The masts and towers on which KNMI shall mount the Wind Sensor are properly grounded. | M C | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
80 | <D> | Compliance by the Wind Sensor and related parts with European standard EN-IEC 62305 concerning lightning protection and failure of electrical installations, or equivalent, is desired. | D C | 2 | Yes = Score of 5 No = Score of 0 | Select your answer: | No answer selected. | #VALUE!
81 | <M> | The Wind Sensor shall be able to operate either (case 1) using a power supply voltage that falls somewhere in the range of 20VDC to 50VDC with a maximum current drawn of 10 Ampere, or (case 2) a 230VAC power supply voltage with a maximum current drawn of 10 Ampere. | M I | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
82 | <M> | The minimum power use of the wind sensor with all heaters active at 100% shall be at least 140 Watt. | M C | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
4.3 Mechanical Requirements
83 | <M> | The Wind Sensor and related parts shall have water and dust protection IP55 or better according to norm EN-IEC 60529, or equivalent. | M C | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
84 | <M> | The Wind Sensor shall be resistant to shock and vibration in normal operating conditions in all locations including North Sea platforms, as well as all transport conditions. | M C | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
85 | <D> | Compliance by the Wind Sensor with tests listed in European standard EN-IEC 60068-2 concerning shock and vibration resistance, or equivalent, is desired. | D C | 2 | Yes = Score of 5 No = Score of 0 | Select your answer: | No answer selected. | #VALUE!
86 | <M> | The material of the Wind Sensor and related parts shall be suitable for use in maritime and industrial environments. Any corrosive effects shall not have a negative impact on the functioning of the Wind Sensor. | M C | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
87 | <D> | Compliance by the Wind Sensor with European standard EN-IEC 60068-2-52 concerning suitability for use in marine environments, or equivalent, is desired. | D C | 2 | Yes = Score of 5 No = Score of 0 | Select your answer: | No answer selected. | #VALUE!
4.4 Installation Requirements
88 | <M> | The proper true north-alignment of the sensor shall be mechanically guaranteed by the mounting and/or connector, such that the mechanical orientation of the Wind Sensor is kept fixed within ±0.5° when replacing the Wind Sensor. Note: This may be achieved by use of a custom mounting solution. KNMI shall itself ensure the proper alignment of the mast head relative to true north. | M C | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
4.5 Sensor Connections Requirements
89 | <M> | The Wind Sensor shall be provided with both the power cable and serial connection cable, and with connectors attached. It must be possible to manually disconnect the connector at the sensor end of the cable without use of tools. These cables shall be suitable for use under exterior conditions and shall be available on the European market. It is acceptable if the power and serial connection are combined in one cable. The length of the cables and type of the connectors shall be specified by KNMI in consultation with Tenderer before delivery. The cable length will not exceed 12 meters for all but five sensors. For five sensors the cable length will not exceed 22 meter. | M C | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
90 | <M> | The cables supplied with the Wind Sensor must be: • Usable in the temperature range of -25°C to +50°C, and: • Oil and chemical resistant, and: • UV radiation resistant, and: • Precipitation resistant. | M C | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
91 | <D> | Compliance of the cables and connectors with European standard EN-IEC 61010, or equivalent, is desired. | D C | 2 | Yes = Score of 5 No = Score of 0 | Select your answer: | No answer selected. | #VALUE!
4.6 Operating Environment Requirements
92 | <M> | The Wind Sensor shall be able to operate according to the requirements (including accuracies measured in a wind tunnel) in all environmental conditions in the Netherlands, including the North Sea platforms and the Caribbean Netherlands. Specifically, it shall be able to operate according to the Requirements in the relative humidity range of 0% to 100% and the ambient temperature range of –25°C to +50°C. | M C | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
93 | <M> | The Wind Sensor shall not be damaged by environmental conditions in the Netherlands, including North Sea platforms and the Caribbean Netherlands. Damage caused by extreme weather events such as hurricanes and the debris associated with these events is exempt. | M C | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
-
5 Software Requirements
94 | <M> | Tenderer shall inform KNMI of each new software/firmware release for the Wind Sensor during the entire Contract period. This includes updates with additional features as well as bug fixes. Information shall be provided on the contents of the update, its consequences for the operation of the Wind Sensor, and whether or not the update is critical. | M C | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
95 | <M> | Tenderer shall supply new software/firmware updates in such way that the installation of updates or upgrades can be performed by adequately trained KNMI staff during 2nd line maintenance without effect on warranty. | M C | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
96 | <M> | Tenderer agrees that KNMI shall be free to choose not to install any updates or upgrades without loss of warranty or 3rd line maintenance. In the case that an update is not installed, KNMI accepts the shortcomings reported by the manufacturer which are solved in the new software/firmware version. | M C | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
97 | <M> | Tenderer shall safeguard the Wind Sensor to avoid lockup situations in which components or the whole system becomes inoperative or unresponsive. | M C | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
-
6 Reliability Requirements
98 | <M> | The Wind Sensor shall be COTS (Commercial Off The Shelf). | M C | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
99 | <M> | The Wind Sensor shall be designed and able to operate continuously 24 hours a day, 7 days a week. | M C | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
100 | <M> | The availability of the Wind Sensor shall be 99.98% or better, during which the measurements made shall meet the mandatory requirements. Events where insects or birds cause the Wind Sensor to operate incorrectly are included in the 0.02% of measurements that do not meet this requirement. | M C | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
101 | <M> | The number of 10-minute periods with more than 10% missing or invalid measurement samples shall be less than 0.03% of all 10-minute periods, irrespective of the cause of the problem, e.g. insects, birds, turbulence, vibration, electrical or acoustic noise, lightning interference, etc. | M C | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
102 | <M> | The Wind Sensor shall continue to meet the mandatory requirements during periods of precipitation. In no more than 0.01% of the total number of 10-minute periods shall more than 3.0 seconds of missing or invalid measurement samples occur due to precipitation. Note: all 10-minute periods will be counted, including those without precipitation. | M C | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
103 | <D> | It is desired that the availability of the Wind Sensor is 99.99% or better, during which the measurements made meet the mandatory requirements. Substantiation of this availability by means of a field test report is highly desired, ideally including analysis of the factors that risk downtime and their mitigation in the Wind Sensor, and by providing detailed information including a field report on availability of the Wind Sensors including downtime due to exchange of the sensor for maintenance. Effects such as insect or bird disturbance, weather effects such as precipitation and any other sources of downtime should be discussed and accounted for. | D O | 14 | Please provide a substantiation of no more than 2 A4 pages. The substantiation will be assessed and a score is given between 0 to 5 maximum. The assessment is based on expert judgement with special attention given to the following assessment criteria: - Inclusion of field tests substantiating the availability in all seasons covering climate conditions comparable to the Netherlands, including North Sea platforms and the Carribean Netherlands, is highly desired. - Influence of insects and birds should be discussed and taken into account in the substantiation. - Influence of weather effects such as precipitation should be discussed and taken into account. - Other sources of missing raw samples should be clarified and their reduction should be discussed. Less samples lost is better. - The actual availability percentage is very important in determining the final score. Higher availability is better. | Select applicable option: | The score will be assigned by the KNMI assessment team. | #VALUE!
104 | <M> | The life time of the Wind Sensor shall be at least 10 years, including first, second and third line maintenance. | M C | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
105 | <M> | The Mean Time Between Failure (MTBF) of the Wind Sensor shall be at least 43830 hours (=60 months), based on all environmental conditions in the Netherlands, including North Sea platforms and the Caribbean Netherlands and continuous 24/7 operation. | M C | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
-
7 Support and Maintenance Requirements
7.1 Support Requirements
106 | <M> | A helpdesk with skilled personnel fluent in Dutch and/or English shall be available and reachable by telephone and e-mail. The helpdesk shall reply to questions with a satisfactory answer within two working days. | M C | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
7.2 Maintenance Requirements
107 | <M> | The design of the wind sensor and mounting connectors shall be such that the cleaning and 1st line maintenance can be performed by one person. | M C | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
108 | <M> | Wind Sensors or parts that have been sent to the Tenderer for repair shall be returned after repair or replaced within 2 months (including FAT and transport). | M C | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
109 | <M> | The serviceability of the Wind Sensor shall be either (case 1) based on the concept of Line Replaceable Units (LRU), which means that it is designed to be repaired quickly at KNMI by either replacement of modular components, or (case 2) the Wind Sensor as a whole is considered the LRU. For case 1, Tenderer shall provide a list with recommended LRU items including recommended amounts. For case 2, the number of recommended spare Wind Sensors shall be specified. Note that these spare Wind Sensors are not included in the ~10% spares that KNMI requires for performing 2nd line maintenance. | M I | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
110 | <M> | The Wind Sensor shall be able to operate in accordance with the requirements in the field for at least 12 months in between scheduled 1st line maintenance. | M C | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
7.3 Delivery Requirements
111 | <M> | Tenderer shall commit to the delivery schedule and support required during the verification test as described in Annex 07 - Verification Phase & Delivery. | M C | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
-
8 Documentation Requirements
112 | <M> | Tenderer shall include a comprehensive and complete set of manuals and technical documentation written in Dutch or English with instructions for operation, installation, maintenance and verification of the Wind Sensor. This set shall at least include as much information as possible on: • Technical data/specifications; • Data output description incl. measurement examples; • Status information output description; • Data communication options; • Installation instructions; • Mechanical and electronic connections, incl electronic wiring schematic diagrams; • Electrical specifications of digital outputs, power supply, power consumption and temperature range of the instrument; • Maintenance guidelines (on site and indoors); • Verification Protocol; • Calibration guidelines; • Safety instructions; • Troubleshooting. This documentation set shall be delivered directly after receiving an "Intention to Award" decision, and at the latest before the Final Award decision will be published. | M C | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
113 | <M> | Tender shall supply new or updated documentation when new firmware releases or upgrades of the Wind Sensor are delivered to KNMI, or when incorrect or incomplete information in the current documentation set is found. | M C | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
114 | <M> | All documentation files shall be formatted in Portable Document Format (PDF). | M C | n/a | Yes=Compliant to Requirement No=Failure to comply | Select your answer: | No answer selected. | n/a
115 | <D> | It is desired to have the documentation set specifying how to implement, use and maintain the Wind Sensor included with the tender. The level of detail of this documentation set is very important, and is desired to be as high as possible. | D O | 6 | Provide the documentation package in a separate file. The documentation will be assessed and a score is given between 0 to 5 maximum. The assessment is based on expert judgement with special attention given to the following assessment criteria: - Detail: does the documentation provide enough detail to both implement the sensor and understand as well as execute first & second line maintenance? - Scope: does the documentation include all the necessary topics. - Clarity: Is the documentation written in clear and concise language without unnecessary jargon, are instructions clearly written in a step-by-step structure, are illustrations used to explain complex instructions. - Structuring: Is the index clear and can desired information of varying kinds be easily found within the entire documentation package? | Select your answer: | The score will be assigned by the KNMI assessment team. | #VALUE!

## Summary of Quality criteria
Summary of Quality criteria points
Chapter | # of desirables | Sub-total points achieved | Total points achieved
|  | (without KNMI-assessed desirables) | (with KNMI-assessed desirables)
1 - Functional | 13 | #VALUE! | #VALUE!
2 - Maintainability | 3 | #VALUE! | #VALUE!
3 - Technical | 1 | #VALUE! | #VALUE!
4 - Hardware | 6 | #VALUE! | #VALUE!
5 - Software | 0 | 0 | 0
6 - Reliability | 1 | 0 | #VALUE!
7 - Support and Maintenance | 0 | 0 | 0
8 - Documentation | 1 | 0 | #VALUE!
Totals | 25 | #VALUE! | #VALUE!
(maximum) |  | (200) | (800)
Please note: this column will only be filled in after the assessment by KNMI is completed.

## Technical Area (hidden)
This sheet contains technical structures to assist with Excel Data Validation. Tenderers do not need to take any actions with respect to this sheet.
For all Mandatory Closed & Desirable Closed
Select your answer:
Yes
No
Substantiations Yes/No
Select applicable option:
Yes, Substantiation is provided.
No, Substantiation is not submitted.
Desirable 15 (max wind speed range)
Select the applicable wind speed range:
0-65 m/s
At least 0-70 m/s
At least 0-75 m/s
At least 0-80 m/s
At least 0-85 m/s
At least 0-90 m/s
Desirable 115 (Documentation package)
Select your answer:
Yes, documentation package is included separately.
No
For mandatory information, case-based strict OR
Select your answer:
Yes, Case 1 is applicable
Yes, Case 2 is applicable
No
For mandatory information, case-based AND/OR
Select your answer:
Yes, Case 1 is applicable
Yes, Case 2 is applicable
Yes, both cases are applicable
No
Desirable 4: higher sampling rate
Select the actual sampling rate:
≤ 12 Hz
12-90 Hz
90-160 Hz
160-220 Hz
220-270 Hz
> 270 Hz