## Introduction
Step | Wetsus Electrospinner | Dear Supplier,
4 steps are introduced in this Tender.
The steps are explained below.
Step 1 | Knock-out phase | This step contains knock-out questions. If one of the questions is answered negatively, then the offer is not eligible
Step 2 | Tender questions | This step contains additional Tender questions: here points can be scored.
Step 3 | Factory Acceptance Test | This step contains the Factory Acceptance Test
Step 4 | Site Acceptance Test | This step contains the Site Acceptance Test

## Knock-out
Supplier name
Contact person name
Telephone number contact person
Email address contact person
Drop-down menu with Yes/No
Category | Instrument | Type | Answer Supplier
FUNCTIONALITY
Electrospinning 3 different polymers at 3 different positions | Can the instrument electrospin 3 different polymers such that these polymers do not mix? See sheet Not Mixing Materials. | Knock-out | No
Electrospinning 3 different polymers at a common position | Can the instrument electrospin 3 different polymers at a common position? See sheet Mixing Materials. | Knock-out | No
Electrospinning with single needles | Can the instrument electrospin with single needles for each polymer solution for Research mode? | Knock-out | No
First high voltage source (positive) | Has the instrument a first differential voltage source of 0-60kV? | Knock-out | No
Second high voltage source (positive) | Has the instrument a second differential voltage source of 0-60kV? | Knock-out | No
Third high voltage source (positive) | Has the instrument a third differential voltage source of 0-60kV? | Knock-out | No
Fourth high voltage source (negative) | Has the instrument a fourth differential voltage source of 0-60kV? | Knock-out | No
Environment Control Unit - Temperature | Can the temperature be controlled in the range of 18-45 degrees Celsius in the electrospinning chamber with a stability of +/- 1 degree Celsius? | Knock-out | No
Environment Control Unit -Relative Humidity | Can the Relative Humidity be controlled from 10% to 80% in the electrospinning chamber with a stability of +/- 3%, depending on temperature? | Knock-out | No
Temperature control, monitoring and logging, display | Can the instrument control, monitor, log, and display the Temperature (Celsius)? | Knock-out | No
Relative Humidity control, monitoring, logging, display | Can the instrument control, monitor, log, display the Relative Humidity (%)? | Knock-out | No
Recipe storage | Can the instrument store recipes with all instrument settings? | Knock-out | No
Recording electrospinning proces with digital camera | Can the instrument record the electrospinning process with video in MP4 or MPEG? | Knock-out | No
Roll-to-roll conveyor belt mechanism | Has the instrument a bi-directional conveyor belt mechanism with forward and backward winding with 1-100mm linear speed? | Knock-out | No
Roll width | Is the roll width between 300-500 mm? | Knock-out | No
Controlling emitter-collector distance | Has the instrument an automated regulator for emitter-collector distance of 0-300mm? | Knock-out | No
Research and pliot scale facilities within one instrument | Can instrument fulfill the operating range from R&D up to pilot scale with large-volume liquid feeding system up to 2L? | Knock-out | No
Dosing speed for the pumps | Can instrument handle flow range for the dosing speed per pump between 0.5-10 mL/hour? | Knock-out | No
Gas/min during operation | Can the instrument operate with <30L nitrogen gas/min during operation with inert gas? | Knock-out | No
Volume of the dosing liquid | Can instrument store between 0.1- 60ml for each volume of the 3 different liquids used for electrospinning? | Knock-out | No
Rotating speed drum collector | Can instrument handle range between 200-1000 rpm for the rotating drum collector? | Knock-out | No
SAFETY |  | Knock-out | No
Safety features | Does the instrument shut down all rotating and dosing equipment before opening cabin? | Knock-out | No
Safety features | Does the instrument ventilate 10X the cabin volume before opening cabin? | Knock-out | No
Safety features | Does the instrument have a diffuse LED lightning? | Knock-out | No
Safety features | Has the instrument the possibility to work with inert gases, in terms of safety, like nitrogen gas? | Knock-out | No
Safety features | Does the instrument follow European CE standard? | Knock-out | No
Safery features | Is the user protected against any electrical shock? | Knock-out | No
Noise level | Does the instrument produce <70 dB(a) at 1m distance under working conditions? | Knock-out | No
BOUNDARY CONDITIONS
Gas supply | Can the instrument operate with <30L nitrogen gas/min during operation with inert gas? | Knock-out | No
Power connection | Does the instrument require a power connection with a maximum of 32A/400V AC? | Knock-out | No
Power connection | Will an eventual instrument frequency controller(s) interference with earth leakage circuit breaker, be solved by the Supplier? | Knock-out | No
Restriction in height | Is instrument maximal 2300mm high? | Knock-out | No
Restriction in dimensions | Has the instrument a maximal floor surface of 1300mm x 2100mm? | Knock-out | No
Restriction in floor load | Has instrument a maximal floor load of 600 kg/m2? | Knock-out | No
Delivery period | Is delivery of 100% Operational Instrument within 6 months after purchase order guaranteed? | Knock-out | No
Maximum price of total package = Electrospinner + ECU + delivery + instruction + 5 years of warranty + replacement parts (excluding Nice to have Options, see Tender questions) | Is the maximum price of the Electrospinner + ECU + delivery + instruction + 5 years of warranty + replacement parts (excluding Nice to have Options, see Tender questions) less than 360.000 EURO excl.VAT? | Knock-out | No
Yes
No

## Tender questions
Requirements | All questions below must be answered; If the indicated range is not reached, the answer "No" must be selected. | Knock-out | If range is extended, give the range (and unit) of the instrument | You can receive a maximum of 250 points per question | Select "Yes" of "No"
Must have requirements | Instrument | Range | Range | Points | Answer Supplier | Remark Wetsus | Final points scored
Dosing speed for the pumps | Can instrument handle flow range for the dosing speed per pump in mL/hour? | 0.1-10 | Dosing speed 0.1-100 | 200 | No | Only one 'Yes' is allowed for this requirement. | 0
|  |  | Dosing speed 0.1-1000 | 250 | No |  | 0
High voltage power supplies | Can instrument handle differential voltage range for the high voltage supplies in kVolt? | 0-60 | 0 - 80 | 200 | No | Only one 'Yes' is allowed for this requirement. | 0
|  |  | 0- 100 | 250 | No |  | 0
Gas volume/min during operation | Can the instrument operate with nitrogen gas during operation at <30 Liter nitrogen gas/min? | <30 L nitrogen gas/min | <= 20 | 250 | No | Only one 'Yes' is allowed for this requirement. | 0
|  |  | > 20 and <= 30 | 200 | No |  | 0
Volume of the dosing liquid | Can instrument handle the range given for the stored volume of the 3 different dosing liquid in mL? | >=0.1 and <= 60 | >= 0.1 and <= 200 | 200 | No | Only one 'Yes' is allowed for this requirement. | 0
|  |  | >= 0.1 and <= 1000 | 250 | No |  | 0
Rotating speed drum collector | Can instrument handle range given for the rotating drum collector in rpm? | >= 200 and <=1000 | >= 200 and <= 2000 | 200 | No | Only one 'Yes' is allowed for this requirement. | 0
|  |  | >= 200 and <= 2500 | 250 | No |  | 0
TOTAL POINTS |  |  |  | 1250
0
Nice to have (optional) requirements | [6 empty] | 0
0
Incorporation of an open loop for inert nitrogen gas | Additional feature of maximal 5000 euro excl. VAT | Please state price in pricing sheet |  |  |  |  | 0
Real-time thickness monitoring system for rotating drum collector | Additional feature of maximal 5000 euro excl. VAT | Please state price in pricing sheet |  |  |  |  | 0
Real-time thickness monitoring system for roll-to-roll collector | Additional feature of maximal 5000 euro excl. VAT | Please state price in pricing sheet |  |  |  |  | 0
Two high-pressure syringe pumps | Additional feature of maximal 25000 euro excl. VAT | Please state price in pricing sheet |  |  |  |  | 0
Needleless slit injectors | Additional feature of maximal 5000 euro excl. VAT | Please state price in pricing sheet |  |  |  |  | 0
0
[6 empty] | TOTAL NUMBER OF POINTS SCORED | 0

## FAT
The FAT Factory Acceptance Test test is conducted at the facilities of the supplier.
The technical results of the FAT test are important and must be shared with Wetsus
| Category | Electrospinner - Factory Acceptance Test | Answer Supplier
| Test dosing rate pumps of 0.5-10 mL/h range | Determine dosing rate in ml/h of pumps
| Test high voltage power supplies of 60kV differential voltage range | Determine high voltage range of 4 differential voltage sources
| Test conveyor belt speed range in forward position of 1-100mm/s range | Determine linear speed range in mm/s in forward position
| Test conveyor belt speed range in backward position of 1-100mm/s range | Determine linear speed rsnge in mm/s in backward position
| Test rotating rate range of the rotating drum collector of 200-1000rpm range | Determine rpm rate of the rotating drum
| Test Relative Humidity range in % (10-80) and Temperature range in degrees Celsius (18-45) of the Electrospinning chamber | Determine RH and temperature for the given ranger of RH and temperature
| ADDITIONAL TESTS CONDUCTED BY SUPPLIER | RESULT OF THE ADDITIONAL TEST CONDUCTED BY SUPPLIER
Category | Instrument CHECK
FUNCTIONALITY |  | TEST TYPE IS KNOCK-OUT | IS TEST SUCCESFUL AT SITE?
Electrospinning 3 different polymers at 3 different positions | Can the instrument electrospin 3 different polymers such that these polymers do not mix? See sheet Not Mixing Materials. | Knock-out | No
Electrospinning 3 different polymers at a common position | Can the instrument electrospin 3 different polymers at a common position? See sheet Mixing Materials. | Knock-out | No
Electrospinning with single needles | Can the instrument electrospin with single needles for each polymer solution for Research mode? | Knock-out | No
First high voltage source (positive) | Has the instrument a first differential voltage source of 0-60kV? | Knock-out | No
Second high voltage source (positive) | Has the instrument a second differential voltage source of 0-60kV? | Knock-out | No
Third high voltage source (positive) | Has the instrument a third differential voltage source of 0-60kV? | Knock-out | No
Fourth high voltage source (negative) | Has the instrument a fourth differential voltage source of 0-60kV? | Knock-out | No
Environment Control Unit - Temperature | Can the temperature be controlled in the range of 18-45 degrees Celsius in the electrospinning chamber with a stability of +/- 1 degree Celsius? | Knock-out | No
Environment Control Unit -Relative Humidity | Can the Relative Humidity be controlled from 10% to 80% in the electrospinning chamber with a stability of +/- 3%, depending on temperature? | Knock-out | No
Temperature control, monitoring and logging, display | Can the instrument control, monitor, log, and display the Temperature (Celsius)? | Knock-out | No
Relative Humidity control, monitoring, logging, display | Can the instrument control, monitor, log, display the Relative Humidity (%)? | Knock-out | No
Recipe storage | Can the instrument store recipes with all instrument settings? | Knock-out | No
Recording electrospinning proces with digital camera | Can the instrument record the electrospinning process with video in MP4 or MPEG? | Knock-out | No
Roll-to-roll conveyor belt mechanism | Has the instrument a bi-directional conveyor belt mechanism with forward and backward winding with 1-100mm linear speed? | Knock-out | No
Roll width | Is the roll width between 300-500 mm? | Knock-out | No
Controlling emitter-collector distance | Has the instrument an automated regulator for emitter-collector distance of 0-300mm? | Knock-out | No
Research and pliot scale facilities within one instrument | Can instrument fulfill the operating range from R&D up to pilot scale with large-volume liquid feeding system up to 2L? | Knock-out | No
Dosing speed for the pumps | Can instrument handle flow range for the dosing speed per pump between 0.5-10 mL/hour? | Knock-out | No
Gas/min during operation | Can the instrument operate with <30L nitrogen gas/min during operation with inert gas? | Knock-out | No
Volume of the dosing liquid | Can instrument store between 0.1- 60ml for each volume of the 3 different liquids used for electrospinning? | Knock-out | No
Rotating speed drum collector | Can instrument handle range between 200-1000 rpm for the rotating drum collector? | Knock-out | No
SAFETY |  | Knock-out | No
Safety features | Does the instrument shut down all rotating and dosing equipment before opening cabin? | Knock-out | No
Safety features | Does the instrument ventilate 10X the cabin volume before opening cabin? | Knock-out | No
Safety features | Does the instrument have a diffuse LED lightning? | Knock-out | No
Safety features | Has the instrument the possibility to work with inert gases, in terms of safety, like nitrogen gas? | Knock-out | No
Safety features | Does the instrument follow European CE standard? | Knock-out | No
Safery features | Is the user protected against any electrical shock? | Knock-out | No
Noise level | Does the instrument produce <70 dB(a) at 1m distance under working conditions? | Knock-out | No
BOUNDARY CONDITIONS |  | Knock-out | No
Gas supply | Can the instrument operate with <30L nitrogen gas/min during operation with inert gas? | Knock-out | No
Power connection | Does the instrument require a power connection with a maximum of 32A/400V AC? | Knock-out | No
Power connection | Will an eventual instrument frequency controller(s) interference with earth leakage circuit breaker, be solved by the Supplier? | Knock-out | No
Restriction in height | Is instrument maximal 2300mm high? | Knock-out | No
Restriction in dimensions | Has the instrument a maximal floor surface of 1300mm x 2100mm? | Knock-out | No
Restriction in floor load | Has instrument a maximal floor load of 600 kg/m2? | Knock-out | No

## SAT
| Category | Electrospinner - Site Acceptance Test | How to determine | Checked by Wetsus | Notes from Supplier if SAT not reached
| Test dosing speed range of pumps with range 0.5-10mL/h | Determine if range is measred | Collect liquid and weigh
| Test voltage high voltage source of 60kV differential voltage range | Determine if range is measred | Measure with high voltage meter
| Test conveyor belt speed in forward position of 1-100mm/s speed range | Determine if range is measred | Measure foil distance transported
| Test conveyor belt speed in backward position of 1-100mm/s speed range | Determine if range is measred | Measure foil distance transported
| Test rotating rate of the rotating drum collector of 200-1000 rpm range | Determine if range is measred | Measure rotating speed
| Test Relative Humidity in % (10-80) range and Temperature in degrees Celsius (18-45) range of electrospinning chamber | Determine if range is measred | Measure RH and temperature
Category | Instrument CHECK AT WETSUS
FUNCTIONALITY |  | TEST TYPE IS KNOCK-OUT | IS TEST SUCCESFUL AT SITE?
Electrospinning 3 different polymers at 3 different positions | Can the instrument electrospin 3 different polymers such that these polymers do not mix? See sheet Not Mixing Materials. | Knock-out | No
Electrospinning 3 different polymers at a common position | Can the instrument electrospin 3 different polymers at a common position? See sheet Mixing Materials. | Knock-out | No
Electrospinning with single needles | Can the instrument electrospin with single needles for each polymer solution for Research mode? | Knock-out | No
First high voltage source (positive) | Has the instrument a first differential voltage source of 0-60kV? | Knock-out | No
Second high voltage source (positive) | Has the instrument a second differential voltage source of 0-60kV? | Knock-out | No
Third high voltage source (positive) | Has the instrument a third differential voltage source of 0-60kV? | Knock-out | No
Fourth high voltage source (negative) | Has the instrument a fourth differential voltage source of 0-60kV? | Knock-out | No
Environment Control Unit - Temperature | Can the temperature be controlled in the range of 18-45 degrees Celsius in the electrospinning chamber with a stability of +/- 1 degree Celsius? | Knock-out | No
Environment Control Unit -Relative Humidity | Can the Relative Humidity be controlled from 10% to 80% in the electrospinning chamber with a stability of +/- 3%, depending on temperature? | Knock-out | No
Temperature control, monitoring and logging, display | Can the instrument control, monitor, log, and display the Temperature (Celsius)? | Knock-out | No
Relative Humidity control, monitoring, logging, display | Can the instrument control, monitor, log, display the Relative Humidity (%)? | Knock-out | No
Recipe storage | Can the instrument store recipes with all instrument settings? | Knock-out | No
Recording electrospinning proces with digital camera | Can the instrument record the electrospinning process with video in MP4 or MPEG? | Knock-out | No
Roll-to-roll conveyor belt mechanism | Has the instrument a bi-directional conveyor belt mechanism with forward and backward winding with 1-100mm linear speed? | Knock-out | No
Roll width | Is the roll width between 300-500 mm? | Knock-out | No
Controlling emitter-collector distance | Has the instrument an automated regulator for emitter-collector distance of 0-300mm? | Knock-out | No
Research and pliot scale facilities within one instrument | Can instrument fulfill the operating range from R&D up to pilot scale with large-volume liquid feeding system up to 2L? | Knock-out | No
Dosing speed for the pumps | Can instrument handle flow range for the dosing speed per pump between 0.5-10 mL/hour? | Knock-out | No
Gas/min during operation | Can the instrument operate with <30L nitrogen gas/min during operation with inert gas? | Knock-out | No
Volume of the dosing liquid | Can instrument store between 0.1- 60ml for each volume of the 3 different liquids used for electrospinning? | Knock-out | No
Rotating speed drum collector | Can instrument handle range between 200-1000 rpm for the rotating drum collector? | Knock-out | No
SAFETY |  | Knock-out | No
Safety features | Does the instrument shut down all rotating and dosing equipment before opening cabin? | Knock-out | No
Safety features | Does the instrument ventilate 10X the cabin volume before opening cabin? | Knock-out | No
Safety features | Does the instrument have a diffuse LED lightning? | Knock-out | No
Safety features | Has the instrument the possibility to work with inert gases, in terms of safety, like nitrogen gas? | Knock-out | No
Safety features | Does the instrument follow European CE standard? | Knock-out | No
Safety features | Is the user protected against any electrical shock? | Knock-out | No
Noise level | Does the instrument produce <70 dB(a) at 1m distance under working conditions? | Knock-out | No
BOUNDARY CONDITIONS |  | Knock-out | No
Gas supply | Can the instrument operate with <30L nitrogen gas/min during operation with inert gas? | Knock-out | No
Power connection | Does the instrument require a power connection with a maximum of 32A/400V AC? | Knock-out | No
Power connection | Will an eventual instrument frequency controller(s) interference with earth leakage circuit breaker, be solved by the Supplier? | Knock-out | No
Restriction in height | Is instrument maximal 2300mm high? | Knock-out | No
Restriction in dimensions | Has the instrument a maximal floor surface of 1300mm x 2100mm? | Knock-out | No
Restriction in floor load | Has instrument a maximal floor load of 600 kg/m2? | Knock-out | No
Note on F2: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Is deze kolom echt nodig?