## Instructions
Programme of Requirements
version: | V1.0 Date 29/07/24 (Draft version)
KNMI Calibration Laboratory Pressure Measurement Standard Replacement (PMS)
IMPORTANT INFORMATION
This Program of Requirements and Desirables forms an integral part of the Tender. It details the functional and non-functional Requirements for the KNMI Calibration Laboratory Pressure Measurement Standard Replacement (PMS). A Requirement is either a Mandatory requirement (M) or a Desirable (D). The Tenderer must answer all Mandatory requirements (Yes/No) and can add short remarks in this document (The requirements sheet). Note: Your remarks may not act as a qualification of your answer. For some Mandatory requirements a substantiation is required in order for KNMI to score the award criterion. The same applies for all Desirables. For more information on the award criteria, see the Descriptive Document and its Annex 6. In general, the assessment is done using the guidelines described in the Descriptive Document. Additional focus points are mentioned in the Evaluation Criteria next to each Desirable, indicating the specific aspects that are of interest to the KNMI assessment team. Read these carefully, because for the Desirables these evaluation criteria can contain guidelines on what is considered necessary to achieve the minimum score. This document also functions as the formal declaration of the Tenderer to comply with the Requirements and to indicate to what extent the Tenderer complies with the Desirables. For this purpose, fill in the Requirements and Desirables sheet of this excel file (Requirements sheet) for all orange cells. These cells typically contain a drop-down menu that can be used to select the proper option. Tenderer may not change any other cells within this excel file even if technically possible. For the purpose of answering the open Desirables with a substantiation, Annex 6 is provided with a numbering of the Desirables that require substantiations. Tenderer may use this annex to structure the substantiations. Tenderer may also use an own template or file to structure the substantiations. In all cases, the text in this Program of Requirements is leading for all Requirements. The amount of space available for the substantiations is given in terms of A4 pages. This is done deliberately to allow for the inclusion of figures & graphs which tenderer may use to support the substantiation. The text in these substantiations must be in "Calibri" font with font size 11pt or larger. Any text within graphics must be of similar (or larger) size and readability. Any substantiation text or graphics which are outside the allowed size will be ignored during the assessment.

## Requirements
Programme of Requirements
version: | V1.0 Date 29/07/24 (Draft version)
KNMI ‘Pressure standard for absolute gas pressure’ with Reference No. TN 514478 - 31203349
Tenderer response
NOTICE: Not all Mandatories completed!
No. | Requirement | Type | Clarification | Meets requirement Yes/No | Remarks | Status
A | Technical Specifications
1 | The operation range of the PMS must be at least (900 - 1100) hPa absolute. | M |  | Select your answer: |  | No answer
2 | The accuracy of the PMS is 0,0035 % (35 ppm) of the generated pressure or better. | M | To prove the accuracy of the PMS, an example of a calibration certificate of a accredited calibration laboratory should be available upon request. | Select your answer: |  | No answer
3 | The pressure transmitting medium used by the PMS must be dry, oil free air. | M |  | Select your answer: |  | No answer
4 | The resolution of the digital output of the PMS is 0,001 hPa or better. | M |  | Select your answer: |  | No answer
5 | The PMS is equipped with USB and/or RS-232 communication data port, it must be fit to install our own software (Labview). | M |  | Select your answer: |  | No answer
6 | The dimensions of the PMS is not larger as 1,25 m x 0,75 m x 0,75 m (l x w x h). | M |  | Select your answer: |  | No answer
7 | The total weight of the PMS is maximum 100 kg. | M |  | Select your answer: |  | No answer
B | Configuration and Delivery aspects
1 | The PMS must be supplied with all peripheral equipment (total set-up) required to generate absolute pressures between 900 hPa and 1100 hPa ( vacuum pump, vacuum sensor, pressure generator, pipes, electrical peripheral equipment, cables, manuals). | M |  | Select your answer: |  | No answer
2 | The Pressure Measurement Standard must be delivered in working order on site (KNMI De Bilt, NL), you must be able to prove after installation that it achieves the desired parameters. | M |  | Select your answer: |  | No answer
3 | The PMS must be delivered in solid transport cases (e.g. Pelicase), that can be reused for shipping/transport. | M |  | Select your answer: |  | No answer
4 | Delivery period of the PMS is not more than 3 months. | M |  | Select your answer: |  | No answer
5 | The software for directly communication and controlling the PMS is included. | M |  | Select your answer: |  | No answer
6 | The software for directly communication and controlling the PMS must be in English or Dutch. | M |  | Select your answer: |  | No answer
7 | The PMS is supplied with an initial calibration certificate performed by an ISO/IEC 17025:2017 accredited calibration laboratory. | M |  | Select your answer: |  | No answer
8 | The initial calibration certificate is not older than 3 months when the PMS is delivered on site. | M |  | Select your answer: |  | No answer
9 | The calibration certificate of the PMS is in the English and/or Dutch language. | M |  | Select your answer: |  | No answer
C | Technical Performance and additional aspects
1 | Automatic calibration points between 900 hPa and 1100 hPa absolute generated by the PMS are 0,1 hPa or smaller. | M |  | Select your answer: |  | No answer
2 | The PMS does not use applications in the Cloud (to avoid information security problems). | M |  | Select your answer: |  | No answer
3 | The PMS is suitable for routine use (approximately 16 hours/week). | M |  | Select your answer: |  | No answer
4 | The critical parts of the PMS can be dismantled for maintenance and/or cleaning without assistance from the manufacturer/supplier. | M |  | Select your answer: |  | No answer
D | Condition of Use of the PMS
1 | The PMS must function at an ambient temperature of (21 ± 5) °C. | M |  | Select your answer: |  | No answer
2 | The PMS must function at a relative humidity: (55 ± 30) % rh. | M |  | Select your answer: |  | No answer
3 | The PMS can function in an industrial calibration laboratory. | M |  | Select your answer: |  | No answer
E | Warranty and maintenance (vacuum pump, vacuum sensor, pressure generator, pipes, electrical peripheral equipment, cables, manuals)
1 | The warranty period of the PMS is at least 2 years > this includes repair, service, parts replacement and shipping/transport. Limitations or Exclusions must be transparently stated. | M |  | Select your answer: |  | No answer
2 | Description on how to Obtain Warranty Service, the Customer Service Setup and Aftersales. | M | Documentation upon request. | Select your answer: |  | No answer
3 | The serviceability and availability of spare parts of the PMS: is at least 22 years. | M |  | Select your answer: |  | No answer
4 | The maximum period for solving defects, replacing parts and repairs regarding the PMS by manufacturer is 2 months. | M |  | Select your answer: |  | No answer
5 | Remote support of the supplier is needed for technical questions such as information about the maintenance of the pressure standard, using the additional components of the configuration, calibration of the standard and question about the user manual. Remote support must be available for the entire lifetime of the pressure standard. This also applies to the additional components of the pressure standard such as the vacuum pump, pressure regulator and associated measuring equipment. | M |  | Select your answer: |  | No answer
F | Documantation and Training Requirements
1 | The delivery of the PMS includes a compact user training that at least covers the operation of the equipment, 1st and 2nd line maintenance, solving minor malfunctions and assembling and disassembling the components of the setup in the event of sending for recalibration. The training must be available for at least 4 persons. | M |  | Select your answer: |  | No answer
2 | A complete user manual of the PMS must be available (preferably digital). | M |  | Select your answer: |  | No answer
3 | The training regarding the PMS must be provided in English and/or Dutch. | M |  | Select your answer: |  | No answer
4 | The training regarding the PMS must be offered within 6 weeks after delivery on site (exact time, date and location is discussable). | M |  | Select your answer: |  | No answer

## Used abbreviations
List of used abbreviations
Abbreviation | Explanation
PMS | Pressure Measurement Standard
l | length
w | width
h | height
ppm | parts per million
m | meter
kg | kilogram
hPa | hectopascal (mbar)
USB | Universal Serial Bus
KNMI | Royal Dutch Meteorology Institute

## Inputdata (hidden)
On-premise
SaaS
Select your answer:
Yes
No