## Requirements
Annex A04 - Compliance to Requirements - Answers to Preferences Applicable to the European public Procurement Procedure for a femtosecond laser source system, Reference number WS2359773965
The Tenderer demonstrates the ability to meet the requirements by answering Yes/No to the questions, which must be structured in accordance with the requirements given in Chapter 8. Tenders that are incomplete and/or fail to meet all the formal requirements will be declared invalid by TNO and discarded. Missing answers, documents, or data will result in exclusion if there is no rectifiable omission. TNO assesses the Tenders against the Programme of Requirements and Preferences. The Programme of Requirements and Preferences includes Minimum Requirements, Performance Requirements and Preferences. The Minimum Requirements must be met on penalty of invalidity at the time of submission of the Tender, the Performance Requirements only at the time of performance of the Contract and the Preferences must be worked out by Tenderers as part of the award of the Contract.
Requirement | Requirement description | Confirmation of fullfillment of requirement Yes/No | Remarks
Requirement 8.1.05 | Output Power: Average power ≥ 75 Watt (fundamental wavelength), ≥ 30 W (2nd harmonic) and ≥ 5 Watt (4th harmonic).
These power requirements are defined by the high throughput application of interconnect scribing.
Requirement 8.1.10 | Fundamental wavelength is defined as between 1000 and 1100 nm (IR).
Shorter and/or longer fundamental wavelengths are not accepted.
Requirement 8.1.15 | Wavelength: Laser must deliver a software selectable output at fundamental (IR) wavelength, second harmonic (Green, 2nd harmonic) and UV wavelength (4th harmonic).
The Laser must deliver only the selected wavelength.
Requirement 8.1.20 | Second and fourth harmonic wavelength is selected by polarisation based on rotation of waveplates.
The exit angle of the beams remain within the beam pointing stability of the laser system.
Requirement 8.1.25 | Repetition rate: Highest pulse frequency for all three wavelength ≥ 2 MHz, lowest pulse frequency single shot.
Scan velocities of at least 10 m/s are needed for high volume throughput. In this case, with a minimum pulse to pulse distance of 5 micron the minimum laser frequency will be 2 MHz.
Requirement 8.1.30 | Pulse energy: Maximum pulse energy > 0.75 mJ/pulse at IR wavelength.
The minimum pulse energy for this system is defined by multiplexing the laser into up to 30 multiple beams of each 25 µJ.
Requirement 8.1.35 | Pulse selection: Pulse on demand, including single shot.
To study effect of single pulses, the pulse on demand feature is crucial.
Requirement 8.1.40 | Pulse duration: Adjustable pulse width 250 fs – 10 ps, software selectable.
The effect of the pulse width selects different processes. For through substrate processing the maximum selectable pulse width should be 10 ps. For multi photon absorption the minimum pulse width should be 250 fs.
Requirement 8.1.45 | Pulse to pulse stability < 1 % at all three wavelengths within specified repetition rate and power output ranges for 12 hours.
The pulse to pulse stability of a pulse train is crucial for large scale scribing. For long term stabilisation an active control system is allowed.
Requirement 8.1.50 | Beam quality (M2) of the IR laser beam ≤ 1.25
Requirement 8.1.55 | Height output beam ≤ 113 mm.
We have standardised our beam path height at 113 mm above the optical table surface. Every laser in our lab is brought with a laser table (supplied by TNO) to this height output. A larger output height would require at least two additional mirrors to bring the laser to the standard height. This standard height is defined by the largest optical mount being the 5 axis beam expander manipulator from OptoSigma placed on a router and rails. All other optics are also brought to this same height.
Requirement 8.1.60 | Laser output beam of IR, green and 4th at same height.
The optics used to combine the three wavelengths beam paths into one beam path are placed in one horizontal plane.
Requirement 8.1.65 | The femtosecond laser source system must have a trigger/sync output for detection system.
The requirement of a trigger output is justified by the need of synchronizing the laser pulses with other devices, such as detectors, PSO and cameras.
Requirement 8.1.70 | All 3 outputs of the femtosecond laser source system must have a linearly polarized output which should be specified, (Linear, vertical; 1 : 1000).
In the optical setup in which the femtosecond laser is intended to be used, the laser beam will be split using a polarizer beam splitters. Therefore the polarization state of the output must be linearly polarized.
Requirement 8.1.75 | Volume available at optical table is limited to maximum length 1100 cm, maximum width 60 cm and maximum height 50 cm.
The laser will replace the laser at the setup and must fit within the available space in the system. The area is cleared by TNO before the installation of the laser.
Requirement 8.1.80 | The femtosecond laser source system is able to operate on non-vibration isolated optical table.
As the predecessor of the laser performs well for more than 10 years on this table.
Requirement 8.1.85 | The offered femtosecond laser source system and associated (necessary) accessories comply at all times with the applicable Dutch and European legal requirements.
These include, (but not limited to): NEN, EN and CE standards and their use including the working conditions - legislation. (93/42 / EEC).
Requirement 8.1.90 | Valid CE marking.
The Tenderer declares that goods are always provided with a valid CE marking upon delivery. The CE marking complies with all current and future European and Dutch legislation and regulations, even if the scope of the legislation and regulations changes in the period between the Contract awarding and the final delivery date.
Requirement 8.3.05 | Factory Acceptance Test (FAT)
The FAT should consist of a test report describing all requirements including:
- Laser output power as function of diode current for the three wavelengths.
- Pulse width measurements.
- Beam quality measurements at specified settings ( 10% , 50% , 100% of maximum laser power) and for all wavelengths.
- Indicative beam profiles for all three wavelengths.
Requirement 8.3.10 | Site Acceptance Test (SAT)
The SAT contains the following checkpoints:
- The femtosecond laser system, will be checked for damages.
- The proper functioning of the equipment will be checked:
o Laser output.
o Electrical safety- and switching functions.
o Single pulse trigger.
- User interface(s), including software.
- All requirements, and all possible additional preferences will be checked.
- The presence and quality of the documentation will be checked.
The specifics of the SAT will be determined collaboratively by TNO and the selected Tenderer. However, the Tenderer is required to include a draft SAT acceptance protocol in their Tender.
SAT ensures that the femtosecond laser operates correctly in its actual working environment. This includes verifying that the installation was done properly and confirm that the laser meets all specified performance criteria in the working conditions. Any issues that arise during the SAT can be addressed immediately by the supplier’s experts. Moreover, SAT provides an opportunity for the supplier to offer on-site training and support, explaining how to operate and maintain the laser effectively.
In the event that the deliverables do not meet the user acceptance criteria for the SAT, the selected Tenderer must document these failures in a written statement, which both parties will sign. The selected Tenderer must then provide TNO with a corrective action plan for these failures within fourteen (14) calendar days from the date of signing the statement. The selected Tenderer must remedy these deficiencies at no additional cost within the timeframe specified in the corrective action plan, but not exceeding three months. A new SAT will then be conducted.
Requirement 8.3.15 | Terms and conditions of payment:
- 70% after delivery at location TNO Eindhoven
- 30% after installation and commissioning On-site and Site Acceptance Test (SAT), approved by TNO, management Solar Technology and Applications and Procurement.
Requirement 8.3.20 | When supplying the femtosecond laser source system, the Tenderer shall provide a Mill Test Certificate (MTC), or similar certificate such that it meets the requirements and obligations to establish the
country of origin (further specified in paragraph 5.2.4).
Requirement 8.3.25 | The Tenderer has the capability to perform remote diagnostics, troubleshooting with the help of on-site TNO personnel and Team-viewer. The tenderer can guide on-site personnel through repairs and maintenance if necessary.
Requirement 8.3.30 | User manual with detailed description of the hardware and software, including any custom made options. Rationale: To provide clear understanding of the system's functionalities, facilitate
effective utilization and help with the troubleshooting of the instrument.
Requirement 8.3.35 | The Contract is to be effected by means of a purchase order, including the Purchasing Conditions for Goods TNO 2022, and according Incoterms Delivered Duty Paid (DDP), 2020. Tenderer must be aware that the Tender will be a contract document if the Contract is awarded to the Tenderer.
Requirement 8.3.40 | US-origin low-technology lasers classified as “EAR99” are free to use within all locations of TNO without further restrictions in end-use(r). In case the femtosecond laser does not classify as EAR99, it depends on the pulse range, the laser source, wavelength of the ECCN specifications which end-use(r) restrictions apply. ECCN should be provided by the reseller or US manufacturer to avoid under- or overclassification.
Requirement 8.3.45 | TNO expects the delivery and installation of the system to take place in the first half of 2026.
Requirement 8.3.50 | TNO expects contractor to agree with the TNO Supplier Code of Conduct (Annex C03).

## Preferences
Annex A04 - Compliance to Requirements - Answers to Preferences Applicable to the European public Procurement Procedure for a femtosecond laser source system, Reference number WS2359773965
The Tenderer demonstrates the ability to meet the requirements by answering Yes/No to the questions, which must be structured in accordance with the requirements given in Chapter 8. Tenders that are incomplete and/or fail to meet all the formal requirements will be declared invalid by TNO and discarded. Missing answers, documents, or data will result in exclusion if there is no rectifiable omission. TNO assesses the Tenders against the Programme of Requirements and Preferences. The Programme of Requirements and Preferences includes Minimum Requirements, Performance Requirements and Preferences. The Minimum Requirements must be met on penalty of invalidity at the time of submission of the Tender, the Performance Requirements only at the time of performance of the Contract and the Preferences must be worked out by Tenderers as part of the award of the Contract.
Preference |  |  | Confirmation of fullfillment of preference Yes/No | Reference in Tenderer's documentation where Tenderer indicates in which way he can fulfil the preference | Remarks
Preference 8.2.05 | Highest pulse frequency (HP Freq). | Validation method: T | Fill in: Yes or No
Max. 70 points
In accordance with requirement 8.1.25, the highest pulse frequency for all three wavelength ≥ 2 MHz.
TNO values a higher pulse frequency of the laser system. As such, the following scores will be awarded:
Assessment criteria: | HP Freq. < 3 MHz: 0 points
3 MHz > HP Freq. < 5 MHz: 50 points
HP Freq. > 5 MHz: 70 Points
Preference 8.2.10 | Highest pulse energy at fundamental wavelength (HP energy). | Validation method: T | Fill in: Yes or No
Max. 70 points
In accordance with requirement 8.1.30 the maximum pulse energy > 0.75 mJ/pulse at IR wavelength.
A higher pulse energy would enable a larger number of laser beam when multiplexed. As such, the following scores will be awarded:
Assessment criteria: | HP energy ≥ 0.85 mJ/pulse: 50 points
HP energy ≥ 1 mJ/pulse: 70 points
Preference 8.2.15 | Fourth harmonic femto laser power output. | Validation method: T | Fill in: Yes or No
Max. 250 points
In accordance with requirement 8.1.05 the power output for the 4th harmonic ≥ 5 Watt. The laser will be used for application research of the 4th harmonic laser wavelength. A larger output is appreciated. As such, the following scores will be awarded:
Assessment criteria: | Power > 9 Watts: 250 points
Power > 8 Watts: 200 Points
Power > 7 Watts: 150 points
Power > 6 Watts: 100 points
Power < 6 Watts: 0 points
Preference 8.2.20 | Electrical power consumption of the laser. | Validation method: C | Fill in: Yes or No
Max. 70 points
TNO values a lower power consumption of the laser system to limit the amount of energy needed to produce solar cells in production. This will also reduce the environmental footprint of the laser.
The answer must contain the power consumption of laser system operation at its highest power output. As such, the following scores will be awarded:
Assessment criteria: | Power consumption < 2.5 kW: 70 Points
Power consumption > 2.5 kW: 0 Points
Preference 8.2.25 | Closed loop water-water cooling system. | Validation method: C | Fill in: Yes or No
Max. 140 points
TNO values that the femtosecond laser has a cooling system without air cooling into the lab. The water-water cooling is preferred to water-air cooling due to noise. The answer should specify the type and specification of the included cooling system. As such, the following scores will be awarded:
Assessment criteria: | Closed loop water-water cooling system without air flow: 140 points
Cooling system based on air flow: 0 points
Preference 8.2.30 | Selection of the wavelength including power up. | Validation method: T | Fill in: Yes or No
Max. 50 points
When wavelength is changed the laser will need some time to adjust to the new setting. During test programs wavelength is used and as the scanner system also has to move to a new position the change can take up to 30 seconds. The complete answer for Preference 8.2.30 must contain the time to switch from fundamental wavelength to the second and fourth harmonic and between the second and fourth harmonic, back-and-forth. As such, the following scores will be awarded:
Assessment criteria: | Selection of wavelength while machine is on < 30 seconds: 50 points
Selection of wavelength while machine is on > 0 seconds: 0 points
Preference 8.2.35 | Internal shutter. | Validation method: I | Fill in: Yes or No
Max. 50 points
The laser is expected to operate safely when enabled, without having to switch off the laser to prevent laser instabilities (e. g. thermal instability). The shutter could be either internal or external. The shutter may be operated using the laser software or using an external hardware/software. The type, working principle and integration of the shutter is valued. As such, the following scores will be awarded:
Assessment criteria: | Shutter available internal system: 50 Points
Shutter available external: 25 Points
No shutter in system: 0 Points

## Sheet1 (hidden)
Click the drop-down box to select
Yes, Tenderer meets the Preference
No, Tenderer does not meet the Preference
Click the drop-down box to select
Yes, Tenderer meets the Requirement
No, Tenderer does not meet the Requirement

## Blad2 (hidden)
Yes
No