## Appendix A05
Appendix A05, Conformity list 'Programme of Requirements' Guidelines for the European Public Tender for an X-ray CT scanner, ref. TNO 2024 FPL-INK 086 dated 11-September-2024
By the following Programme of Requirements TNO will assess if Tenderer meets the relevant requirements as laid down in Section 6 and 8. The Tenderer must indicate for each item, the conformity for the requirements in the table below.
If to any of the requirements, as stated in section 8, are not met by Tenderer, he will be excluded for any further assessment. By answering a question with “compliant” will mean that in the event of awarding the contract, non-conformity to the requirement shall count as failure to meet an obligation from the Agreement to be concluded.
Requirement | Decription | Tenderer is compliant to the requirement | Reference in Tenderer's documentation where compliances with the requirement is described | Remark
General
8.1.1 | The Tenderer ensures that they understand the purpose for which TNO is acquiring the products and associated services and agrees to perform the deliveries and services as an expert in such a way that, at the time of delivery and throughout the maintenance period, the requirements are met. | Click the drop-down box to select
8.1.2 | The Tenderer guarantees that the system and all accessories to be delivered are available on the Dutch market, that it is not a prototype, that it incorporates the latest state of the art technology, and that it includes the most recent version of the relevant hardware and software. | Click the drop-down box to select
8.1.3 | Components of the instrument with a limited expected lifespan can be replaced. | Click the drop-down box to select
8.1.4 | The X-ray CT scanner consists of an X-ray source, a rotation mechanism, a detector, a sample transport mechanism, a reconstruction algorithm, and a control system. | Click the drop-down box to select
8.1.5 | The X-ray CT scanner must be CE certified and compatible with Dutch occupational safety regulations (ARBO). | Click the drop-down box to select
8.1.6 | The X-ray CT scanner must comply with the requirements set out in the Dutch Nuclear Energy Act (ANVS). | Click the drop-down box to select
8.1.7 | Full (hardware and software) accessibility is required for TNO to enable the integration of the CT scanner into a custom-made automated system (in other words, "Engineering access" is required). | Click the drop-down box to select
8.1.8 | The operating language of the system must be English. | Click the drop-down box to select
8.1.9 | The obtained CT scan data remains confidential to TNO and will only be shared with the Tenderer and other parties in consultation with TNO. | Click the drop-down box to select
8.1.10 | Under normal use, the system must remain operational for at least 10 years, during which it must continue to be fit for purpose and meet the agreed specifications. | Click the drop-down box to select
Technical specifications
8.2.1 | Automated sequential and individual CT scans must both be possible. | Click the drop-down box to select
8.2.2 | The 3D reconstruction/scan must cover the entire specimen (with a length of up to 3 meters or more). | Click the drop-down box to select
8.2.3 | The maximum external dimensions of the system are 12 x 4 x 4 meters (l x w x h). | Click the drop-down box to select
8.2.4 | The power supply input must be compatible with the Dutch power grid, for example, 230VAC or 400VAC, 50/60Hz, <16A or <63A. | Click the drop-down box to select
8.2.5 | The energy consumption must be less than 10 kW. | Click the drop-down box to select
8.2.6 | Maximum floor load of 2500 kg/m². | Click the drop-down box to select
8.2.7 | Radiation intensity at the surface and at the specimen input and output must be less than 1 micro-Sv/hour. | Click the drop-down box to select
X-ray source
8.3.1 | The specifications of the imaging system, specifically the source and the detector, must be provided at the time of submission. | Click the drop-down box to select
8.3.2 | Stability of the source voltage to the X-ray tube must be within 1% of the set value. | Click the drop-down box to select
8.3.3 | Variable source voltage must be at least between 60 and 150 kV. | Click the drop-down box to select
8.3.4 | Variable source current must be at least up to 2 mA. | Click the drop-down box to select
X-ray detector
8.4.1 | Bit depth of at least 12 bit. | Click the drop-down box to select
Scan quality and resolution
8.5.1 | 3D resolution of at least 0.5 mm. | Click the drop-down box to select
8.5.2 | The scan quality (resolution, contrast, artifacts) must be sufficient to clearly detect metal fasteners (nails, screws, staples, etc.) with a diameter of 0.5 mm or larger, as well as the wood grain rings (if detectable with the given resolution) within a single scan. | Click the drop-down box to select
8.5.3 | There must be no loss in scan resolution or quality during the first 8 hours of scanning. | Click the drop-down box to select
Specimen specifications and transport
8.6.1 | Maximum specimen weight of 75 kg. | Click the drop-down box to select
8.6.2 | Maximum specimen cross-section of 0.2 x 0.1 m. | Click the drop-down box to select
8.6.3 | Maximum specimen length of 3 meters. | Click the drop-down box to select
8.6.4 | Scanning speed up to 5 meters/min. | Click the drop-down box to select
8.6.5 | Adjustable scanning speed. | Click the drop-down box to select
Reconstruction software
8.7.1 | Visual output files must be exportable to (stacked) images with at least one of the following file extensions: “.tif”, “.png”, “.jpg”, “.bitmap”or similar. | Click the drop-down box to select
Delivery, placement en installation
8.8.1 | Delivery is all-inclusive, covering all costs, including transport, overhead, installation, licenses, training (offered), warranty, maintenance, and any other (unforeseen) expenses. | Click the drop-down box to select
8.8.2 | The delivery, placement, and installation of the system at the location in the Delft region must be completed within a maximum of 12 months after the award of the contract. Faster delivery and installation will be rewarded as described in the preference (8.8.1) below. | Click the drop-down box to select
8.8.3 | The CT scanner must be delivered on call. | Click the drop-down box to select
Training
8.9.1 | On-site training must be provided during and within two months after the installation of the system, covering the system, hardware, and software. Between 2 to 5 (TNO) staff members will attend the training. | Click the drop-down box to select
8.9.2 | Engineers must be demonstrably trained and certified for the system. | Click the drop-down box to select
Acceptation
8.10.1 | After the awardance, a plan will be developed in consultation with TNO by the supplying party detailing how this will be executed (including a timeline, what is needed for installation, facilities, etc.), taking into account the requirements specified in Subsection 2.2.23 Site Acceptance Test (SAT). | Click the drop-down box to select
Warranty, service en documentation
8.11.1 | The system must be delivered with a warranty period of at least 12 months, starting from the date of acceptance of the SAT. This warranty must include parts, labor, and travel costs for engineers. | Click the drop-down box to select
8.11.2 | The system must come with comprehensive documentation for both hardware and software, written in English or Dutch. Documentation must be provided both digitally and in paper format. | Click the drop-down box to select
8.11.3 | On-site service must be available within 1 week of an incident. | Click the drop-down box to select
8.11.4 | A service contract must be available. | Click the drop-down box to select
8.11.5 | Full product support, including maintenance and repairs, must be provided for a period of at least 5 years. | Click the drop-down box to select
Invoicing schedule
8.12.1 | The billing schedule should be as follows: - 30% after the order is placed (with a bank guarantee); - 50% after delivery (with a bank guarantee); - 20% after the completion of installation and acceptance. | Click the drop-down box to select

## validatie (hidden)
Click the drop-down box to select
Yes, compliant to requirement
No, not compliant to requirement
note: dit blad 'verbergen' indien je dit gaat delen intern/extern (behoudens aan inkoop collega's)