## 1. Frontpage
Appendix A04: Conformity list
Belongs to the Public Procurement Guidelines for the European public procurement for the delivery of a Double Climate Chambers
WS2538799081
Date:
Tenderer company name:
Name:
Signature:

## 2. Instructions
Instructions | Tenderer is required to provide this document and a signed pdf version in the offer
Secur | The white boxes are secur and Tenderer can not change these fields.
Click the drop-down box to select | The drop-down boxes within Tab "Conformaty list" Tenderer is requested to select whether Tenderer is compliant to the requirment/ preference or not.
Entry box | The Yellow boxes within Tab "Conformaty list" Tenderer is requested to enter the reference in Tenderer's documentation and own remarks.

## 3. Conformity list
Annex A04, Conformity list 'Programme of Requirements and Preferences' Guidelines for the European Public Tender for Double Climate Chambers, ref. TNO WS2538799081, dated July 30, 2025.
By the following Programme of Requirements TNO will assess if Tenderer meets the relevant requirements as laid down in Section 6 and 8. The Tenderer must indicate for each item, the conformity for the requirements in the table below.
If to any of the requirements, as stated in section 8, are not met by Tenderer, he will be excluded for any further assessment. By answering a question with “compliant” will mean that in the event of awarding the contract, non-conformity to the requirement shall count as failure to meet an obligation from the Agreement to be concluded.
Requirement/ preference | Decription | Tenderer is compliant to the requirement | Reference in Tenderer's documentation where compliances with the requirement or preference is described | Remark
Tenderer is required to fill in #documentname, #page and #paragraph
8.1 | General
Requirement 8.1.1 | The Tenderer ensures that they understand the purpose for which TNO is acquiring the products and associated services and agrees to perform the deliveries and services as an expert in such a way that, at the time of delivery and throughout the maintenance period, the requirements are met. | Click the drop-down box to select
Requirement 8.1.2 | The Tenderer guarantees that the system and all accessories to be delivered are available on the Dutch market, that it is not a prototype, that it incorporates the latest state of the art technology, and that it includes the most recent version of the relevant hardware and software. | Click the drop-down box to select
Requirement 8.1.3 | Components of the instrument with a limited expected lifespan can be replaced . | Click the drop-down box to select
Requirement 8.1.4 | The Climate Chambers consist of (1) a chamber that can resemble outdoor climate condition (e.g. NEN 5060), (2) a chamber that can resemble indoor climate conditions, (3) two façade frames that can be installed between the Climate Chambers, capable of supporting façade and roof components (min. load capacity: 2000 kg), enabling secure installation and sealing of building envelope fragments for realistic testing. | Click the drop-down box to select
Requirement 8.1.5 | The Climate Chambers must be CE certified and compatible with Dutch occupational safety regulations (ARBO) or equivalent. If Tenderer states to equivalent safety regulations, Tenderer adds prove of equivalency. | Click the drop-down box to select
Requirement 8.1.6 | TNO requires unrestricted, real-time access to all data generated by the Climate Chambers, along with export functionality (e.g. CSV format), to enable direct integration into a custom-developed automated system. | Click the drop-down box to select
Requirement 8.1.7 | The operating language of the system must be English. | Click the drop-down box to select
Requirement 8.1.8 | Under normal use, the system must remain operational for at least 10 years, during which it must continue to be fit for purpose and meet the agreed specifications. | Click the drop-down box to select
8.2 | Technical Specifications
Requirement 8.2.1 | The power supply input must be compatible with the Dutch electrical grid, typically 230VAC or 400VAC at 50Hz, with current ratings below 16A or 63A. Given the potential for high inrush currents, particularly from components like humidity control systems, the appropriate type and rating of circuit breaker must be specified. | Click the drop-down box to select
Requirement 8.2.2 | Maximum floor load of 2500 kg/m². | Click the drop-down box to select
Requirement 8.2.3 | Sustainable refrigerants are part of the new policy within TNO. It must be specified which refrigerant is being used. Requirement: GWP < 150 must be achieved. | Click the drop-down box to select
Requirement 8.2.4 | The maximum external height of all assembled components is 5000 millimetres. The maximum external width of all assembled components is 3250 millimetres. The maximum external length of all assembled components is 5500 millimetres. | Click the drop-down box to select
Requirement 8.2.5 | The internal and external dimensions must be specified, including width, depth, and height. | Click the drop-down box to select
Requirement 8.2.6 | Personnel door: each chamber must have a personnel door. | Click the drop-down box to select
Requirement 8.2.7 | Doors: must be protected against freezing shut. | Click the drop-down box to select
Requirement 8.2.8 | Viewing window: there must be a window for visibility inside, for both climate chambers. | Click the drop-down box to select
Requirement 8.2.9 | Lighting: Both chambers must be equipped with lighting suitable for temperatures ranging from -30 to +50 degrees Celsius, which can be used when chamber is in use. | Click the drop-down box to select
Requirement 8.2.10 | Lighting Requirement: The Climate Chamber must have sufficient lighting to safely perform construction tasks such as the assembly and disassembly of test objects. | Click the drop-down box to select
Requirement 8.2.11 | At least one of the Climate Chambers must be mounted on a rail system (or an equivalent solution), allowing TNO employees to move it. This mobility enables the placement of thicker façade or roof elements, of at least 1000 mm, between the chambers. The façade or roof test element will be mounted in a movable frame, similar to a trolley system, which allows it to be positioned between the chambers. Please describe how the test element is to be installed, fixed or clamped into this movable frame. If the rail and frame system provides sufficient space (≥1000 mm) and mobility, installation and operation can be handled by TNO staff. | Click the drop-down box to select
Requirement 8.2.12 | Cable Routing and Moisture Protection: Cables must be easy to replace. Cable entry points must be sealable in an airtight manner and designed to minimize moisture ingress. While heat loss is a consideration, the primary concern is the accumulation of moisture when the chamber operates at freezing temperatures. This moisture can freeze, potentially disrupting the proper functioning of the installation. | Click the drop-down box to select
Requirement 8.2.13 | Power Supply: There must be power available on the inside of both chambers with 220V two (2) double outlets and passages to the outside. | Click the drop-down box to select
Requirement 8.2.14 | The total system must be designed to be sealed airtight with minimal effort. It must be possible for two people to complete the airtight sealing process within approximately one hour, using a straightforward and repeatable procedure. The supplier must propose a design solution to achieve this. | Click the drop-down box to select
Requirement 8.2.15 | The tenderer must ensure that the heat transfer through the biobased façade test object, mounted on the façade frame, is accurately measured. It must be guaranteed that no heat loss occurs through the outer walls of either of the climate chambers, heat transfer should occur only through the test specimen. The tenderer must describe how this will be achieved. | Click the drop-down box to select
Requirement 8.2.16 | The Climate Chambers must be capable of supporting the performance of test methods as described in the following standards and guidelines: – ISO 8990 – ASTM C1363 – ISO 12567 | Click the drop-down box to select
Requirement 8.2.17 | Water Outlet Positioning: Water discharge points (e.g., from dehumidifiers or condensate drains) must be positioned high enough to allow gravity-based drainage into the TNO building’s existing system without the need for pumps. | Click the drop-down box to select
8.3 | Climate Chamber A (Outdoor)
Requirement 8.3.1 | The system must be capable of maintaining an adjustable temperature range from -30°C to +50°C with a control accuracy of ±0.5°C. This range is required to support durability testing under various climatic conditions. Any deviations from the setpoint must remain within the specified tolerance and may not persist for longer than 2 hours. The supplier must indicate the system's temperature control strategy and confirm compliance with these stability. | Click the drop-down box to select
Requirement 8.3.2 | Heating and Cooling Speed: At least 5 degrees per hour. | Click the drop-down box to select
Preference 8.3.3 | Simulated Solar Radiation: the objective is to test solar radiation on the test specimen in the façade frame. To achieve higher temperatures and simulate the effects of solar radiation, the system should include a removable heat source that can realistically replicate solar gain on building components. The method of heat generation is open and may be proposed by the supplier; the energy consumption is not a limiting factor. The system must be designed so that the heat source can easily be installed and removed from the chamber. The supplier is expected to provide a proposal, including the required power capacity (in kW), to achieve effective and realistic solar simulation. | Click the drop-down box to select
Requirement 8.3.4 | Fire Safety and Overtemperature Protection: The system must be equipped with a temperature sensor located near or on the control panel or critical components. This sensor must trigger a warning if the temperature exceeds 70°C. If the temperature rises above 80°C, a safety mechanism, such as an automatic system shutdown or power cut-off, must be activated to prevent fire risk or equipment damage. The supplier must describe the safety logic and response actions in case of temperature escalation. | Click the drop-down box to select
Requirement 8.3.5 | The Climate Chamber must be capable of maintaining relative humidity (RH) levels within the range of 10% to 90%, depending on the internal temperature. The following conditions apply: - Below 2°C: Active humidity control is not required. However, the supplier must estimate the expected drying performance per hour under these conditions. At 2°C, an RH equivalent of approximately 10 g/m³ is acceptable. - Above 20 °C up to 35 °C: The system must provide full RH control across the 10% to 90% range. A reference chart based on NEN 5060 is included below to illustrate the desired combinations of temperature and RH. The supplier must: - Specify the achievable RH range across the entire temperature spectrum. - Describe the control strategy and system capabilities for meeting these criteria. - Provide an estimate of the largest expected air leakage and associated thermal loss, under controlled conditions. | Click the drop-down box to select
Preference 8.3.6 | The reference graph based on NEN 5060 that is provided above (Requirement 8.3.5) illustrates the desired combinations of temperature and relative humidity (RH) that the Climate Chamber must be capable of achieving. The tenderer must: - Indicate which RH–temperature combinations can be reliably achieved by the proposed system. - Create a comparable chart (in a similar format and layout) that clearly shows the achievable combinations within the operational range of the system. | Click the drop-down box to select
Requirement 8.3.7 | The system must be capable of both humidifying and dehumidifying the air. It must be able to maintain the target RH and temperature conditions even when air leakage occurs, up to a rate of 40 litres per minute. | Click the drop-down box to select
Requirement 8.3.8 | Change in Relative Humidity (RH): The RH must be able to change by at least 20% per hour. | Click the drop-down box to select
Requirement 8.3.9 | RH must not be measured based on the wet bulb temperature, but capacitively. RH must be measured in a manner that does not require user intervention. | Click the drop-down box to select
Requirement 8.3.10 | Air Pressure Control and Leak Testing: The Climate Chamber must allow the test object (e.g., façade or building envelope component) to be subjected to controlled overpressure and underpressure to enable the detection of air leakage through joints, penetrations (e.g., electrical outlets), or unintended openings. The system must: - Allow adjustable pressure differentials across the test object (e.g., ±50 Pa or more). - Ensure airtight sealing of the frame and chamber boundaries during testing. - Be compatible with standard air permeability or airtightness test methods, and support classification in accordance with EN 12207 (e.g., pressure steps of 50 Pa, 100 Pa, 150 Pa, and 600 Pa). This allows for proper evaluation and classification of the test object's airtightness performance. | Click the drop-down box to select
Requirement 8.3.11 | Rainwater Simulation: The system must include a method for simulating (wind-driven) rain to assess the test object's resistance to rainwater infiltration under realistic conditions. The simulation must replicate the impact of rain. The tenderer must describe in detail how this simulation is achieved, including: - The method of water application (e.g. nozzles, spray angle, droplet size), - The integration of airflow or wind pressure, - The control of water flow rate and direction, - And any provisions for drainage or water collection. | Click the drop-down box to select
8.4 | Climate Chamber B (Indoor)
Requirement 8.4.1 | Temperature and Humidity Control Requirements: The temperature inside the Climate Chamber must be adjustable between 0°C and 50°C. Relative humidity (RH) must be controllable within a range of 10% to 90%, depending on temperature. - Below 5°C: The RH does not have to go below 20%. - Above 20 °C up to 35 °C: The system must provide full RH control across the 10% to 90% range. The supplier must specify the achievable RH range across the full temperature spectrum and describe how humidity control is managed under low-temperature conditions. | Click the drop-down box to select
Requirement 8.4.2 | Heating and Cooling Speed: At least 5 degrees per hour. | Click the drop-down box to select
Requirement 8.4.3 | Change in Relative Humidity (RH): The RH must be able to change by at least 20% per hour. | Click the drop-down box to select
Requirement 8.4.4 | Relative humidity must be measured and regulated using capacitive sensors, not based on wet bulb temperature. The system must be low-maintenance, requiring no routine user maintenance to ensure long-term accuracy and reliability. | Click the drop-down box to select
Requirement 8.4.5 | Energy Monitoring and Passive Performance Assessment: The system must provide clear and accessible readout of the total energy input required to maintain the chamber at the set temperature. This is essential for testing the thermal and moisture performance of façade elements. Additionally, it must be possible to disable active humidity control, allowing observation of natural moisture and heat transfer through the test object. | Click the drop-down box to select
8.5 | Façade frame
Requirement 8.5.1 | Dimensions of the Façade Frame: The frame must accommodate prefabricated façade elements with dimensions ranging from a minimum of 1200 mm x 1200 mm (W x H) to a maximum of 2200 mm x 2600 mm (W x H). | Click the drop-down box to select
Requirement 8.5.2 | It must carry a weight of 2000 kg, and structurally strong enough to support a facade, roof, or floor. | Click the drop-down box to select
Requirement 8.5.3 | The frame must move via swivel wheels for maximum movement flexibility. The façade frame must stand independently: This is a separate frame in which the facade is placed and air tight mounted. | Click the drop-down box to select
Requirement 8.5.4 | Number: There must be two façade frames, so that one can be prepped while the other is used during a test. | Click the drop-down box to select
Requirement 8.5.5 | Sensors and Cabling: The system must allow for easy integration of sensors and cables within the structural frame to enable accurate measurement of the test object. The sensors must be capable of measuring temperature, relative humidity (RH), and heat flux. The design must facilitate straightforward installation, removal, and repositioning of these sensors without compromising the airtightness or thermal performance of the system. | Click the drop-down box to select
Requirement 8.5.6 | See: 8.1.4 - Two structural façade frames, to be installed between the Climate Chambers, must be capable of supporting façade and roof components (minimum load capacity: 2000 kg), allowing for the secure installation and sealing of building envelope fragments for realistic testing. | Click the drop-down box to select
8.6 | Software
Requirement 8.6.1 | Camera: Each chamber (A and B) must be equipped with a camera. A screen must be mounted on the outside of the chamber and be easily connectable to a transmitter for remote monitoring. This allows observation of internal conditions, including lighting status. | Click the drop-down box to select
Requirement 8.6.2 | Thermal Imaging Requirement: TNO requires a thermal imaging camera as part of the system to support the visual assessment of temperature distribution across the test specimen. This is essential for detecting thermal bridges, insulation defects, or unexpected heat loss during testing. The supplier must ensure that the system allows for easy integration and positioning of the camera for both indoor and outdoor climate sides, if needed. | Click the drop-down box to select
Requirement 8.6.3 | Temperature and RH Control: Must be able to be controlled both manually and automatically. | Click the drop-down box to select
Requirement 8.6.4 | Data from the actual measured room temperature and RH can be stored and uploaded to the computer in CSV file format. | Click the drop-down box to select
Requirement 8.6.5 | Temperature and RH Sensors in the chambers: These must be in the form of separate probes, so they can be moved within the room and positioned as close as possible to the test object. | Click the drop-down box to select
Requirement 8.6.6 | Visual output files must be exportable to videos with at least one of the following file extensions: “.avi”, “.mp4”, “.wmv” or similar. | Click the drop-down box to select
8.7 | Safety
Requirement 8.7.1 | The Double Climate Chambers must be equipped with at least one emergency stop which is easily accessible to the operator. | Click the drop-down box to select
Requirement 8.7.2 | The Double Climate Chambers must have a safety switch that prevents unintended operation of the Double Climate Chambers outside testing periods. | Click the drop-down box to select
Requirement 8.7.3 | The Double Climate Chambers must have a warning light that indicates at least the following distinctive states: “in operation” and “not operating”. The warning light is clearly visible from at least 5 meters distance from any angle and from the operator’s position. | Click the drop-down box to select
Requirement 8.7.4 | The Double Climate Chambers must be compliant to “Directive 2006/42/EC on machinery” also known as the “Machinery directive”. | Click the drop-down box to select
Requirement 8.7.5 | The Double Climate Chambers must operate safely without any manual intervention after the operator has started a test. | Click the drop-down box to select
8.8 | Delivery and after-sales
Requirement 8.8.1 | After the final award, a plan will be developed in consultation with TNO by the supplying party detailing how this will be executed (including a timeline, what is needed for installation, facilities, FAT, SAT, etc). | Click the drop-down box to select
Requirement 8.8.2 | All documentation accompanying the Double Climate Chamber must be delivered in English. The documentation consists at least of an operational manual, error codes including possible solutions by the operator, and technical drawings of the machine. | Click the drop-down box to select
Requirement 8.8.3 | The supplier must perform a Factory Acceptance Test (FAT) as well as a Site Acceptance Test (SAT) together with representatives from TNO as final stage of the delivery. | Click the drop-down box to select
Requirement 8.8.4 | The Double Climate Chambers must be delivered on call. The delivery, placement, and installation of the system at the location in the Delft region must be completed within a maximum of 12 months after the award of the contract. | Click the drop-down box to select
Requirement 8.8.5 | The Double Climate Chamber must have a warranty period of at least 2 years starting from the SAT. | Click the drop-down box to select
Requirement 8.8.6 | Supplier must provide an on-site basic training of at least 1 day within a month after the SAT. Supplier adds information about this to the supplier’s submission to the tender. Between 2 to 5 (TNO) staff members will attend the training. Engineers must be demonstrably trained and certified for the system. | Click the drop-down box to select
Requirement 8.8.7 | The supplier provides an advanced maintenance contract for the Double Climate Chamber for 5 and 10 years after approved SAT which states the maintenance activities and costs. The supplier describes its calamity procedure within and outside the warranty period including response time and complaint registration. The supplier clearly describes the differences between the all-inclusive maintenance contract and the standard maintenance contract as mentioned in paragraph 1.7. Maintenance performed by a third party appointed by the supplier is allowed. | Click the drop-down box to select
Requirement 8.8.8 | Delivery and installation of the Double Climate Chamber is done by the supplier. | Click the drop-down box to select
8.9 | Export Control and TNO Supplier Code of Conduct 2024
8.9.1 | The supplier must immediately notify TNO if it is prescribed by export control laws and regulations within the context of the Agreement that an export or import license is required from a government or any other government body or if the execution of the Agreement is otherwise restricted or prohibited by export control laws and regulations. 'Export control laws and regulations' as referred to in this article should be interpreted as all EU and USA laws and regulations applicable at the time for the regulation of export and import of military goods and dual-use goods (together also referred to as 'strategic goods') and chemical substances and also countries or persons against which sanctions have been imposed by the Security Council of the United Nations (SCUN), by the European Union (EU) or the Dutch government. | Click the drop-down box to select
8.9.2 | Tenderer confirms itself to the TNO Supplier Code of Conduct 2024 (annex C04) | Click the drop-down box to select

## validatie (hidden)
Click the drop-down box to select | Click the drop-down box to select | Click the drop-down box to select
Yes, compliant to requirement | Yes, compliant to preference | Yes
No, not compliant to requirement | No, not compliant to preference | No
No answer
note: dit blad 'verbergen' indien je dit gaat delen intern/extern (behoudens aan inkoop collega's)