## Validation method
Validation method
Test (T) | The operation of the system, or part of the system, using instrumentation or other special test equipment to collect data for later evaluation.
Demonstration (D) | The operation of the system, or a part of the system that relies on observable functions, does not require the use of the instrumentation, special test equipment, or subsequent analysis.
Analysis (A) | The processing of accumulated data obtained from other qualification methods. Examples are reduction, interpolation, or extrapolation of test results.
Inspection (I) | The visual examination of system components, documentation, etc.
Certification (C) | A declaration by a designated stakeholder, usually the supplier or developer.
Review (R) | Detailed review of the specification received from the Tenderer of the equipment .
Moment
Tender Doc (TD) | Compliance to requirement should lead from documentation submitted during the tender process.
Customer Demo (CD) | Compliance to requirement should be confirmed during Customer Demonstration at vendor site (or other site where similar system is installed and accessible) after all documentation has been submitted
FAT | Compliance to requirement should be confirmed at Factory Acceptance Test
SAT | Compliance to requirement should be confirmed at Site Acceptance Test

## Requirements
Program of Requirements Plasma 3 - Polymer RIE etch If to any of the requirements, are not met by Tenderer, The Tenderer will be excluded for any further assessment. By answering a question with “compliant” will mean that in the event of awarding the contract, non-conformity to the requirement shall count as failure to meet an obligation from the Agreement to be concluded.
Requirement | Type | Description | Validation method & moments | Compliant to requirement? | Reference in Tenderer's documentation where compliance with the requirement is demonstrated
Etch Specification
R-5000-005 | BCB (Benzocyclobutene) Etch with Resist mask | The following specifications shall be achieved: - Depth: 2000..5000 nm; - BCB Etch speed (bulk etch): >500 nm/min; - BCB Etch speed reproducibility (run to run): <3%; - Selectivity: ≥1.1:1 (BCB: Resist); - Wall Angle: 95°..130°; - Non-Uniformity : full map depth uniformity < 3%; - Surface roughness: <10 nm; - Residues/defects/by product redeposition: No significant etch residues that cannot be cleaned with soft chemical cleaning shall remain on etched surface or mask layer; - Process-stop by end-point detection(EPD). | R - TD, D - SAT | select
R-5000-010 | BCB (Benzocyclobutene) slow etch | The following specifications shall be achieved: - Etch Thickness: ≤ 10 nm; - BCB Etch speed: 10..50 nm/min; - BCB Etch speed reproducibility (run to run): <3%; - Non-Uniformity: full map depth uniformity < 3%; - Surface roughness: < 5 nm; - Residues/defects/by product redeposition: No significant etch residues that cannot be cleaned with soft chemical cleaning shall remain on etched surface or mask layer. | R - TD, D - SAT | select
R-5000-015 | Stripping- O2+Halogen based | The following specifications shall be achieved: - Resist Etch speed: 60..120 nm/min; - Non-Uniformity: full map depth uniformity < 2%; - Surface roughness: < 5 nm; - Residues/defects/by product redeposition on metal surface: No significant etch residues that cannot be cleaned with soft chemical cleaning shall remain on etched surface or mask layer. | R - TD, D - SAT | select
R-5000-020 | Descum- O2 based only | The following specifications shall be achieved: - Resist Etch speed: 60..120 nm/min; - Non-Uniformity: full map depth uniformity < 2%; - Surface roughness: < 5 nm; - Residues/defects/by product redeposition on metal surface: No significant etch residues that cannot be cleaned with soft chemical cleaning shall remain on etched surface or mask layer; - Photo pattern / pattern feature dimension changes: < 5 nm. | R - TD, D - SAT | select
System scope
R-5000-025 | Gas Handling system | The equipment shall have a gas stick with Mass Flow Controller (MFC) for the following process gases: 1. Argon 2. Chlorine 3. Tetrafluoromethane (CF4) 4. Oxygen 5. Sulfur hexafluoride (SF6) 6. Spare 7. Spare The equipment shall have an automatic gas shut off valve to prevent leaks. The equipment shall have a purging system (N2) that ensures safe chamber venting and prevents unwanted gas residues. | R - CDR | select
R-5000-030 | Mass Flow Controllers (MFC's) | Each gas line shall be individually controlled by a dedicated MFC. The MFC's shall be compatible with the specific gas type and shall support the flow range required by the process. The exact MFC type will be discussed during design review. | R- CDR | select
R-5000-035 | Process control | The equipment shall have process control and data-logging of at least: - Source power Generator with Low and High-Power control; - Platen power generator with Low and High-power control; - Wafer Stage Temperature Control; - APC valve system to control pressure for process repeatability; - End point detection (EPD) and Process Monitoring (Data logging & alarms); - Wall and Shower head Temperature Control. | R - TD, CDR, D - SAT | select
R-5000-040 | Chuck | The equipment shall have a chuck with the capability of controlling chuck temperature between 0..100 °C | D - FAT, SAT | select
R-5000-045 | EPD (end point detection) | The equipment shall have and EPD system that has: - The ability to detect the EPD signal from an open area of 10%..90%; - The capability to reprocess (simulate) the EP data with different EP conditions. | D - SAT | select
R-5000-050 | Waferless Auto Clean (WAC) | The equipment shall have the possibility to perform wafer less cleaning after every wafer run. | D - SAT | select
R-5000-055 | Recipe editor | The following specifications shall be achieved: - The software shall be able to create etch recipe with cyclic settings; - The software shall allow to have over etching step with different. Chemistry based on previous etch step's end point detection time. | D - SAT | select
Productivity Requirements
R-5000-060 | Throughput process | The equipment shall have a throughput of 7 wafers per hour, assuming an average process time of 490 sec. The Tenderer shall show comparable process times for a representative process to ensure the required throughput can be met. | R - TD, D - SAT | select

## validatie (hidden)
select
Yes
No