## Sheet1
Annex D - Quality Criteria
Tenderer:
Upload this document in Excel-format with your Tender in TenderNed. It is not allowed to upload this document in PDF or any other format than Excel.
The yellow cells must be filled in by the Tenderer.
Means of proof
Tenders must provide documentation with the quotation to show that the offered solution meets the minimum requirements and/or there is surplus offered on the minimum requirements. The purpose of this table is to avoid misunderstandings in the verification process we carry out on the minimum requirements and the extent to which the surplus on the minimum requirement is offered. For this reason, please fill out the form carefully. Factsheets are preferable to brochures, but this in itself does not affect the result of the verification.
# | Criteria | Description | Evaluation criteria | Max. score | Your offer includes | Your score | Your fictitious deduction in € | Name of the document | Page/Paragraph
1. Wishes regarding measurement techniques in UHV: Scanning Tunneling Spectroscopy (STS)
1 |  | The possibility of measuring height as a function of bias, z(V). | Included: € 4.500 (3 points) Not included: € 0 (0 points) | 3 |  | 0 | 0
2 |  | The possibility of applying a bias pulse (of at least +- 10V) in order to condition the tip. Ideally this bias pulse can be applied while the tip is removed from the measurement location, while the tip can be automatically brought back to the measurement location after the reconditioning has taken place. | Included: € 12.000 (8 points) Not included: € 0 (0 points) | 8 |  | 0 | 0
2. Wishes regarding measurement techniques in UHV: Atomic Force Microscopy (AFM)
3 |  | A ‘non-contact’ tapping mode in which minimal adhesive interaction between tip and sample takes place while high resolution (topographical) measurements can be performed as described in requirement 42. | Included: € 9.000 (6 points) Not included: € 0 (0 points) | 6 |  | 0 | 0
4 |  | Calibration of the torsional stiffness of the cantilever through the thermal noise spectrum. | Possible: € 9.000 (6 points) Not possible: € 0 (0 points) | 6 |  | 0 | 0
5 |  | Possibility of applying bias larger than +- 10 V during the measurements described in requirement 45. | Possible: € 16.500 (11 points) Not possible: € 0 (0 points) | 11 |  | 0 | 0
6 |  | Possibility of applying bias larger than +- 100 V during the measurements described in requirement 45. | Possible: € 9.000 (6 points) Not possible: € 0 (0 points) | 6 |  | 0 | 0
7 |  | Possibility of measuring, within an array of points, in a single measurement, adhesion, stiffness and KPFM (see requirement 48). | Possible: € 7.500 (5 points) Not possible: € 0 (0 points) | 5 |  | 0 | 0
8 |  | KPFM measurements (see requirement 48) through frequency modulation. | Included: € 6.000 (4 points) Not included: € 0 (0 points) | 4 |  | 0 | 0
9 |  | The possibility of programming measurements, for example a script within which the normal force is linearly decreased or increased during a scan line in contact mode. Another example could be to keep the tip stationary in contact for a controllable time during a line scan, before completing the stroke (slide-hold-slide, see for example https://doi.org/10.1038/nature10589). Finally, another example would be control over the contact time that is applied before the tip is retracted for an adhesion measurement. | Possible: € 22.500 (15 points) Not possible: € 0 (0 points) | 15 |  | 0 | 0
3. Wishes regarding the optical microscope
10 |  | Possibility in the software to select a location on the sample, as seen through the microscope, and to automatically navigate the tip to this location for measurements. | Possible: € 16.500 (11 points) Not possible: € 0 (0 points) | 11 |  | 0 | 0
4. Wishes regarding tips, samples and connections to the system
11 |  | Storage of tips and samples as described in requirement 58, for a larger number of tips and samples. For each extra storage slot € 1.500 will be deducted, with a maximum of € 9.000. | Storage slots as described in requirement 58: €0 (0 points) 1 extra storage slot: €1.500 (1 point) 2 extra storage slots: € 3.000 (2 points) 3 extra storage slots: € 4.500 (3 points) 4 extra storage slots: € 6.000 ( 4 points) 5 extra storage slots: € 7.500 (5 points) 6 extra storage slots: € 9.000 (6 points) | 6 |  | 0 | 0
5. Other wishes
12 |  | Control over the system through the pc to engage and retract the tip, to coarse position the sample and to align the laser and cantilever for cantilever based AFM. | Included: € 9.000 (6 points) Not included: € 0 (0 points) | 6 |  | 0 | 0
13 |  | Delivery of the SPM system, as described in the program of requirements, in a period shorter than 12 months after purchase agreement. For every month of reduced delivery time € 4.500 will be deducted, with a maximum of € 27.000. | Delivery in 12 months: € 0 (0 points) Delivery in 11 months: € 4.500 (3 points) Delivery in 10 months: € 9.000 (6 points) Delivery in 9 months: € 13.500 (9 points) Delivery in 8 months: € 18.000 (12 points) Delivery in 7 months: € 22.500 (15 points) Delivery in 6 months or sooner: € 27.000 (18 points) | 18 |  | 0 | 0
14 |  | Tenderer has an externally audited environmental management system, such as the ISO-14001 certificate (or equivalent). | Tenderer has ISO-certificate (or equivalent): € 4.500 ( 3 points) Tenderer does not have ISO-certificate (or equivalent): € 0 (0 points) | 3 |  | 0 | 0
15 |  | Ergonomical operation of the instrument: A person with a height of 1.60cm should be able to operate the setup and transfer samples/tips without additional tools/aids and in a natural posture. | Yes: € 3.000 (2 points) No: € 0 (0 points) | 2 |  | 0 | 0
16 |  | Low-noise operation: at regular operating conditions of the vacuum setup and the microscope the noise level should be below 50dB. | Yes: € 1.500 (1 point) No: € 0 (0 points) | 1 |  | 0 | 0
17 |  | Warranty period during which the supplier guarantees functioning of the equipment and solves any issues without additional cost. A minimum warranty period of one year is required. € 3.750 will be deducted for each additional year of warranty, with a maximum of € 15.000. | 1 year of warranty: € 0 (0 points) 2 years of warranty: € 3.750 (2,5 points) 3 years of warranty: €7.500 (5 points) 4 years of warranty: € 11.250 (7,5 points) 5 years or more of warranty: € 15.000 (10 points) | 10 |  | 0 | 0
18 |  | Pressure interlocks that switch off devices and set voltages to zero once pressure thresholds are exceeded. | Included: € 7.500 (5 points) Not included: € 0 (0 points) | 5 |  | 0 | 0
19 |  | Protection of the vacuum level in the case of power outages, for example by valves that close automatically. | Included: € 3.000 (2 point) Not included: € 0 (0 points) | 2 |  | 0 | 0
|  |  | Total score | 128 |  | 0 | 0
|  |  | Price on offer (exl Vat) |  |  | 0
|  |  | Price per quality point |  |  | #DIV/0!

## Sheet2 (hidden)
1 | Included | 3
| Not included | 0
2 | Included | 8
| Not included | 0
3 | Included | 6
| Not included | 0
4 | Possible | 6
| Not possible | 0
5 | Possible | 11
| Not possible | 0
6 | Possible | 6
| Not possible | 0
7 | Possible | 5
| Not possible | 0
8 | Included | 4
| Not included | 0
9 | Possible | 15
| Not possible | 0
10 | Possible | 11
| Not possible | 0
11 | Storage slots as described in requirement 58 | 0
| 1 extra storage slot | 1
| 2 extra storage slots | 2
| 3 extra storage slots | 3
| 4 extra storage slots | 4
| 5 extra storage slots | 5
| 6 extra storage slots | 6
12 | Included | 6
| Not included | 0
13 | Delivery in 12 months | 0
| Delivery in 11 months | 3
| Delivery in 10 months | 6
| Delivery in 9 months | 9
| Delivery in 8 months | 12
| Delivery in 7 months | 15
| Delivery in 6 months or sooner | 18
14 | Tenderer has ISO-certificate (or equivalent) | 3
| Tenderer does not have ISO-certificate (or equivalent) | 0
15 | Yes | 2
| No | 0
16 | Yes | 1
| No | 0
17 | 1 year of warranty | 0
| 2 years of warranty | 2.5
| 3 years of warranty | 5
| 4 years of warranty | 7.5
| 5 years or more of warranty | 10
18 | Included | 5
| Not included | 0
19 | Included | 2
| Not included | 0