## PvE
Programma van Eisen Openbare procedure aanschaf GC-IRMS
Yes
No
Toelichting: In het programma van eisen staan alle eisen omschreven voor de Openbare procedure aanschaf GC-IRMS. U dient in de kolom D, middels "Yes" of "No" aan te geven of u aan deze eis kunt voldoen. Alle eisen, behalve Informatieve vragen 13.1 en 13.2, worden beschouwd als knock-outcriteria. Dit betekent dat u, indien u niet aan de eis voldoet of slechts voorwaardelijk voldoet, wordt uitgesloten van de aanbesteding. Als er om nadere toelichting wordt gevraagd, moet er een PDF-file worden opgevoerd met een korte uiteenzetting of bewijslast die toeziet op de eis. De antwoorden mogen gebundeld worden in één PDF-file.
1. GENERAL Requirements
1.1 | The supplier agrees to the draft purchase agreement.
2. REGULATORY COMPLIANCE Requirements
2.1 | Qualification work books/PDFs must be available that can be used by a suitably qualified engineer to record in detail the installation, operation and periodic servicing of the mass spectrometer. These must be used as part of a qualification service provided by the vendor
2.2 | Qualification work books/PDFs must be available that can be used by a suitably qualified engineer to record in detail the installation, operation and periodic servicing of the GC system. These must be used as part of a qualification service provided by the vendor
2.3 | An IQ (Installation Qualification) is part of the installation and is performed by the Supplier
2.4 | Adequate training and familiarization of all the operators of the stable isotope lab is part of the agreement and included in the costs
2.5 | The supplier is willing to measure samples (supplied by the lab) to assess quality of the offered system
3. GENERAL TECHNICAL Requirements
3.1 | The mass spectrometer must be a magnetic sector IRMS
3.2 | Enrichments of hydrogen, carbon, nitrogen and oxygen must be measurable, and the set-up must allow for on-line determination of 13C/12C and D/H isotope ratios in organic mixtures separated by gas chromatography
3.3 | The supplier must have demonstrated applications or proven use of the system in a medical setting
3.4 | The laboratory is equipped with infrastructure for the following gases: He 5.0, N2 6.0, CO 4.7, CO2 4.5, O2 5.0, as well as compressed air; however, no cryogen infrastructure is available. Pressure for all gases is 5 bar. The proposed system must be fully compatible with and operable using this existing setup.
3.5 | Ability to remotely start sequences, troubleshoot and adjust parameters such as pressures without being physically present at the unit
3.6 | Instrumental parameters of the system are stable in a working environment with temperature fluctuations of ±5 °C/24hours
3.7 | The system must be capable of analyzing large batches of complex biological extracts/derivates without maintenance or performance degradation
3.8 | The system must have the same capabilities and serve as a back-up of our current GC-IRMS (Delta V) installation and vice versa. Sensitivity and specificity must be identical or improved.
3.9 | The system offered must have dimensions that are compatible with the actual site of the lab, footprint of an instrument (including computer and monitor, and peripherals) should fit within a rectangle of 3.25 m (length) x 1.10 m (depth)
3.10 | The instrument must be deliverable to the lab for installation through a standard door opening, with a maximum width of 89 cm.
3.11 | If not a floor-standing model, suitable furniture must be included in the offered quote.
3.12 | The supplier must designate a single contact responsible for installation, maintenance, troubleshooting, and support for the entire system (autosampler, GC, interface, IRMS).
3.13 | Since we already have a Thermo Finnigan TCEA in the lab, it must be possible to connect this to the new configuration, both hardware-wise and software-wise, with equal specifications for the stable isotope measurements.
3.14 | Service and spare parts must be available for at least the next 10 years after acceptance of the instrument
3.15 | The supplier shall provide a starter pack of consumables essential for the initial installation, operation, and routine troubleshooting of the system. This must include all commonly replaced items such as septa, ferrules, O-rings, inlet liners, seals, and other consumables required for injector, detector, and ion source maintenance
4. AUTOSAMPLER Requirements
4.1 | The number of positions in autosampler for liquid samples must > 100, for exetainer breath tubes > 90
4.2 | The autosampler must be able to perform both head space and liquid injection. Conversion between the configurations must be simple.
4.3 | The autosampler must have the capability of automatic tool exchange
4.4 | The liquid sampler must be able to inject volumes ranging from 0.1 µl to 10 µl at a controlled and programmable speed. Headspace injections up to 100 µl.
4.5 | There must be an adequate number of reservoirs for several different types of washing fluids, with sufficient rinsing volume to allow up to 4 days without the need to supplement the solvents
4.6 | It must be possible to perform head space injections from exetainer breath tubes
4.7 | Fully integrated control of autosampler must be provided within the MS acquisition software
5. GC Requirements
5.1 | The GC must be equiped with a PTV or MMI, and SSL injection system
5.2 | It must be possible to install a different type of inlet after initial installation of the GC
5.3 | Range of temperature control for both column and inlet > 30-350°C
5.4 | The system must allow for column exchange without requiring full venting or extended cool-down periods. It must return to a stable, operational state within max. 30 minutes after a column change.
6. Continuous flow interface Requirements
6.1 | The IRMS must use continuous flow techniques, and is capable of pyrolysis as well as combustion experiments
6.2 | The configuration must include a fully automated continuous flow interface, with the possibility to simultaneously connect at least 5 reference gasses
6.3 | At least two peripherals must be connected simultaneously, and switching must be fully software controlled
6.4 | It must be possible to automatically dilute samples from EA and GC, without manual intervention, in a range 1-50
6.5 | Water needs to be removed by a permeable membrane/nafion tubing
6.6 | The interface must have a gas save mode for all reference gases, so that there is no consumption of reference gases which are not needed for the analysis method
7. IRMS Requirements
7.1 | The IRMS must be equiped with an EI source
7.2 | The precision of enrichment measurements at natural enrichment levels must be 0.2 per mille (in delta) absolute
7.3 | The mass resolution must be greater than 110m/Δm (10% valley)
7.4 | The IRMS must be capable of automatic mass jump within analysis (e.g. between H and O measurement using TCEA)
7.5 | The tuning and mass calibration of the mass spectrometer must be automatic and software-controlled
7.6 | The system stability must be better than 10 ppm on mass scale
7.7 | The precision of the reference/monitoring gas must be at least: 13C (CO2) 0.06 ‰ for 5 nA signal of 1.5V (SD 1 for 10 consecutive pulses); 18O (CO2) 0.08 ‰ for 5 nA signal of 1.5V (SD 1 for 10 consecutive pulses): tested and confirmed on site
7.8 | The linearity of the reference/monitoring gas must be at least: 13C (CO2) 0.15 ‰ SD 1 for 0.03-0.3 nmol (2-20 nA); 13C (CO2) 0.02 ‰ for 0.03-0.9 nmol (2-60 nA): tested and confirmed on site
7.9 | Memory is virtually eliminated by construction and design of the ion source and vacuum system
8. DETECTOR Requirements
8.1 | The system must have the possibility of automatic adjustment of H3+ factor
8.2 | The system must have the possibility of automatic re-analysis of H3+ factor
8.3 | The m/z cup 3 must have a filter to eliminate 4He+ ions
8.4 | The H3+ factor must be smaller than 10 ppm/nA, and the stability better than 0.03 ppm/nA/hour
8.5 | The sensitivity in continuous-flow mode must be minimally 1100 molecules per ion
8.6 | The detector must have a measurement range of at least 50 V
8.7 | The detector must operate without loss of overall sensitivity for > 10 years during normal practice and maintenance. If this requirement could not be met, a realistic span of lifetime and costs for replacement must be specified in the offer and will be scored accordingly. Tenderawarding will be on basis of total cost of ownership. Costs for detector replacements need to be specified in the pricing sheet. These numbers need to be realistic, therefore we will hold you to the offer, i.e. if the detector needs to be replaced more times than specified, this will have to be without additional costs.
9. VACUUM SYSTEM Requirements
9.1 | The roughing pump must have a longevity/warranty of >10 years for during normal practice and maintenance. If this requirement could not be met, a realistic span of lifetime and costs for replacement must be specified in the offer and will be scored accordingly. Tenderawarding will be on basis of total cost of ownership. Costs for pump replacements need to be specified in the pricing sheet. These numbers need to be realistic, therefore we will hold you to the offer, i.e. if the detector needs to be replaced more times than specified, this will have to be without additional costs.
9.2 | The turbo pump must have a longevity/warranty of >10 years for during normal practice and maintenance. If this requirement could not be met, a realistic span of lifetime and costs for replacement must be specified in the offer and will be scored accordingly. Tenderawarding will be on basis of total cost of ownership. Costs for pump replacements need to be specified in the pricing sheet. These numbers need to be realistic, therefore we will hold you to the offer, i.e. if the detector needs to be replaced more times than specified, this will have to be without additional costs.
9.3 | The instrument must have a fail-safe vacuum system protection. It must automatically shut-down after e.g. power failures
9.4 | The roughing pump(s) must be dry/oil free vacuum pump(s).
10. SOFTWARE/COMPUTER Requirements
10.1 | The software must have direct and easy data transfer to popular processing programs such as Excel, Word, Powerpoint
10.2 | The MS acquisition software must allow full system control over the GC
10.3 | Fully integrated control of the MS must be provided in the MS acquisition software
10.4 | Vacuum readbacks and system vent/pump cycles must be digitally monitored and controlled
10.5 | The MS control software must be a windows based platform and have the ability to control the MS, GC and autosampler devices
10.6 | The system must be able to operate uninterrupted under conditions where the internet and/or the linked external servers are not available or become unavailable during an analytical run
10.7 | The computer must have at least enough computational power for smooth data processing during real time data monitoring
10.8 | The software license must allow for installation and simultaneous use of multiple PC's or a server that will primarily be used for processing and reporting purposes. In practice, the license must allow for software installation on 10 different PCs with a minimum of 5 simultaneous users at a time.
10.9 | The software must be compatible with installation in a citrix/one view environment
10.10 | Costs for software and licenses are included in the package offered
10.11 | All hard- and software upgrades and updates that are neccesary for the system warranty on the performance are installed free of charge during the term of this agreement
10.12 | After installation of a new release of the software or OS, all previous records are completely available
11. INSTALLATION Requirements
11.1 | The offered equipment must be able to meet the Acceptance criteria described in the Acceptance prototol.
11.2 | Specify the electrical requirements, including required voltage (230–380V), plug type | Answer by PDF upload
12. SERVICE PLAN Requiremets
12.1 | Response time from notification via telephone/e-mail/on-line is within 8 business hours with a suitably qualified engineer.
12.2 | On-site response time from notification is within 10 business days.
12.3 | Fix time from notification is within 15 business days.
12.4 | An uptime of at least 95%, based on 9 hours of operation per day, 260 working days per year.
13. Informative questions
13.1 | The supplier provides a document that demonstrates the daily energy consumption in kilowatt-hours (kWh). The document must specify the energy consumption in operational mode (after the system has been started up and equilibrated) and include the energy consumption of the interface, IRMS and pumps. Due to the complexity of calculating the GC's energy consumption and its dependency on the temperature gradient, this component can be ignored in this document. | Answer by PDF upload
13.2 | The supplier provides a document demonstrating the helium consumption [in liters] under the specified standard conditions: - In Operation Mode: Column flow: 1 ml/min (constant flow); Runtime: 40 minutes; Column specifications: 30m length, 0.25mm internal diameter, 0.50µm film thickness; Interface operation flows must be included. Report the helium consumption in liter per run - In Standby Mode: Column flow: 1 ml/min; Column specifications: 30m length, 0.25mm internal diameter, 0.50µm film thickness; Interface operation flows must be included. Report the helium consumption in liter per hour | Answer by PDF upload
Ondergetekende rechtsgeldig vertegenwoordiger van Inschrijver verklaart deze naar waarheid te hebben ingevuld. Inschrijver verklaart akkoord met het Programma van Eisen. Op het moment van Inschrijving kan Inschrijver voldoen aan de gestelde eisen, welke gelden als Knock Out.
Naam | [invullen]
Functie | [invullen]
Onderneming | [invullen]
Handtekening | [invullen]
Plaats en datum | [invullen]
✓