## Explanation
1 |  | Program of Requirements Document
Automation of hiPSC culture
Overview of Categories:
| 1 | General Conditions
| 2 | Automation of hiPSC Culture
| 3 | Other(IT, Service & Purchase)
| 4 | Wishes
MUST The Tenderer its offer must unconditionally meet the requirements set by the LUMC for the performance of the contract. This means that all requirements must be answered with “yes”. When the LUMC concludes that the Tender does not comply with the requirements, the offer will be rejected and the Tenderer will be excluded from any further participation (knock-out). WISH (AND DESCRIBE) Wishes must be answered to potentially earn points. If no points are scored, this will not immediately lead to exclusion of the Tenderer. The answer to each whish should be submitted in this document and/or separately as evidence documents, whereby each answer must contain a clear reference to the Wish to which answer applies. Be short and concise in your answer. Wishes are assessed in accordance with the method indicated in the resprective wish. The number of points that can be achieved for each Wish is visible under the category ''Wishes''.

## General
1. GENERAL
To be filled in by Tenderer
Pos | Specification | Must | Wish | Describe / Specify | YES | NO | Comments/Proof
| The platform must be designed in such a way that future compatibility with ISO 9001 certification is possible | X
| The analysis equipment / manufacturer / supplier and consumables must demonstrably comply with the European directive (EU) 2017/746 when it comes into force. | X
| The platform should have at minimum a biosafety level I enclosure providing sample and user protection and fully tested to (BS EN 12469 2000 – Performance criteria for microbiological safety cabinets). | X
| The platform enclosure and instruments are compatible with UV decontamination. |  | x | Give specifications
| The enclosure should provide sufficient space to accommodate future upgrades and the addition of extra equipment. See Future upgrades for list of apparatus we would like to integrate in the future. |  | x | Please elaborate and add proof of concept
| It would be desirable to have real time visual feedback of both process and the safety status. |  | x
| This platform is aimed at automating cell line development of gene-edited human induced pluripotent stem cell (hiPSC) lines | x
| The ability to automatically culture undifferentiated hiPSCs including medium changes, coating of cell culture plates with extracellular matrix, passaging/expansion and potentially freezing of hiPSCs. | x
| The platform should be able to preform the workflow as indicated in Tab "Workflow" | x |  | Please add configaration and explanation how workflow is succesfully preformed on platform
| The platform should be able to refresh and/or expand a minimum of 24 hiPSC lines with a maximum of three 6-well plate or two 96 well plate per line simultaneously within 24 hours (see Appendix workflow). | x |  | Please add configaration and calculation of platforms theoretical througput based on the workflow
| The system must be able to handle multiple hiPSC lines simultaneously without any risk of cross contamination or mixup. | x |  | Provide documentation
| The platform should include at minimum following devices: Liquid handler, Incubator, Robot arm, Cytena UP.SIGHT (supplied by LUMC), (BSL-I) Enclosure | x |  | Please add list of equipment
| The platform must be an integrated whole of the beforementioned devices. The Workflow is controlled from a single operating system. | x
| Preferably, the platform should include a cooled storage for culture reagents. |  | x | Please add list of equipment
| The system must be designed for long-term reliable and continuous operation over a timespan of at least 12 hours without operator support. | x |  | The vendor must submit documented proof, such as certifications, test reports, or case studies, to demonstrate compliance
| Ideally, the platform should be able to reliably run autonomously for at least 60 hours, and during that period it can be monitored remotely by the lab personel. |  | x | The vendor must submit documented proof, such as certifications, test reports, or case studies, to demonstrate compliance
| The BSL-I enclosure airfiltration system should maintain a 100% coverage of the entire workarea, minimising risk of contamination entering the system. And allowing cell culture without addition of standard antibiotics. | x
| The platform should be a recirculating system that does not require external connection to the site HVAC. | x
| The platform should have internal environment monitoring with full real time data logging of temperature & humidity. |  | x
| The platform must have an illuminated visual alert system that indicates the system status | x
| The platform should have anti-microbial worksurfaces to reduce potential sources of contamination. | x
| The platform should have a large portal near the UP.SIGHT to enable operator access and standalone use without impacting on the internal environment enclosure | x |  | Provide documentation
| The platform should have a large portal near the liquid handler to enable operator access and standalone use without impacting on the internal environment enclosure | x |  | Provide documentation
| The platform should include bins for waste tips, liquids and labware, equipped with an operator alert to indicate when they are full. | x
| The platform should have integrated power and service connections, including a main power panel for single-phase connection, a central mains switch, and accessible fuses. | x
| The plaform should have one or more incubator(s)(37C, 5%CO2) with the capacity of at least 60 standard multiwell plates. | x |  | Give specifications
| The platform must include drivers and control interfaces for selected instruments | x
| The Individual instruments must be able to operate as stand-alone in the system without interference to the entire platform | x
| The individual instruments must be reachable for usage without interference or interruption to the platform. | x
| The platform is equipped with a PC and monitor to provide an interface for system control, data processing, and user interaction. | x
| Vendor must include installation and support of the system at LUMC, Mirai House location | x
| Vendor should provide a full design review including hardware and software specifications | x
| Vendor will provide system training to 4 employees of the Leiden hiPSC Centre Team | x
| Vendor will provide user manuals for the complete system | x
| The vendor must perform a Site Acceptance Test (SAT) to validate that the system meets all performance specifications and safety requirements. | x
| The maximum weight of the platform is 500kg/m2 | x |  | Provide documentation
| The hight of the platform cannot be more than 2400mm; the system will fit inside the space indicated in Red of tab "Lab Location", leaving space for maintanance and operator acces. | x |  | Provide documentation
| Transportation/ installation of single components via standard size elevators/stairs/doors should be possible | x |  | Provide documentation
| The platform should be powered by a 400VAC, 32A, IEC60309 power plug or similar | x |  | Provide documentation
| System cabelling should be managed, labelled separately. | x
| The system should ensure that plates that are being moved are not dropped during a brief power interruption. | x
| It should be possible to integrate artificial intelligence or learning technologies to make culture decisions based on relevant data |  | x | Give examples how you have done this in the past
| We envision the platform to be able to report on phenotype quality by measuring the ratio of undifferentiated cells versus spontaneously differentiated cells for each well. |  | x | Give explanation how
It would be desirable to have remote monitoring and control capabilities of the platform and individual instruments.
The system should avoid waste of cell culture reagents due to 'empty' volumes.

## Automation of hiPSC culture
2. Automation of hiPSC culture
To be filled in by Tenderer
Pos | Specification | Must | Wish | Describe / Specify | YES | NO | Comments/Proof
| The platform should be able to maintain undifferentiated hiPSCs, including imaging and reporting confluency, should be able to perform medium exchange, and passaging | x |  | Provide documentation
| Required is a platform to automate the following workflows; Dispensing hiPSCs as single cells by use of the Cytena UP.SIGHT into 96 or 384-wells plates, Culturing (refreshing/passaging) hiPSCs in 384, 96, 12 and 6 well format, Harvesting gDNA from hiPSCs by use of QuickExtract. See tab Workflow | x |  | Provide documentation
| The platform should be able to perform passaging of undifferentiated hiPS cells as small, multicellular aggregates ('clumps'). See SOPs. | x |  | Provide documentation
| The platform should be able to perform single cell passaging of undifferentiated hiPSCs. See SOPs | x |  | Provide documentation
| The platform should able to dissociate cells and transfer to 15ml tube or 1,5ml reaction tubes for further manual handling (see workflow). | x
| The platform should have the capacity to work with an overlap of similar workflows that are not started on the same day | x |  | Provide documentation
| The system should be able to continuously expand hiPS lines in a structured manner, ensuring a continous workflow during the expansion of many lines. | x
| The platform should automate manual tasks, enabling parallel operations that enhance both throughput and quality. It should also provide improved control and monitoring to ensure process consistency and repeatability. | x
| The system must be designed for long-term reliable and continuous operation over a timespan of at least 12 hours without operator support. | x |  | Provide documentation
| The system must be designed for long-term reliable and continuous operation over a timespan of at least 10 weeks without interference due to system maintenance. | x |  | Provide documentation
| The automation platform must be flexible and scalable to meet both current and future research needs. The ability to rapidly reconfigure (ongoing) workflows and system layouts will be an important advantage. | x
| The system should be designed for scalability to enhance throughput and support growing operational needs. |  | x | Give examples
| There must be an audit trail of events in the system. | x
| Should record telemetric data from the automated platform, including temperature and humidity measurements. |  | x
| Should offer Gantt charts to show process progress. |  | x
| The user interface should be logically structured and require a training of maximal three days for an average technician. | x
| Should be able to accommodate error handling on the fly without impacting on the entire process runs. | x |  | Provide documentation
| The System must notify users via Email, App or similar that system is finished with certain steps and user should take over with manual handling. (see Workflow) | x
| The liquid handerler gripper arm must be able to work with both standard multiwell plates and standard tubes, or automatically switch (without manual user intervention) between different types of gripper-fingers (for plates, tubes etc.) | x
| The pipetting arm of the liquid handler must support a pipetting range from 1 µl to 2000 µl without requiring manual exchange of parts. | x
| The pipetting arm of the liquid handler should be capable of dispensing volumes with a minimum of 1 µl in free-dispense mode with high consistency and reliability, ensuring precise and reproducible liquid handling for demanding applications. | x
| The pipetting arm of the liquid handerler must have a protection measure, to ensure no liquid can contaminate/damage and/or enter the pipetting channel. | x
| The liquid handler should ensure high accuracy and reliability in liquid handling and be able to automatically monitor and detect incorrect volumes. | x
| The platform has tilt stations to support efficient medium refreshment and passaging of plates. | x
| The platform should have one or more temperature blocks allowing warming of culture plates up to 37C | x
| The supplier must ensure that the liquid handler has sufficient capacity to accommodate all the lab culture ware required for performing every expansion step outlined in the desired Workflow. | x
| It would be desirable for the liquidhandler to have a separate UV lamp for decontamination |  | x
| Preferably the liquidhandler has an integrated automated thermal cycler |  | x
| The liquid handler software should be user friendly, launch any protocol clearly, quickly, easily, and intuitively | x
| The liquid handler deck can be optimally reached by the robot arm. | x
| It would be desirable for the platform to be capable of cold storage of medium/reagents for at least 60 hours. |  | x
| It would be desirable for the platform to be capable of storing half-liter aliquots of medium. |  | x
| It would be desirable for the platform to be able to warm up half-liter aliquots of medium to room temperature within 30 to 60 minutes before medium refreshment or passaging cells. |  | x
| The platform should include a robot arm. | x
| The robot arm must be collaborative and safe to use | x
| The robot arm must be easy to configure new labware and instruments on the system | x
| The robot arm should have 360° usability to optimize workflow | x
| Robot arm must be able to grip: Standard multi-well plates, with and without lids. 96-well PCR plates | x
| The robot arm should be able to perform cross movements to evenly distribute the hiPSC clumps after passaging in the well. | x
| The robot arm should have smooth movement profile to avoid uneven attachment of hiPS clumps because of movement. | x
| The robot arm should be fully capable of accessing all critical positions on the liquid handler's deck, ensuring efficient handling of time-consuming tasks like the passaging of hiPSCs. | x
| We wish for the robot arm to include integrated sensors or feedback mechanisms that can detect and recover from errors autonomously, minimizing downtime. |  | x
| We wish for the platform to support adaptive scheduling, allowing the robot arm to dynamically adjust tasks and priorities based on real-time workflow requirements. |  | x
| We wish for a compact, modular robot arm design to allow for easy upgrades, scalability, and adaptation to different laboratory configurations. |  | x
| We wish for the robot arm to be compatible with a wide range of laboratory instruments, enabling diverse applications without additional customization. |  | x
| The plaform should have one or more cell culture incubator(s), temperature 37C and 5% CO2 | x
| The incubator(s) should have a total storage of minimum 60 multiwell plates. | x
| The incubators should have adjustable temperature (30C-40C) and O2(1%-25%) |  | x
| The incubators must have a decontamination protocol and proven to be contamination free. | x
| The incubators must be accessible to the robot arm to load and unload plates. | x
| The incubators must be accessible to manually remove plates from the system as well as for cleaning/ decontamination. | x
| Removing/ moving of plates from/into the incubator does not cause a change of position for other plates (possibly negatively affecting cellular attachment) or prolonged periods of changes in temperature and/or CO2 |  | x
| The incubator is seamlessly welded, corrosion-resistant, and equipped with a stainless steel chamber with rounded edges for easy and thorough cleaning. |  | x

## Other
3. Other
To be filled in by Tenderer
Pos | Specification | Must | Wish | Describe / Specify | YES | NO | Comments/Proof
| The software must provide a well-structured and visually clear interface for selected instruments accessible via dedicated software, web browsers and/or mobile devices. | x
| The software must Include tools for graphical planning of samples, devices, and resource utilization, alongside real-time simulation capabilities. | x
| The software must provide the ability to define and adjust protocols during (ongoing) experiments, automatically managing changes across all subsequent steps. | x
| The software must ensure the system can seamlessly scale to accommodate growth in users and workload without compromising on speed or efficiency. | x
| The software must incorporate mechanisms to recover ongoing experiments and samples even in the event of critical failures like power outages. | x
| The software must automatically re-start after a power cycle. | x
| The software must enable creation, assignment, and modification of tasks and events, with resource tracking, conflict resolution, and prioritization options. | x
| The software must deliver immediate updates on schedule changes with notifications to relevant users. | x
| The software must support generation of tailored reports and dashboards, alongside detailed performance metrics and time tracking. | x
| The software must allow protocols to be done with diverse labware and ajust protocols automatically. | x
| The software must include option to pause new processes while allowing in-progress tasks to conclude efficiently. | x
| The Software must provide seamless import/export options for data, scripting Application Programming Interfaces (APIs) for system integrations. | x
| The software must enable direct input of experiment and curated lists into the software for streamlined processing. | x
| Software updates and new versions of the software required for the operation of the device are provided free of charge until end of life. | x
| Operators must be able to independently develop scheduling protocols after installation without help from Vendor. | x
| It would be desirable to add (AI-based) optimization to adjust schedules in real-time for efficiency, resource usage, or cost reduction. |  | x
| Data import/export capabilities, supporting multiple file formats such as CSV, Excel, XML, PDF etc. |  | x
| Ability to set up different scheduling views (e.g., daily, weekly, resource-specific). |  | x
| Intuitive, user-friendly interface for scheduling, task prioritization, and easy access to features. |  | x
| Features for tracking and reporting on costs associated with scheduling (e.g., overtime, resources used, equipment). |  | x
| Cost optimization algorithms for reducing operational expenses through efficient scheduling. |  | x
| User support portal with FAQs, troubleshooting guides, and helpdesk integration. |  | x
| Regular software updates and patches to ensure the system is up to date. | x
| The instrument can be remotely controlled. |  | x | Please elaborate in which form: Wired Remote Control, Wireless Remote Control, Software-Based Control, App-Based & Mobile Control
| It is desirable the warranty begins from the moment the equipments are successfully integrated and operational at Leiden hiPSC Centre. | x
| In the event of a system failure, it is desirable the vendor responds within 2 hours and resolve critical failures within 24 hours. |  | x | Please elaborate
| The operator receives remote warnings. |  | x | Explanation/Example
| Evidence should be provided for availability of spare parts and required installation time. | x |  | overview/table
| Spare parts should be available for a minimum of 10 years after installation. | x |  | Please elaborate
| The vendor is expected to provide detailed suggestions for a comprehensive maintenance strategy (executed by LUMC, due to budgetairy constraints for maintenance we do not expect to get a full service contract), including the frequency and scope of routine maintenance (e.g., monthly, quarterly, or annually) to ensure optimal platform performance, a preventative maintenance plan designed to minimize downtime and extend equipment lifespan, and a system for maintaining accurate logs of all maintenance activities, including repairs and inspections. | x |  | Please elaborate
| The vendor is required to provide an escalation model. This model must clearly outline the steps to be taken when escalating issues or incidents, including the parties involved, timelines, and responsibilities. | x |  | Please elaborate
| It would be desirable that the vendor provides support for equipment and software even after the end-of-life period, with clear terms for continued service or parts availability (equipment supplied by the automation specialist). |  | x | Please elaborate
| It would be desirable if the vendor can provide any upgrade or replacement conditions if equipment becomes obsolete or incompatible with future technological developments (equipment supplied by the automation specialist). |  | x | Please elaborate
| The tenderer shall perform a release run/system check after each repair, relocation or change to the software, after which the supplier shall officially release the system for use. | x
| A user forum/community to exchange experience, tips and tricks is present |  | x | Please elaborate
| The System can be approched by other systems by use of DNS name resolving | x
| The system makes use of DHCP | x
| The system fuctions on the basis of dynamic IP-adresses | x
| The System makes use of central LUMC network services (DHCP, DNS, NTP, Radius, AD, Mail, Printen, etc) No own network services | x
| The system makes us of the lumc Infrastructure (no own switches, accespoints, routers, firewalls etc) | x
| Network cable connenction must be a minimal of 200 Mb Full duplex | x
| In case of WiFi connection, on the basis of 802,1x with WPA2 en AES with a minimal signal strength of -55DB | x
| The system supports IPv4 en IPv6. | x
| The system functioned on a routed network and does not use broadcast | x
| The system does not make use of VPN Tunnels | x
| The system has well functioning Anti-virus software that is updated indefitnely. | x
| The system has well functioning operating system that is kept up-to-date | x
| The system updates will be done via LUMC-netword (no USB, CD, DVD or other fysical system) | x
| Client software on the system is and stays up-to-date. | x
| Data is stored as minimally and shortly as possible on the system. | x
| The system will keep functioning when the ICT infastructure failes (network, Servers, etc) | x
| Acces to the system is only possible by authorised personel. Authentication via central AD directory service. | x
| The system can be tested/scanned for vulnerabilities without impact on the system. | x

## Workflow and SOP
General workflow of each project
| P+1 | P+2 | P+3 | P+4 | P+5
| On day 0 (electroporated) hiPSCells are manually inserted into the UP.SIGHT to be dispensed (directly!) into prepared 384 well plate. | 96 colonies from 384-well plate are replica splitted into two 96-well plates. Inbalanced replica splitting of each clone; 25% in plate A and 75% in plate B. | Up to 12 colonies are passaged to 12-well plate. Each clone in a separate well. | Clones are passaged to 3 wells of 6-well plates. 3 clones, Each clone will have its own plate. | Each Clone is passaged to 3 wells of a 6well plate.
When dispensing is done, plates are directly transferred to incubator.
| Day 4, 6, 8 and 10 plate is imaged using the UP.Sight Bottom imager. Pictures uploaded to C.studio software | Plate A is imaged with UP.SIGHT bottom imager 2, 4 and 6 days after passaging. | Plate is imaged with UP.SIGHT bottom imager on day 4, 6 and 8 after passaging. | Plates are imaged with UP.SIGHT bottom imager on day 5 and 7 after passaging. | Plates are imaged with UP.SIGHT bottom imager on day 5 and 7 after passaging.
| Day 2 after dispensing medium is added, afterwards refreshing of medium every 2 days | Plates are refreshed every 2 days, beginning the day after replica-splitting. | 12-well plate is refreshed every 2 days, beginning the day after passaging. | 6-well plates are refreshed every 2 days, beginning the day after passaging. | 6-well plates are refreshed every 2 days, beginning the day after passaging.
|  | 2-3 days after passaging, plate B is harvested with Quick-Extract and removed from system for genetic screening |  | on day 6 after passaging 1 well is dissociated and transferred to individual 1,5mL or 15mL tubes for further analysis (DNA isolation) | on day 6 after passaging 1 well is dissociated and transferred to individual 15mL tubes for further analysis (FlowCytometry Analysis)
| Colonies are passaged when >70% confluentcy is reached. | Clones from plate A are passaged when genetic screening is complete and selected wells have >70% confluency. | When clones reach <70% confluency. Based on morphology 3 clones are selected for passaging. | On day 7, 1 well of each clone is selected for passaging based on morphology
| Leftover clones on 384-well plate are discared | Leftover clones on 96-well plate are discared | 12-well plate is ejected from system for further manual handling(Freezing) | After passaging, plates are ejected from system for further manual handling (Freezing) | On day 7, plates are ejected from system for further manual handling (Freezing)
SOP Passaging of hiPSCs
Clump passaging (6- and 12-well plate)
1. | Aspirate culture medium
2. | add 1ml/well Gentle cell dissociation reagent
3. | incubate at RT for 5 min
4. | remove GCDR
5. | add 1ml/well culture medium
6. | mechanicaly dislodge cells(Scraping or pipetting)
7. | Transfer cell suspension to new plate containing culture medium. Dilution 1:10(100uL)-1:50(20uL)
8. | Shake plate in cross motion(forward-back+left right) for even distribution of cells
9. | Transfer plate(s) to incubator 37C
Singe Cell passaging (384- and 96-well plate)
1. | Aspirate culture medium
2. | Wash with PBS-/-
3. | Add 40uL/well Accutase
4. | Incubate at 37C for 7 min
5. | Add 60uL/well culture medium
6. | Completely dislodge cells by pipetting repeatetly.
7. | Transfer Cell suspension to new plate
a. 384 to 96 well --> Transfer 25 uL to plate A, Transfer 75 uL to plate B
b. 96 to 12 well --> Transfer 50uL to 12 well plate, discard rest
8. | Transfer plate(s) to incubator 37C
SOP Quick Extract DNA purification
1. | Remove (96-well) plate from incubator
2. | Aspirate Culture Medium
3. | Wash cells once with PBS, Aspirate PBS.
4. | Add 30µL of Quick-Extract solution to each well
5. | Incubate at 37C for 5 min.
6. | Repeatedly pipet up and down to lyse the cells
7. | Transfer 30uL to 96-well PCR plate(Cat#???)
If Heatblock is included in setup: Place PCR plate in Heatblock and run protocol
If no heatblock is included in setup: Place plate in hotel and notifiy user that plate is ready for pick-up

## Lab Location
Red | Space available instrument
A | 5,90m
B | 4,30m
C | 3,80m
D | 3,00m
E | 1,70m
Total area:
~23m2
B
C
A
D
E

## Descriptive text
Background
The Leiden hiPSC Centre LUMC is part of the hDMT INFRA StemCells infrastructure. The hDMT INFRA StemCells infrastructure serves to be used to conduct research and promote innovations related to human stem cells.
An important component of hDMT INFRA’s capacity is realized by the (robotics-based) automation of several core services, including the mass generation of ‘healthy’ and ‘diseased’ human induced Pluripotent StemCells(hiPSCs). Lines generated from diseased donors will be genetically modified to remove disease mutations. These isogenic lines that can be used as controls in disease models. Via automation, we will be able to substantially improve the efficiency of our services and to standardize our services.
Our protocol consists of transfecting cells with Cas9 and HDR template. When the cells have recovered they will be cloned by use of the UP.SIGHT(Cytena). Cloning of the transfected lines is a necessity to achieve a genomic pure population. The clones are screened and correctly modified clones are expanded further for full QC analysis and freezing of stocks.
The current throughput limitation is the cell culture during the cloning and expansion phase of the project. During cloning multiple 96-well plates have to be combined to a single plate and replica split. Automating this step might make it possible to work with 384-well plates, necessitating less plates and culture media. This also is true for the expansion phase of the culture. By a more consistent way of culture less material will be lost and less consumables will be needed in general.
This Program of Requirement is based on a budget of 970.000 incl VAT given as a Grant by NWO for the project hDMT INFRA StemCells