## Sheet1
No. | About | Question | Answer
1 | General tab PvE | The platform must be designed in such a way that future compatibility with ISO 9001 certification is possible. How will this be assessed? Upgrades and improvements to the system will be possible. | The applicant must demonstrate/proof: - ISO 9001:2015 certification, OR - An equivalent quality management system Evidence: - ISO 9001 certificate, OR - Quality management system description which addresses at least: * Quality policy and objectives * Process monitoring and control * Deviations and complaints procedure * Internal audits and management reviews * Continuous improvement process
2 |  | The analysis equipment / manufacturer / supplier and consumables must demonstrably comply with the European directive (EU) 2017/746 when it comes into force. Our integrated solutions are designed for Research Use Only. While individual components may carry IVDR certification, the complete system is not certified as a whole under IVDR. Would this meet your requirements? | Yes
3 |  | The enclosure should provide sufficient space to accommodate future upgrades and the addition of extra equipment. See Future upgrades for list of apparatus we would like to integrate in the future. The enclosure requirements will vary depending on the specific equipment involved. We were unable to locate your list of apparatus intended for future integration. Could you please provide this information? | Will be provided in Tender when published
4 |  | The platform should have anti-microbial worksurfaces to reduce potential sources of contamination. Would this be acceptable if it only applies to the liquid handler? | All surfaces where plates and media are placed should be anti-microbial. This would mean the liquid handler, but also plate hotels and similar
5 |  | The vendor must perform a Site Acceptance Test (SAT) to validate that the system meets all performance specifications and safety requirements. What are the requirements for the SAT? | Will be provided in Tender when published
6 |  | We envision the platform to be able to report on phenotype quality by measuring the ratio of undifferentiated cells versus spontaneously differentiated cells for each well. Is the image analysis performed by the UP.SIGHT software or is a custom image analysis solution expected? | The UP.SIGHT software can only do confluency measurements. A custom solution would be necessery. Based on either the UP.SIGHT pictures or a secondairy imaging system
7 | Automation tab PvE | The platform should have the capacity to work with an overlap of similar workflows that are not started on the same day Could you provide more details about this requirement? For example, are you referring to batch processing, queue management, or protocol scheduling capabilities? Are you looking for the system to handle multiple instances of the same protocol running concurrently but initiated on different days? What types of workflows need to overlap, and what scheduling or resource management features are you envisioning? | Primarily we are looking for a system that can handle the same protocol(see workflow) initiated on different days. Secondary we would be like to the system to handle different protocols simultaneosly
8 |  | The system should be able to continuously expand hiPCS lines in a structured manner, ensuring a continous workflow during the expansion of many lines. Can you elaborate on this? Are you looking for the system to automatically manage passaging schedules for multiple hiPSC lines at different growth stages? Does “structured manner” refer to standardized protocols, automated decision-making, or coordinated resource allocation across cell lines? How many hiPCS lines do you typically need to expand simultaneously, and what’s the expected throughput requirement? | We would like a system that can culture many different hiPSC lines for multiple weeks without the need for manual handling. Automated decisions based on e.i. confluency would be ideal but not a necessity.
9 |  | There must be an audit trail of events in the system. Our system generates comprehensive log files for tracking operations. Would this satisfy your audit trail requirements, or do you need specific information captured that we should clarify? | What you descibe seem adiquate. We would like logs per project, what happend to what line on which moment. Also error and repair logs.
10 |  | The liquid handerler gripper arm must be able to work with both standard multiwell plates and standard tubes, or automatically switch (without manual user intervention) between different types of gripper-fingers (for plates, tubes etc.) Based on the workflow requirements, direct tube handling may not be necessary. Our system can accommodate tubes by placing them in SBS-format tube adapters on the liquid handler deck. The gripper arm handles all SBS-format labware, so tubes in SBS-compatible racks can be processed automatically. Would this approach meet your requirements? | Yes
11 |  | The pipetting arm of the liquid handler must support a pipetting range from 1 µl to 2000 µl without requiring manual exchange of parts. Could you clarify the necessity for both extreme ends of this range? Liquid handlers are typically optimized for either high precision at low volumes (limiting large volume performance) or large volume dispensing (compromising low-volume accuracy). Which capability is more critical for your workflow, and would you consider separate solutions for large volume dispensing? | Both high(6-well) and low(384-well) volume plates are used in the system. To be able to handle(refresh) both well types in a timely maner this range of pipeting might be neccecery. If other solutions are possible we are open to look at it. Will amend line to reflect possible high and low volume pipetting. Could you give an overview of your solution for review?
12 |  | It would be desirable for the platform to be capable of storing half-liter aliquots of medium. How many different types of media need to be stored simultaneously on the system? | At least 3, preferably 5. And storage for at least 3 smaller volumes(50ml)
13 |  | The robot arm should be able to perform cross movements to evenly distribute the hiPSC clumps after passaging in the well. Only as part of standard approach path to certain instruments. Would this meet your cell distribution needs? | If the robot arm can make a small side-ways and forward-backward(cross-figure) movement before placing the plate in the incubator, then yes.
14 |  | We wish for the robot arm to include integrated sensors or feedback mechanisms that can detect and recover from errors autonomously, minimizing downtime. Our robot arm has no autonomous recovery capabilities. The system can detect errors and provide alerts but requires operator intervention for resolution. Is this level of error handling sufficient for your workflow? | Yes
15 | Other tab | The software must Include tools for graphical planning of samples, devices, and resource utilization, alongside real-time simulation capabilities. Partially, not for individual samples. Would this level of functionality meet your needs? | We would like software that helps with planning of projects. Graphical overview of protocols now running and seeing where there is time to initialise new projects. Simulation helps with adapting excisting protocols and creating new ones and seeing how they effect already initialised/planned projects. Although this is for the system as a whole, single samples/projects should be able to be visualised as part of the whole.
16 |  | The software must provide the ability to define and adjust protocols during (ongoing) experiments, automatically managing changes across all subsequent steps. There is no automatic changing of subsequent steps. We can’t meet this requirement as specified. Are there specific experimental scenarios where mid-experiment protocol changes are essential, or could pre-planned protocol branches address your needs? | Changing of ongoing protocols would be a necessity when cells grow very different then expected. If other solution is possible please give solution
17 |  | The software must incorporate mechanisms to recover ongoing experiments and samples even in the event of critical failures like power outages. Our software does not include mechanisms to recover experiments after critical failures such as power outages. Rather than recovery capabilities, we can provide UPS systems to prevent power-related interruptions. Would this prevention-based approach satisfy your requirement? | The system will be connected to emergency power that will kick-in approximatly 60 seconds after failure. If a UPS system can bridge the gap that would be acceptable.
18 |  | The software must automatically re-start after a power cycle. Similar to the previous recovery requirement, our software lacks automatic restart capabilities after power cycles. The UPS solution we mentioned would also address this concern by preventing power interruptions. Is this acceptable? | The software should be able to pick up protocols at the place it was after a powercycle.
19 |  | The software must enable creation, assignment, and modification of tasks and events, with resource tracking, conflict resolution, and prioritization options. We cannot fulfill this requirement. Is this task management functionality critical for your hiPSC culture workflows, or could external project management tools be used alongside our automation software? | External tools could work, please provide details about this system
20 |  | The software must deliver immediate updates on schedule changes with notifications to relevant users. This functionality is not available in our software.
21 |  | The software must support generation of tailored reports and dashboards, alongside detailed performance metrics and time tracking. This functionality is not available in our software.
22 |  | The software must allow protocols to be done with diverse labware and adjust protocols automatically. This requirement isn’t clear for us. Can you elaborate what you are envisioning, | That when starting a protocol, there can be choice what plates to use. For instance: use 6-well instead of 12 well because colonies were bigger then expected.
23 |  | The software must include option to pause new processes while allowing in-progress tasks to conclude efficiently. This functionality is not available in our software.
24 |  | The software must enable direct input of experiment and curated lists into the software for streamlined processing. This depends on input location and format. Can you elaborate further? | Input of line names and other project details. This would be done on the running PC itself or done remotely via network.
25 |  | The System can be approched by other systems by use of DNS name resolving Not by default | This is a current IT Department criterium, we will discuss with them what other options are possible
26 |  | The system makes use of DHCP Our system does not use DHCP. Is DHCP a mandatory requirement for your network infrastructure, or can static IP addressing work for your setup? | This is a current IT Department criterium, we will discuss with them what other options are possible
27 |  | The system fuctions on the basis of dynamic IP-adresses Our system does not support dynamic IP addressing. | This is a current IT Department criterium, we will discuss with them what other options are possible
28 |  | The System makes use of central LUMC network services (DHCP, DNS, NTP, Radius, AD, Mail, Printen, etc) No own network services Our system operates on a dedicated automation network behind a hardware firewall or gateway PC, separate from central network services. Would this isolated network architecture be compatible with your requirements? | This is a current IT Department criterium, we will discuss with them what other options are possible
29 |  | The system makes us of the lumc Infrastructure (no own switches, accespoints, routers, firewalls etc) See above. | This is a current IT Department criterium, we will discuss with them what other options are possible
30 |  | The system supports IPv4 en IPv6. IPv4 only, IPv6 is not supported. Can IPv4 only operation work in your environment? | This is a current IT Department criterium, we will discuss with them what other options are possible
31 |  | The system will keep functioning when the ICT infrastructure fails (network, Servers, etc) Scheduling: yes, data handling depends on instrument settings and local IT. Is this meeting your requirement? | Yes
32 |  | The system can be tested/scanned for vulnerabilities without impact on the system. No additional software tools supported except Windows Defender. Are external vulnerability scans a mandatory requirement for your IT security compliance, or can Windows Defender-based protection meet your needs? | This is a current IT Department criterium, we will discuss with them what other options are possible
33 | Market consultation document | 6. Questions and Answers (page 6) Item 10: We intend to incorporate a user test in the Tender We would appreciate clarification on your expectations. Our solutions are purpose-built rather than standardized platforms, so we don't have a ready system with all devices listed in your requirements (Excel General tab, row 16). We can offer component demonstrations of our liquid handler and robotic arm technology and software. | Demonstrations of the separate components in different platforms would be sufficient. We are more interested in your answer of the second part of the questions: what objective criteria can be used during the user test?
34 | PvE | The platform should able to dissociate cells and transfer to 15ml tube or 1,5ml reaction tubes for further manual handling (see workflow). Is a 15mL tube fixed, or are there possibilities to change to 6mL tubes that fit in an SBS rack. | If 15mL tubes are not possible, we would prefer 1,5ml reaction tubes(eppendorfs). If also not possible we should have tubes with a V-bottom and are can be centrifuged.
35 |  | The system must be designed for long-term reliable and continuous operation over a timespan of at least 12 hours without operator support. what are the maximum number of plates that are required during a run. How many tips/plates should be stored in plate hotels. | Enough tips to refresh a maximum of 60 plates(either 6 well or 384) and enough plates to passage about 1/3 of that number
36 |  | The liquid handerler gripper arm must be able to work with both standard multiwell plates and standard tubes, or automatically switch (without manual user intervention) between different types of gripper-fingers (for plates, tubes etc.) For what application should tubes be moved? The solution to do so is an expensive option. | If tubes are easily accessable for removal after completion of workstep this would also be acceptable. Other wise tubes should be moved to easily accesable location
37 |  | The pipetting arm of the liquid handler must support a pipetting range from 1 µl to 2000 µl without requiring manual exchange of parts. Why 1 - 2000ul, is pipetting twice with a 1000ul channel not sufficient? | Speed mostly. The pipetting arm needs to move back and forth 12 times to refill one 6-well plate the cells would be without medium for to long(30sec or more) If other solution is possible please let me know.
38 |  | The pipetting arm of the liquid handler should be capable of dispensing volumes with a minimum of 1 µl in free-dispense mode with high consistency and reliability, ensuring precise and reproducible liquid handling for demanding applications. Why is free-dispense mode required? | Dispensing additional components to the media without touching the destination liquid(contamination risk). If other solutions are possible to achive this please provide details
39 | Budget | Is het beschikbare budget een vast gegeven of kan dit eventueel worden uitgebreid? | This budget is a fixed amount.
40 | Marktconsultatie | Ter verduidelijking, jullie verwachten een volledig ingevuld program of technical requirements met bijbehorende data? Of is antwoord op de vragen in hoofdstuk 6. van de marktconsultatie voldoende? Waar en hoe kan dit worden geupload in TenderNet? | We would like the answers to the questions in Chapter 6. The programme of requirements is for information purposes only.
41 | NDA | Zijn vragen zichtbaar voor andere inschrijvers? Hoe wrodt omgegaan met NDA mbt deze marktconsultatie? | The information provided may be used in the tender. If so, it will be anonymized. See also the explanation in the market consultation document. When submitting questions during the Information Note, you can choose to ask individual questions (see also the tender documents later).