## Requirements
Annex A04, Compliance with Requirements - Answers to Preferences Applicable to the European open tender for a Bifacial LED based Solar Simulator, Ref. WS2682288826, dated 3 December 2025
The Tenderer demonstrates the ability to meet the requirements by answering Yes/No to the questions, which must be structured in accordance with the requirements given in Chapter 8. Tenders that are incomplete and/or fail to meet all the formal requirements will be declared invalid by TNO and discarded. Missing answers, documents, or data will result in exclusion if there is no rectifiable omission. TNO assesses the Tenders against the Programme of Requirements and Preferences. The Programme of Requirements and Preferences includes Minimum Requirements, Performance Requirements and Preferences. The Minimum Requirements must be met on penalty of invalidity at the time of submission of the Tender, the Performance Requirements only at the time of performance of the Contract and the Preferences must be worked out by Tenderers as part of the award of the Contract.
Requirement | Description | Confirmation of fullfillment of (minimum) requirement Yes/No | Reference in Tenderer's documentation where Tenderer states his answer to the Preferences | Optional: Remarks
8.1 | Front side illumination light engine
Requirement 8.1.1 | Pulse length 100ms to continuous. | Fill in here: Yes or No
Requirement 8.1.2 | Illuminated area at least 300⨯300 mm2. | Yes No
Requirement 8.1.3 | Standard Test Conditions (IEC 60904-9 2020) A+A+A+: | etc.
· Reference spectrum AM1.5 (IEC 60904-3 2019);
· Spectral Match using Table 2 (300 – 1200nm) between 0.95 and 1.05;
· Non uniformity A+, < 1%;
· Long-term instability, <0.2%;
· Short-term instability Synchronized;
· Spectral deviation (SPD) < 23%;
· Spectral coverage (SPC) > 98%;
· minimum intensity range between 300 and 1200 W/m2.
Requirement 8.1.4 | > 25 types of LEDs with different wavelengths.
Requirement 8.1.5 | The solar simulator must be equipped with spectral tuning, where the individual wavelengths can be controlled between 10 and 100%.
Requirement 8.1.6 | Integrated photo-spectrometer:
· Output calibrated in W/m2/nm;
· Resolution wavelength range 400-1000nm <= 1.5 nm;
· wavelength range 300 – 1200nm;
· possibility for recalibration.
Requirement 8.1.7 | Spatial non uniformity of individual wavelengths < 2%.
Requirement 8.1.8 | AM0 spectrum 1000W/m2 (300 – 1200nm).
Requirement 8.1.9 | Control of the Light Engine via external software (e.g., Python) must be possible.
Requirement 8.1.10 | Minimum lifespan of LED lamp units at least 5 years in laboratory operation assuming 800 hours of use per year.
Requirement 8.1.11 | For longer flashes or light soaking it is desirable that the intensity and spectrum remain stable. Therefore, TNO requires an active feedback loop to control the intensity (and spectrum) during a measurement.
Requirement 8.1.12 | Since LED based light engines only have accurate illumination at a fixed distance from the light engine, and TNO is planning to measure samples with different heights, an indication of the correct height from sample to light engine is a must. Therefore, TNO requires an intuitive system to indicate the measurement plane.
8.2 | Rear side illumination light engine
Requirement 8.2.1 | Pulse length 100ms to continuous.
Requirement 8.2.2 | Illuminated area at least 160⨯160 mm2.
Requirement 8.2.3 | Standard Test Conditions (IEC 60904-9: 2020) AAA.
Requirement 8.2.4 | Minimum intensity range between 200 and 500 W/m2.
Requirement 8.2.5 | Spatial non uniformity of individual wavelengths < 2%.
Requirement 8.2.6 | Control of the Light Engine via external software (e.g., Python) must be possible.
Requirement 8.2.7 | Minimum lifespan of LED lamp units at least 5 years in laboratory operation assuming 800 hours of use per year.
Requirement 8.2.8 | For longer flashes or light soaking it is desirable that the intensity and spectrum remain stable. Therefore, TNO requires an intensity sensor.
Requirement 8.2.9 | Measurement pane (focus) same as front illumination.
8.3 | Measurement system
Requirement 8.3.1 | Synchronized measurement of Voltage, Current, Irradiance and temperature.
Requirement 8.3.2 | 4-wire measurement (Kelvin sensing).
Requirement 8.3.3 | Measurement electronics based either on electronic load or Keithleys: Electronic load: • High current voltage range at least -5 to +5V; • Low current voltage range at least -10 to +50V; • 4 quadrant operation; • At least 16 bit data acquisition; • At least 4 different current ranges lowest range below 25mA; • Maximum current > 16A; • Noise level / SN ratio less than 0.005%. Keithley: Data acquisition with TSP-Link synchronized Keithley measurement equipment: • max current at least 14A for instance 2 ⨯ 2461 in parallel; • Compatible with all 2400 series Source measure units.
Requirement 8.3.4 | Measurement accuracy for both current and voltage below 0.1% of measurement range. Ranges should be laid out such that they comply with the IEC 60904-1:2020 ‘measurement uncertainty of 0.2% or less of the Voc and Isc’. Ranges for Isc between 5 mA and 14A, and Voc between 0.5 and 50V.
Requirement 8.3.5 | Repeatability for Isc and Voc measurement of less than 0.1%.
Requirement 8.3.6 | Electrical connections via banana plugs.
Requirement 8.3.7 | Auxiliary measurement channels TNO would like to have the possibility to record additional signals during the measurement. For instance, to compare a temperature sensor laminated inside the PV device with one taped to the back of the device. Therefore, TNO requires at least two additional voltage measurement channels.
Preference 8.3.1 | Mitigation against bad contacts (max 150 point): Experience with TNOs current solar simulator, equipped with high current 4 quadrant power supply, is in case of one of the wires being interrupted it can destroy devices. Currently there are two mitigation measures in place; contact check for device under test and a fuse in the reference cell connector. The contact check measures the resistances between Voltage and Current leads, with selectable pass/fail threshold values, before each measurement cycle. Another option would be to have a hard limit to the current such as for example the compliance current in a Keithley source measure unit. That is why TNO would like to know how the tenderer deals with the assessment criteria 8.3.1 a, 8.3.1 b and 8.3.1 c. As such, the following scores will be awarded:
Assessment criteria 8.3.1 a: | Separate fused connection for reference cells, to protect the reference cells during the calibration procedure:
· 0 points if there is no fused reference cell connector;
· Or, 20 points for other facility of 250 mA fast fuse for measuring the reference cells;
· Or, 40 points for separate (LEMO) connector for reference cells with a 250 mA fast fuse in one of the current leads.
Assessment criteria 8.3.1 b: | Automated contact check:
· 0 points if there is no contact check;
· Or, 20 points if separate standalone system;
· Or, 60 points if fully integrated in software and measurement routine.
Assessment criteria 8.3.1 c: | Hard current and voltage limit:
· 0 points if there is no provision against current and voltage spikes in the power supply.
· Or, 50 points if the power supply has a selectable limit for current and voltage that can not be exceeded.
8.4 | Measurement software
Requirement 8.4.1 | The LED based Solar simulator must have following measurement modes:
· Mono and bifacial light IV (LIV), single scan forward and reverse and hysteresis
· Dark IV (DIV)
· Mono and bifacial light soaking
o Mppt tracking, Isc, Voc, and fixed bias voltage, at controlled irradiance and temperature
· Suns–Voc (See also Preference 8.4.2 and 8.4.3)
· Variable irradiance (VIM) and variable temperature LIV measurements
· EQE (See also Preference 8.4.4)
Requirement 8.4.2 | Independent control of both light engines.
Requirement 8.4.3 | The LED based Solar simulator must have spectral measurement and adjustment
· Save + load light engine spectra, name of active spectrum should be prevalent in GUI
· Import reference spectrum (eg from an outdoor measurement), and adjust light engine spectrum to best spectral match
Requirement 8.4.4 | The LED based Solar simulator must allow user configurable recipes and procedures:
· Procedure; settings to perform one measurement (eg 10s lightsoak at Voc, 25°, 500W/m2 AM0)
· Save and load procedures (measurement, light engine settings and temperature)
· Recipe; sequence of procedures
· Save and load recipes
Requirement 8.4.5 | The LED based Solar simulator must have a database for reference cells.
Requirement 8.4.6 | The LED based Solar simulator must have hysteresis without breaking voltage and illumination at changeover.
· Perovskite PV cells and modules act like high capacitive devices. Therefore, they don’t like large (rapid) fluctuations in voltage (I=C×dV/dt). When the voltage is broken at the changeover the voltage over the cell will go rapidly from Isc to Voc and back again. Resulting in unwanted transient behaviour and in some cases even destruction.
Requirement 8.4.7 | The LED based Solar simulator must have control over measurement speed (sweep rate, time delay), number of measurement points, current voltage ranges, current limit, start stop voltage, measurement direction.
Requirement 8.4.8 | The LED based Solar simulator must have storage of all parameters and results in ASCI CSV format. Including raw data (of both light engines) for external conversion to STC and further analysis of measurement data, and information of measurement procedure, including metadata such as aperture area, and name of spectrum used.
Requirement 8.3.9 | The LED based Solar simulator must have determination of Voc, Isc, FF, mpp and slope at Voc and Isc integrated in the software so the user has direct feedback.
Requirement 8.4.10 | The LED based Solar simulator must have automated and manual calibration procedure for single junctions and tandem devices:
· Adjust spectrum and intensity based on EQE difference between device under test and reference cell(s).
Preference 8.4.1 | Spectral bands (max 30 point):
TNO intends to use the system with (mono and bifacial) tandem devices. In order to discriminate between subcells in a multijunction PV device the individual cells should be probed. Therefore, user configurable spectral bands combining multiple wavelengths are desired (for use in suns–Voc, VIR, and temperature dependence).
Assessment criteria | · 0 points if non or only one spectral band can be configured;
· Or, 30 points, definition of at least two user configurable spectral bands with multiple wavelengths;
Preference 8.4.2 | Suns–Voc and suns–Isc (max 80 point):
Changing the illumination intensity while measuring the Voc, or Isc.
Assessment criteria | Suns–Voc single junction:
8.4.2 a | · 0 points if no suns–Voc measurements are possible on single junction devices;
· Or, 15 points for suns–Voc single junction (full spectrum);
Assessment criteria | Suns–Voc multijunction:
8.4.2 b | · 0 points if no suns–Voc measurements are possible on multi junction devices.
· Or, 15 points 1 spectral band fluctuating, other band(s) dark;
· Or, 40 points 1 spectral band fluctuating, other band(s) constant controlled illumination (including darkness).
Assessment criteria | Suns–Isc:
8.4.2 c | · 0 points, if no suns–Isc functionality is included;
· And/or, 10 points for suns–Isc for single junction devices (full spectrum);
· And/or, 15 points, for suns–Isc for multijunction devices (with bias voltage).
Preference 8.4.3 | Suns–Voc and suns–Isc, minimum intensity (max 20 point):
For pseudo Fill Factors above 80%, an intensity of less than 10 W/m² is desired. In case this minimum intensity cannot be reached it would be possible to perform multiflash with for example a 10% greyfilter. Provide in your answer the offered minimum intensity in a suns–Voc measurement.
Assessment criteria | · 0 points for minimum intensity >100 W/m²;
· Or, 10 points for minimum intensity between 10 and 100 W/m²;
· Or, 20 points for minimum intensity <10 W/m².
Preference 8.4.4 | External Quantum Efficiency (EQE) (max 150 point):
Assessment criteria | · 0 points if no EQE functionality is offered;
· Or, 25 points single junction (DC controlled 0V);
· Or, 150 points multijunction with bias illumination and voltage.
Preference 8.4.5 | Monitoring reference cells in parallel (max 20 point):
For multijunction, each junction should be set to a matching reference cell. Therefore, an option to monitor the Isc of two or more cells simultaneous would be nice to have. Especially if this can be used with the automated calibration procedure.
Assessment criteria | · 0 points if non or one reference cell can be connected and monitored;
· And/or, 10 points if multiple reference cells can be connected and monitored in parallel (software and/or SMU/DMM);
· And/or, 10 points if parallel monitoring of multiple reference cells is fully integrated in the automated calibration procedure.
8.5 | System
Requirement 8.5.1 | The LED based Solar simulator must have Lighttight housing.
Requirement 8.5.2 | For monofacial measurements the LED based Solar simulator must have a height adjustable removable vacuum chuck:
· At least 10 cm height adjustment below measurement plane;
· Accuracy of height adjustment less than 0.5 mm;
· Minimum size of 300 x 300 mm2;
· Vacuum for a minimum of 3 different substrate sizes, substrate sizes that should fit are 25x25 mm2., 100x100 mm2 ,and 156x156 mm2.
Requirement 8.5.3 | The LED based Solar simulator must have temperature control for the device under test
· Temperature range 20 – 65°C.
Requirement 8.5.4 | The LED based Solar simulator must have freely accessible height above vacuum chuck of at least 300 mm.
Requirement 8.5.6 | The LED based Solar simulator must have CE Marking and other relevant certificates.
Requirement 8.5.7 | The LED based Solar simulator must have metric fasteners and components.
Requirement 8.5.8 | The LED based Solar simulator must include a computer with accessories for system control compatible with W12.
Requirement 8.5.9 | The LED based Solar simulator must have an instruction manual in English.
8.6 | System
Requirement 8.6.1 | The LED based Solar simulator (including side equipment and free space for access of measurement and service compartments) must have following maximum dimensions (L×W×H) 300×350×299 cm3
· The system (or parts of) has to fit through a door of (W×H) 180×230 cm2.
Requirement 8.6.2 | The LED based Solar simulator must have safety features:
· Electrical safety according NEN 3140;
· Safety interlocks where needed;
· checks for critical utilities such as cooling water and ventilation;
· overtemperature protection of light engine.
Requirement 8.6.3 | Due to the long pulse length cooling of the light engines with water is required.
Preference 8.6.1 | Installation (max 100 point):
System to be installed in the space of the old solar simulator at the TNO location HTC21 Eindhoven. TNO will remove the old system, the new system is preferably installed in less than 5 days, including an on-site training.
Furthermore, the candidate should indicate if the solar simulator requires more utilities than present in the lab. Utilities present are:
· N2, CDA and cooling water (17-19°C 4-5 bar);
· Electrical connection max 3 phase 16A (5 pin CEE);
· Ventilation maximum 500m3/hour.
Assessment criteria | Downtime:
8.6.1 a | · 0 points, installation and training more than 10 working days;
· Or, 30 points, installation and training >= 5 working days < 10 working day;
· Or, 70 points, installation and training < 5 working days.
Assessment criteria | Infrastructure:
8.6.1 b | · 0 points if TNO must provide additional infrastructure for utilities;
· Or, 20 points if no additional infrastructure is required.
Assessment criteria | Cooling of light engines:
8.6.1 c | · 0 points if cooling with external chiller;
· Or, 10 points if cooling through process cooling water.
8.7 | Service and Support
requirement 8.7.1 | 2 year service and maintenance
Included 1 service visit per year for recalibration of light engine and critical components such as photospectrometer. Optional extension of the service and maintenance with periods of 2 year per extension in included.
Preference 8.7.1 | Service and support (max 150 point):
TNO intends to use the system for more than 10 years and expects the system to be supported by the manufacturer during this period.
Specify the service and support, during and after the warranty period of 24 months (as stated in 18.2 General Purchasing Conditions of TNO), addressing the following aspects:
· Repair of instruments and minimized downtime during malfunction;
· Delivery and availability of replacement parts;
· Service and support on-site and off-site to keep the system operational;
· Expected lifespan of the system;
· Response time of service department.
The complete answer for Preference 8.7.1 must not exceed a maximum of 2 pages A4. If the answer is longer than 2 pages, the excess pages will not be included in the assessment.
Assessment criteria | The assessment will be based on the overall impression given by the answer given. The preference does not involve separate assessment aspects or criteria. However, the continuity of operation over a longer timespan than 10 years will be the most important aspects for assessment.
· Maximum 150 points, in the manner as described in par 6.1.

## Preferences (hidden)
Annex A04, Prescribed format for answering questions/fullfilling preferences Guidelines for the European Public procurement tender for: one cell DLIT and one module DLIT, Ref. WS2616908858, dated 00 September 2025
The Tenderer demonstrates the ability to meet the preferences by means of his answers to the various questions/preferences, which must be structured in accordance with the instructions given in Chapter 8.
Preference | Description | Confirmation of fullfillment of preference Yes/No | Reference in Tenderer's documentation where Tenderer indicates in which way he can fulfil the preference | Remarks
8.2.4
8.4.6
8.4.7
8.5.2
8.5.3

## Blad2 (hidden)
Yes
No