## Lot 1
Response form - Program of Requirements and award criteria
Please note; the given terms will be included as a contractual agreement.
Awarding criteria - Lot 1: Nd:YAG laser | [6 empty] | Additional requirements to be submitted with Proposal | Tenderer's response
Type | # | Laser quality consisting of: | Answers | Answer tenderer | Score | % Total score |  | Any additional information which Tenderer wishes to provide. This is also your chance to explain why you checked "No". Be advised: even though a motivation has been given for a "No" score, it is up to the discretion of the KNMI to decided if the give explenation is valid.
KO | S1 | Wavelengths: The laser is a diode pumped Nd:YAG laser, able to emit at 1064 nm, 532 nm and 355 nm, sumulatneously | Y-N |  | Not applicable
KO | S2 | Optical alignment: The laser shall be configured to emit the laser beams at the three wavelengths parallel and spatially separated, fitting the layout of the Caeli emitter configuration (see Figure 1) | Y-N |  | Not applicable |  | A schematic drawing must be provided
KO | S3 | Physical size: The laser head shall fit the existing optical table and lidar emitter set-up with a footprint not larger than 150 cm x 80 cm, excluding cables and hoses. The total height of the laser head shall not exceed about 70 cm. If panels of covers need to be removed for access to parts, the available heigh is 100 cm. | Y-N |  | Not applicable |  | A schematic drawing must be provided
KO | S4 | Physical size: the power supply shall be a unit smaller than 90 nm high, 90 cm deep and 100 cm wide. | Y-N |  | Not applicable |  | A schematic drawing must be provided
W | W1 | Physical weight of the laser head is less or equal to 40 kg. Yes = 1 point No = 0 points | Y-N |  | 0
KO | S5 | Necessary cable lenths for the power supply unit shall be about 4.5 m. If cable lengths need to be shorter, the power supply must fit below the optical table in a 19-inch rack | Y-N |  | Not applicable |  | A schematic drawing must be provided
KO | S6 | Necessary hose lengths for the chiller unit shall be about 4.5 m. Quick connects shall enable disconnecting hoses. | Y-N |  | Not applicable
KO | S7 | A seed laser for stabilisation of the fundamental wavelength is included | Y-N |  | Not applicable |  | A test report showing the wavelength stability must be provided
KO | S8 | Emitted power at 355 nm: the emitted pulse energy at 355 nm shall be 330 mJ per pulse or better at a pulse repetition rate of 100 Hz | Y-N |  | Not applicable
KO | S9 | Emitted pulse energy at 532 and 1064 nm shall be at least 220 mJ and 220 mJ per pulse while delivering at least 330 mJ at 355 nm | Y-N |  | Not applicable
W | W2 | Emitted pulse energy at 532 and 1064 nm shall be optimised and reach better than 250mJ and 300 mJ per pulse while keeping the requirement for the emitted pulse energy at 355 nm. Yes = 8 points No = 0 points | Y-N |  | 0 |  | Optimisation options must be given
KO | S10 | Polarisation: the emitted laser radiation at the fundamental output shall be linearly polarised | Y-N |  | Not applicable
KO | S11 | Polarisation purity: shall be 98% or better at 355 nm | Y-N |  | Not applicable |  | Measurements must be provided
KO | S12 | The beam diameter shall be 10 mm or less | Y-N |  | Not applicable
KO | S13 | Beam divergence shall be less than 0.5 mrad full angle for all wavengths | Y-N |  | Not applicable |  | Measurements must be provided
KO | S14 | The spatial beam profile shall be free of hot spots that may damage any optical components from the primary exit of the laser until a distance of 10 meters. | Y-N |  | Not applicable |  | Demonstrate with beam prints over the mentioned distance
KO | S15 | A soft start/power ramp up system is in place | Y-N |  | Not applicable
KO | S16 | The laser pulse duration shall be in the order of 10 ns | Y-N |  | Not applicable
KO | S17 | The SHG and THG wavelength conversion shall be actively stabilised | Y-N |  | Not applicable |  | A test report showing the power stability must be provided over a period of 24 hours
KO | S18 | A built-in laser energy meters shall be installed for 355 nm. The output can be logged by computer. | Y-N |  | Not applicable
W | W3 | A built-in laser energy meters is installed for 532 nm. The output can be logged by computer. Yes = 1 point No = 0 points | Y-N |  | 0
W | W4 | A built-in laser energy meters is installed for 1064 nm. The output can be logged by computer. Yes = 1 point No = 0 points | Y-N |  | 0
KO | S19 | A spatial beam profiling camera shall be installed inside the laser head. The camera output can be monitored and visualised on a controlling computer. | Y-N |  | Not applicable
KO | S20 | A full datasheet with specifications, factory instrument tests is provided | Y-N |  | Not applicable
KO | S21 | Remote control: it shall be possible to remotely control the laser through a unix system, via command line driven interface | Y-N |  | Not applicable
W | W5 | Manual control: It shall be possible to manually control the laser through a control pad directly connected to the laser, not connected to a control compter. Yes = 1 point No = 0 points | Y-N |  | 0
KO | S22 | Remote control shall be via ethernet | Y-N |  | Not applicable
W | S23 | Laser health parameters can be logged (temperature sensors). Yes = 2 points No = 0 points | Y-N |  | 0
KO | S24 | Spare parts shall be available for at least 10 years after installation | Y-N |  | Not applicable
KO | S25 | Installation on site is included | Y-N |  | Not applicable
KO | S26 | Instructions on site is included for 2 - 4 persons | Y-N |  | Not applicable
KO | S27 | The laser power supply shall work with 240V/AC or 400V/AC | Y-N |  | Not applicable
KO | S28 | The pump diodes shall be warranted for at least 4 billion shots | Y-N |  | Not applicable
W | W6 | Warranty for the laser head shall be extended for a period of at least 3 years Yes = 4 points No = 0 points | Y-N |  | 0
KO | S29 | During a period of 12 months after installation any necessary adjustments to the laser alignment and beam profile shall be performed free of charge in case the laser beam shows hot spots that cause damage to external optics. | Y-N |  | Not applicable
W | W7 | The cooling system is directly water to air. No separate chiller unit is needed. Yes = 4 points No = 0 points | Y-N |  | 0
|  | Subtotal score: (if all knockout criteria are met) | 22 |  | 0 | 73.33%
Type | # | Other criteria: | Answers | Answer tenderer | Score
KO | S30 | Delivery time up to 6 months (180 days) | Y-N |  | Not applicable
W | W8 | Environmentally conscious production process. This must be demonstrable, for example through certification under ISO 14001 (or an equivalent system), and/or a process description that clearly shows whether sufficient attention is given to environmentally responsible manufacturing. Yes = 1 point No = 0 points | Y-N |  | 0
KO | S31 | Warranty Free of charge replacment of faulty components for (at least) 12 months a component is deemed faulty if it causes the station to deviate from the knockout criteria | Y-N |  | Not applicable
W | W9 | Faster delivery up to max. 120 days - 2 points up to max. 140 days - 1 points up to max. 180 days - 0 points | 1. up to max. 120 days 2. up to max. 140 days 3. up to max. 180 days |  | 0
W | W10 | Response time for malfunctions during campaigns within 3 business days during campaigns - 3 points within 1 week during campaigns - 2 points more than 1 weeks - 0 points | 1. within 3 business days during campaigns 2. within 1 week during campaigns 3. more than 1 weeks |  | 0 |  | Campaigns are intenstive observation periods of 1-2 months in duration (continuous). This happens typically every other year. Once per year in exceptional situations.
W | W11 | Response time for malfunctions within 2 weeks - 2 points within 1 month - 1 points more than 1 month - 0 points | 1. within 2 weeks 2. within 1 month 3. more than 1 month |  | 0
|  | Subtotal score:: (if all knockout criteria are met) | 8 |  |  | 26.67%
|  | Total score) | 30 |  |  | 100%

## Blad1
Yes
No
W10
up to max. 120 days
up to max. 140 days
up to max. 180 days
within 3 business days during campaigns
within 1 week during campaigns
more than 1 weeks
within 2 weeks
within 1 month
more than 1 month