## Annex A04
Annex A04 - Conformity list 'Programme of Requirements and Preferences' Guidelines for the European Public Tender for a Dual‑beam FIB/SEM system, ref. TNO WS2715368683, date: 07-01-2026
By the following Programme of Requirements TNO will assess if Tenderer meets the relevant requirements as laid down in Section 6 and 8. The Tenderer must indicate for each item, the conformity for the requirements in the table below.
If to any of the requirements, as stated in section 8, are not met by Tenderer, he will be excluded for any further assessment. By answering a question with “comply” will mean that in the event of awarding the contract, non-compliance to the requirement shall count as failure to meet an obligation from the Agreement to be concluded. If Tenderer does not comply with any of the preferences, as mentioned in point 8, he will not be excluded from further assessment. Answering a preference with "comply with" means that, at the time of award, failure to comply with the preference counts as failure to fulfil an obligation under the Agreement to be concluded.
Requirements
Requirement | Decription | Tenderer is compliant to the requirement | Reference in Tenderer's documentation where compliances with the requirement is described The Tenderer is required to include a reference indicating where in their documentation the detailed justification or elaboration is provided, demonstrating how the Tenderer meets the requirement. | Remark
Requirement 8.100.005 | The Dual‑beam FIB/SEM system shall be a newly build first hand system. | Click the drop-down box to select
Requirement 8.100.006 | TNO is willing to accept offers for refurbished or newly overhauled equipment, provided that strict conditions are met. The refurbishment or overhaul must have been carried out, or be scheduled to be carried out, exclusively by the Original Equipment Manufacturer (OEM). All warranty and maintenance obligations related to the offered refurbished or newly overhauled equipment must be fully guaranteed by the OEM for the duration specified in this tender. The Tenderer must ensure the availability of all necessary spare parts and consumables throughout the intended lifespan of the equipment. Prior to the final award of the contract, and in addition to paragraph 7.1, the Tenderer shall provide written evidence that the refurbished or newly overhauled equipment will be available on the contract award date as described in paragraph 2.1. The Tenderer must clearly indicate whether the proposed system is new, refurbished, or newly overhauled. To enable TNO to verify the actual condition of the offered equipment, the Tenderer must submit appropriate proof. Acceptable forms of evidence include an OEM certificate of refurbishment or overhaul, factory test reports, serial number validation (production or refurbishment), OEM documentation specifying the refurbishment date and scope of work, or a formal OEM statement confirming the equipment’s status. | Click the drop-down box to select
Requirement 8.100.010 | Two-Beam FIB/SEM system The Tenderer shall provide a fully integrated, turnkey two-beam station combining a Focused Ion Beam (FIB) and a Focused Electron Beam (SEM), specifically designed for submicron prototyping and surface functionalization using the FIB. The system must include all components necessary for full operational capability, requiring only the following infrastructure provisions by TNO: - Electrical power - Electrical grounding - Clean Dry Air (CDA) - Clean nitrogen - Vacuum exhaust Remark: An installed phone line in the laboratory and a LAN connection will be provided by TNO. | Click the drop-down box to select
Requirement 8.200.005 | Milling Resolution The system shall support ion beam milling with a minimum feature size of ≤ 20 nm when using Si++, Au++, and Ge++ ion species at a beam energy of 35 keV. Rationale: This resolution is required for the fabrication of quantum technologies. | Click the drop-down box to select
Requirement 8.200.010 | Beam size The system shall support ion beam size resolution of ≤ 12 nm when using Si++, Au++, and Ge++ ion species at a beam energy of 35 keV, determined by the sharpness of intensity scans. Rationale: This resolution is required for the fabrication of quantum technologies. | Click the drop-down box to select
Requirement 8.200.015 | Field size The lithography system shall support a field size of up to 200µm, with automatic correction mechanisms for distortion and stigmation. Rationale: A field size of 200µm shall provide adequate coverage for FIB tasks, while automatic corrections shall ensure consistent alignment during processes. | Click the drop-down box to select
Requirement 8.200.020 | Beam position stability The system shall demonstrate beam position stability of ≤ 400 nm over a continuous period of 8 hours, measured with a stationary (resting) stage using a Si++ ion beam at 35 keV. Rationale: Beam position stability is critical for long-duration patterning and milling operations. | Click the drop-down box to select
Requirement 8.200.025 | Beam current stability The system shall maintain a beam current stability of ≤ 1% per hour, measured using a Si++ ion beam at 35 keV. Rationale: Stable beam current is essential for ensuring consistent material removal rates and reproducible patterning results. | Click the drop-down box to select
Requirement 8.200.030 | Stitching The system shall provide automatic management of larger area work by means of stitching. The field stitching precision, specifically for Si++ at the highest beam energy, in a 100 µm field, shall be less than or equal to 50 nm (mean plus three sigma). This value is determined based on blind stitching, without the use of additional alignment marks. Rationale: Baseline High-precision stitching without the use of additional alignment marks is essential to ensure the functional integrity and performance of devices produced over larger areas. | Click the drop-down box to select
Requirement 8.200.035 | Overlay The system shall provide management of level-by-level alignment strategies (overlay). The software must be capable of automatically recognizing alignment marks and automatically calculating the correction parameters to be applied. Overlay accuracy, specifically for Si++ at the highest beam energy, in a 100 µm field, shall be less than or equal to 50 nm (mean plus three sigma). Rationale: Accurate overlay is essential to ensure that features on successive layers are correctly aligned, which is critical for the functionality and yield of advanced devices. | Click the drop-down box to select
Requirement 8.200.040 | Stitch-free continuous tracks The system shall support stitch-free patterning of lines over large distances (1 mm) by enabling uninterrupted structure fabrication with a quasi-stationary beam during stage movement. Confirmation is reached via straightness and width uniformity tests over a 1 mm long line fabricated on a metal coated silicon sample. For a line with less or equal to 120 nm width, the straightness shall be below or equal to 20 nm and the uniformity shall be below 1or equal to 10 nm. Rationale: This capability is essential to ensure the structural and functional integrity of long distance continuous nanostructures, particularly in applications where stitching defects would compromise device performance, uniformity, or reliability. | Click the drop-down box to select
Requirement 8.200.045 | Stitch-free periodic structures The system shall support stitch-free patterning of periodic structures over large distances (1 mm) by enabling continuous, uninterrupted structure fabrication with a deflected beam during stage movement. Confirmation is reached via patterning of two grating stripes of 2mm length with a measured average pitch of 400nm +/-0.2nm and 400.5nm+/-0.2nm. Fabrication is demonstrated by ion beam exposure of a resist layer on a silicon sample. Rationale: This capability is essential to ensure the structural and functional integrity of large scale periodic nanostructures, particularly in applications where stitching defects would compromise device performance, uniformity, or reliability. | Click the drop-down box to select
Requirement 8.310.005 | Ion column the FIB column shall be positioned orthogonally to the interferometric stage. Rationale: For maximum stability of ion lithography processes, it is vital the main FIB column is mounted orthogonally to the stage. | Click the drop-down box to select
Requirement 8.310.010 | Liquid metal alloy Ion Source The FIB column shall be equipped with a Liquid Metal Alloy Ion Source (LMAIS) compromising of at least gold (Au), silicon (Si) and germanium (Ge). Rationale: This configuration ensures flexible and precise ion beam processing. Heavy atoms (Au) serve a high sputter yield. Elements Si and Ge allow for compatibility with sensitive materials for quantum technologies. | Click the drop-down box to select
Requirement 8.310.015 | Software-Controlled Ion Type Switching The system shall support switching between specific ion types by software means only, with a switching time of less than one minute, provided the operating voltage remains unchanged. Rationale: Fast and software-controlled ion switching is essential for efficient multi-species processing workflows. | Click the drop-down box to select
Requirement 8.310.020 | Beam Voltage Range The system shall support a continuously adjustable beam voltage range from 10 kV to 35 kV. Rationale: This voltage range is essential to accommodate a wide variety of ion beam applications. | Click the drop-down box to select
Requirement 8.310.025 | Beam Blanker The system shall be equipped with an electrostatic beam blanker with a rise time of at least 50 nsec. Rationale: A fast beam blanker is essential for overall process duration. | Click the drop-down box to select
Requirement 8.310.030 | Beam Current Range with Motorized Aperture Selection The system shall support a selectable beam current range between approximately 5 pA and 1 nA, depending on the ion species used. Selection of beam current shall be accomplished via motorized aperture control. Rationale: A wide and precisely selectable beam current range is essential for accommodating diverse nanofabrication tasks. | Click the drop-down box to select
Requirement 8.320.005 | Electron Source Type The system’s electron column shall be equipped with a thermal field emitter of the Schottky ZrO/W type. Rationale: A Schottky ZrO/W thermal field emitter provides high brightness, excellent beam stability, low energy spread and long lifetime. | Click the drop-down box to select
Requirement 8.320.010 | Beam Control and Scan Field Range The system’s SEM column shall include all necessary components for beam focusing and deflection, and shall support a selectable scan field size ranging from 1 µm to approximately 400 µm, depending on beam energy and working distance. Rationale: Precise control over beam focus and deflection is essential for high-resolution imaging and accurate nanostructure inspection. A flexible scan field range enables both detailed analysis of small features and overview imaging of larger areas. | Click the drop-down box to select
Requirement 8.320.015 | Imaging Resolution The system’s SEM column shall achieve an imaging resolution better than 5 nm at an accelerating voltage of 20 kV. Rationale: High-resolution imaging is essential for accurate inspection, metrology, and nanostructure verification in advanced research and development. | Click the drop-down box to select
Requirement 8.320.020 | Beam Voltage Range The system’s SEM column shall support a continuously adjustable beam voltage range from 5 kV to 20 kV. Rationale: This voltage range is essential to accommodate a variety of imaging conditions, from low-voltage surface-sensitive inspection to high-voltage deep-structure visualization. | Click the drop-down box to select
Requirement 8.320.025 | Beam Blanker The system shall be equipped with an electrostatic beam blanker with a rise time of at least 50 nsec. Rationale: A fast beam blanker is essential for overall process duration. | Click the drop-down box to select
Requirement 8.320.030 | Beam Current Range The system’s SEM column shall support a typical beam current range between 10 pA and greater than 600 pA, depending on operating conditions and configuration. Rationale: This current range is essential to accommodate both high-resolution imaging at low currents and faster acquisition or broader-area inspection at higher currents. | Click the drop-down box to select
Requirement 8.330.005 | Sample Stage Configuration The system shall be equipped with a sample positioning stage offering three-axis movement (X, Y, Z). Rationale: Three-axis positioning is essential for precise sample alignment, focus control, and navigation during imaging and nanofabrication processes. | Click the drop-down box to select
Requirement 8.330.010 | Stabilized Laser Source The system shall implement position control in the XY plane using a double interferometer configuration, based on a stabilized laser source. Rationale: Double interferometric control ensures nanometer-level precision and repeatability in stage positioning. | Click the drop-down box to select
Requirement 8.330.015 | XY Stage Travel Range The system shall be equipped with an interferometric stage for the XY plane (sample surface plane), providing a minimum travel range of 100 mm × 100 mm or larger, to ensure compatibility with 4-inch sized samples. Rationale: A large and precise XY travel range is essential for handling standard wafer formats and enabling high-resolution patterning or inspection across extended sample areas. | Click the drop-down box to select
Requirement 8.330.020 | True Mechanical Positioning Resolution The system’s interferometric XY stage shall achieve a true mechanical positioning resolution of 1 nanometer. Rationale: Nanometer-level mechanical positioning is essential for high-precision patterning, stitching, and alignment tasks in advanced nanofabrication workflows. | Click the drop-down box to select
Requirement 8.330.025 | XY Stage Speed The system shall support a stage speed of ≥ 2 mm/s in the XY plane. Rationale: A stage speed of ≥ 2 mm/s ensures efficient processing. | Click the drop-down box to select
Requirement 8.330.030 | Z-Axis Travel Range and Resolution The system’s sample stage shall provide a minimum travel range of 10 mm along the Z-axis, with a resolution of 1 µm or better. Rationale: A vertical resolution of ≤ 1 µm is essential for accurate focusing, sample alignment, and multi-layer processing. A travel range of at least 10 mm offers flexibility in height positioning. | Click the drop-down box to select
Requirement 8.410.005 | Working Chamber Integration with Load Lock The system’s working chamber shall house the sample stage and shall be directly connected to a load lock for sample transfer. Rationale: Integrating the sample stage within the working chamber ensures optimal environmental control. Connection to a load lock enables fast sample exchange without venting the main chamber, which is essential for maintaining stable vacuum conditions. | Click the drop-down box to select
Requirement 8.410.010 | Load-Lock Sample Entry System The system shall include a load-lock mechanism for sample introduction, capable of handling samples up to 100 mm × 100 mm in size and transferring them to the stage without breaking vacuum. Rationale: A load-lock system enables efficient sample exchange, which is essential for maintaining vacuum integrity while. Compatibility with samples up to 100 mm × 100 mm ensures support for standard wafer formats and large-area substrates. | Click the drop-down box to select
Requirement 8.410.015 | Optical Height Sensing System with Automatic Focus Adjustment The system shall include an optical height sensing system, capable of automatically measuring sample height after each stage movement to maintain focus. The system shall operate over a Z-range of 2 mm or greater, with a reproducibility of 1 µm or better. Rationale: Accurate and automated height sensing is essential for maintaining consistent focus when working across samples with varying topography. A Z-range of ≥ 2 mm ensures compatibility with typical sample height variations, while reproducibility of ≤ 1 µm guarantees reliable focus correction. | Click the drop-down box to select
Requirement 8.410.020 | Chamber-Mounted Secondary Electron Detector The system shall include a chamber-mounted secondary electron detector, featuring automatic contrast and brightness control with manual override. Rationale: A secondary electron detector is essential for high detailed surface imaging with excellent signal-to-noise ratio. | Click the drop-down box to select
Requirement 8.410.025 | Built-In Infrared Camera System for Interior Observation The system shall include a built-in infrared camera system with integrated optics, enabling visual control and orientation within the working chamber. Rationale: Real-time visual observation is essential for safe and precise sample handling, alignment, and operational monitoring. | Click the drop-down box to select
Requirement 8.410.030 | Fully Automatic Dry Vacuum System with Column Isolation and UHV Capability The system shall be equipped with a fully automatic vacuum system, incorporating column isolation valves and vacuum gauges for monitoring and control. The vacuum system shall be of the dry type, comprising a pre-pump, a turbomolecular pump, and ion getter pumps, enabling ultra-high vacuum (UHV) generation in the gun region. Rationale: A robust and automated vacuum system is essential for maintaining stable operating conditions and ensuring contamination-free environments. Column isolation valves and vacuum gauges provide operational safety and diagnostic capability. The use of dry pumping technology eliminates oil-related contamination risks, while the combination of turbomolecular and ion getter pumps ensures UHV levels required for optimal emitter performance and beam stability. | Click the drop-down box to select
Requirement 8.410.035 | Integrated Air Damping Vibration Isolation System The system shall incorporate an integrated air damping vibration isolation system within the plinth structure. Rationale: An integrated air damping system within the plinth provides a stable mechanical foundation, reducing the influence of building vibrations and acoustic noise. | Click the drop-down box to select
Requirement 8.410.040 | Integrated Plasma Cleaner The system shall incorporate an integrated plasma cleaner. Rationale: An integrated plasma cleaner is required for cleaning basic contaminants inside the main chamber. | Click the drop-down box to select
Requirement 8.420.005 | PC-Based Multi-User Control Software with Integrated Patterning and Metrology Capabilities The system shall include a PC-based control software platform that supports multi-user access and provides comprehensive functionality for pattern design, processing, imaging, and metrology. The software shall feature a fully integrated GDSII CAD editor, supporting beam control and parameter management through recipe-based configurations, allowing reproducible and automated operation across different user profiles. Rationale: Multi-user capability ensures accessibility and workflow continuity across research teams. The integration of a GDSII CAD editor streamlines design-to-pattern processes, while recipe-based beam control and parameter handling enhance repeatability and reduce setup time. | Click the drop-down box to select
Requirement 8.420.010 | Fully Digital Column Control with Parameter Set Save and Reload The system shall feature a fully digital column control interface, enabling precise management of all column-related parameters. The control system must support the saving and reloading of complete parameter sets. The interface shall be integrated into the main system software and accessible via the multi-user platform. Rationale: Digital column control enhances reproducibility, precision, and ease of use in complex nanofabrication and imaging workflows. The ability to save and reload parameter sets ensures consistent performance across sessions and users, reduces setup time, and minimizes human error. | Click the drop-down box to select
Requirement 8.420.015 | Automation Interface with Python Scripting for Unattended Batch Job Generation The system shall provide an automation interface that supports Python scripting for the creation and execution of unattended batch jobs. This interface must allow users to define, schedule, and run complex workflows without manual intervention, including patterning, imaging, and metrology tasks. The scripting environment shall be fully integrated with the system control software and offer access to all relevant operational parameters and functions. Rationale: Automation through Python scripting enables flexible and scalable control over nanofabrication and imaging processes. It supports reproducibility, reduces operator workload, and allows for advanced customization of experimental protocols. | Click the drop-down box to select
Requirement 8.420.020 | Software Licensing and Offline Data Preparation The system shall include perpetual licenses for the operating system of the PC system and for all software components described in this specification. In addition, the supplier shall provide a total of six offline licenses for data preparation, enabling installation and use on Buyer-owned Microsoft Windows PCs. These offline licenses must support full functionality for pattern design and preparation workflows, independent of the main system. Rationale: Perpetual licensing ensures long-term operational continuity and avoids dependency on subscription renewals or license expirations. Providing offline licenses supports distributed data preparation workflows, allowing multiple users to work in parallel without occupying the main system. | Click the drop-down box to select
Requirement 8.420.025 | Software Updates: Bug Fixes and Minor Improvements The Supplier shall provide software updates free of charge for the entire minimum expected lifetime of the system. These updates shall include, but are not limited to, bug fixes, security patches, and minor functional improvements. If the implementation of a software update requires a hardware modification, the Supplier shall present this option to TNO for prior approval. Should TNO choose to proceed with the hardware modification, no additional license fees shall apply for the software update. The cost of the required hardware shall be itemized separately in accordance with the agreed pricing schedule. Rationale: Continuous access to software updates is essential for maintaining system reliability, security, and performance. Free updates ensure that the system remains aligned with evolving operational requirements and technological standards. The option for hardware modification - subject to TNO’s approval - provides flexibility for future enhancements without introducing hidden costs. This approach supports TNO’s long-term research objectives in photonics, quantum devices, and semiconductor prototyping, where system integrity and adaptability are critical. | Click the drop-down box to select
Requirement 8.420.030 | Software Upgrades: Functional Enhancements and Major Releases The Supplier shall provide software upgrades free of charge to TNO for a minimum period of x years following delivery of the system, as part of regular product development. Software upgrades refer to substantial functional enhancements, new software versions, or performance improvements. Upgrades may include: - Introduction of new features or modules - Performance optimization and scalability improvements - Migration to newer platforms or operating systems (e.g., Windows upgrades) - Re-engineering of software architecture for maintainability and future-proofing If a hardware upgrade is required to implement a software upgrade, the Supplier must notify TNO in advance. TNO reserves the right to approve or reject the proposed hardware upgrade. If approved, no additional license fees shall be charged for the software upgrade. Hardware costs must be specified separately in accordance with the agreed pricing schedule. The Supplier shall also provide technical support for the implementation and operation of software upgrades throughout the minimum expected lifetime of the system. This support shall include assistance with installation, configuration, troubleshooting, and user guidance, and shall be delivered within agreed service levels at no additional cost to TNO. Rationale: Access to functional upgrades ensures that the system remains aligned with evolving technological standards and user needs. Free upgrades promote long-term value and adaptability, while technical support guarantees smooth implementation and operational continuity. These provisions are essential for TNO’s research in photonics, quantum devices, and semiconductor prototyping, where system evolution and reliability are critical to maintaining scientific and engineering excellence. | Click the drop-down box to select
Requirement 8.430.005 | Load-Lock Integration The system shall support sample mounting via dedicated sample holders. These sample holders must be compatible with the system’s stage and introduced through a load-lock mechanism to ensure vacuum integrity and minimize contamination risk. The load-lock shall allow for efficient sample exchange without venting the main chamber, and the holders must provide secure and repeatable positioning for various sample types. Rationale: Controlled and contamination-free sample handling is essential for maintaining ultra-high vacuum conditions and ensuring reproducible results in high-resolution imaging and nanofabrication. The use of sample holders introduced via a load-lock enables fast and safe sample exchange, it also supports operational efficiency and protects sensitive components from environmental exposure. | Click the drop-down box to select
Requirement 8.430.010 | 3D Sample Handling Module with Tilt and Continuous Rotation The system shall include a dedicated 3D module for sample handling, enabling tilt and continuous rotation of small samples. This module must be fully compatible with the system’s air-lock mechanism. Rationale: Tilt and rotation capabilities are essential for advanced imaging, metrology, and nanofabrication tasks that require multi-angle access to sample surfaces. | Click the drop-down box to select
Requirement 8.430.015 | Universal Piece Part Holder for Small Samples The system shall include a universal piece part holder designed to accommodate small samples with dimensions up to 50 mm or 2 inches. The holder must ensure secure and stable mounting of various sample geometries and materials, and be fully compatible with the system’s stage and load-lock mechanism. Rationale: A universal holder for small samples enhances operational flexibility and efficiency by allowing quick exchange and reliable mounting of diverse sample types without the need for custom fixtures. Compatibility with the load-lock ensures vacuum integrity and supports streamlined workflows. | Click the drop-down box to select
Requirement 8.430.020 | 4" Wafer Holder According to SEMI Standard The system shall include a dedicated 4" wafer holder designed in compliance with SEMI standards and be fully compatible with the system’s stage and load-lock mechanism. Rationale: Dedicated holder that is required for integration into standardized 4 inch wafer fabrication platforms. | Click the drop-down box to select
Requirement 8.440.005 | Uninterrupted Power Supply (UPS) System The system shall include an uninterrupted power supply (UPS) system with an autonomy time of at least 30 minutes during the first year of operation. The UPS shall be integrated with the system control software to enable a controlled and safe shutdown procedure when the critical threshold for main power loss is reached. Rationale: A reliable UPS system is essential for protecting sensitive equipment and data during unexpected power interruptions. The 30-minute autonomy ensures sufficient time for safe shutdown or power restoration, minimizing operational disruption. Integration with the system’s control logic allows for automated response. | Click the drop-down box to select
Requirement 8.450.005 | Water-Air Chillers The system shall include water-air chillers as necessary to support stable thermal management of all components requiring active cooling. Rationale: Including water-air chillers ensures that temperature-sensitive components operate within their specified limits, preventing drift, degradation, or failure. | Click the drop-down box to select
Requirement 8.460.005 | Operator Table with Integrated Control Interfaces The system shall include an operator table as part of the delivery. The table shall be equipped with flat panel displays, keyboard, joystick, control panel, and mouse, all integrated to support full system operation. Rationale: A dedicated operator table with integrated control interfaces is essential for efficient and comfortable system operation. | Click the drop-down box to select
Requirement 8.500.005 | System Documentation The delivery shall include comprehensive operation and reference manuals in English, provided in digital form. These manuals must cover all aspects of system setup, operation, maintenance, troubleshooting, and safety procedures. Documentation shall be structured to support both routine use and advanced configuration. Rationale: Clear and complete documentation is essential for safe and efficient system operation, user training, and long-term maintainability. Providing manuals in English ensures accessibility across international teams and supports standardized procedures. | Click the drop-down box to select
Requirement 8.600.005 | Operator Training The delivery shall include five (5) days of basic operator training conducted on site. This training must cover system operation, safety procedures, routine maintenance, and troubleshooting. The training shall be delivered by qualified personnel and tailored to the specific configuration and use cases of the delivered system. In addition, the Supplier shall offer optional advanced or follow-up training modules upon request. Rationale: On-site operator training ensures that users gain hands-on experience with the system in its actual operational environment. It supports safe and efficient use, minimizes startup time, and reduces the risk of operational errors. Optional training modules provide flexibility for future skill development and onboarding of new personnel. | Click the drop-down box to select
Requirement 8.700.005 | Critical design review (CDR) The objective of the CDR is to confirm that the equipment and all deliverables satisfies the requirements. It establishes the compatibility of external & internal interfaces and the compliance of the design with contractual requirements (predicted and/or measured). The successful completion of the CDR permits the Supplier to start the manufacturing of the system. The CDR will contain the following details: - Physical dimensions of the system and system components. - External and internal interface location and dimensions of the system. - Showing the design meets the functional and performance requirements. The acceptance of the CDR shall by achieved using a written document, to be signed by both parties. In the event that the CDR fails to meet the requirements, the supplier or TNO must set out such failure in a written statement, to be signed by both parties. The supplier must then provide TNO with corrective action plan within fourteen (14) calendar days from the date of signing. The supplier must remedy these deficiencies at no additional cost within the timeframe set out in the corrective action plan, but not exceeding 1 month. | Click the drop-down box to select
Requirement 8.700.010 | Site Acceptance Test (SAT) The SAT contains the following checkpoints: - The FIB-SEM will be checked for damages; - The proper functioning of the equipment will be checked; o Mechanical functions, o Electrical safety- and switching functions, o Gas safety and functions, o Vacuum functionality, o User interface(s), including software. - All requirements and preferences, if offered, will be checked; - The presence and quality of the documentation will be checked. A full user acceptance test on-site of the contracted tenderer must be performed by the Tenderer in cooperation with TNO. Following approved user acceptance test on-site, the FIB-SEM shall be at the unrestricted disposal of TNO for one month. If the FIB-SEM works according to TNO’s functional and technical specifications, without faults and/or malfunctions during operational mode, the test period is considered successfully concluded by TNO. If the FIB-SEM is not operating in this period according TNO’s functional and technical specifications, the faults and/or malfunctions will be remedied and another test period of one month must successfully be concluded and approved by TNO, up to a maximum of three test periods of one month each. The details of the site acceptance test shall be agreed between the user and the supplier. Tenderer must submit a draft test protocol for the site acceptance test with its Tender. In the event that the deliverables fail to meet the user Acceptance Criteria for the SAT, the Supplier must set out such failures in a written statement, to be signed by both parties. The supplier must then provide TNO with a corrective action plan for such failures within fourteen (14) calendar days from the date of signing of such statement. The supplier must remedy these deficiencies at no additional cost within the timeframe set out in the corrective action plan, but not exceeding 3 months. A new acceptance test will then take place. | Click the drop-down box to select
Requirement 8.700.015 | Terms and conditions of payment - 30% at order, after receipt of a bank guarantee covering 30% of the total amount. - 50% after delivery at location TNO Delft, STW. At this moment, after a positive visual inspection of the delivered Item, TNO will return the bank guarantee. - 20% after installation and commissioning On-site and Site Acceptance Test (SAT), including test periods, approved by TNO, Management Quantum Technology and Procurement. The Tenderer to whom the Contract will be awarded must provide a bank guarantee to TNO, covering at least 40% of the total price. The Bank guarantee and the bank who issues the Bank guarantee have to be approved by TNO (minimal A rating). TNO must have the right to draw the bank guarantee in any and all cases where Tenderer falls short in fulfilling its obligations under the Contract. The validity of the bank guarantee may not expire until approval of the SAT by TNO Management and Procurement. Alternatively to a bank guarantee, TNO also accepts a Parent Company Guarantee, in which the parent company agrees to the repayment of the progress payment in the event that tenderer, or wholly owned subsidiary, defaults on the supply of an FIB-SEM, which is the subject of this order. | Click the drop-down box to select
Requirement 8.700.020 | MILL Test Certificate When supplying the system, the Tenderer shall provide a Mill Test Certificate (MTC), or similar certificate such that it meets the requirements and obligations to establish the country of origin (further specified in paragraph 5.2.4). | Click the drop-down box to select
Requirement 8.700.025 | Maintenance/services Contract Further specified in paragraph 1.5. | Click the drop-down box to select
Requirement 8.700.030 | Comprehensive Maintenance Facilities within Europe To ensure uninterrupted operation and long-term reliability of the equipment, it is essential that the Tenderer maintains comprehensive maintenance facilities within Europe. These facilities must be capable of handling diagnostics, repairs, and parts replacement to minimize downtime and ensure rapid resolution of technical issues. Strategically located European facilities enable: - Efficient service delivery through proximity to operational sites. - Reduced logistical delays and transportation costs. - Faster turnaround times for maintenance and support. Furthermore, TNO strongly prefers to avoid export restrictions, elevated costs, and administrative burdens associated with cross-border servicing outside Europe. European-based facilities help mitigate risks related to: - Export control regulations that may limit the movement of equipment or parts. - Customs delays and additional fees. - Dependency on non-EU jurisdictions, which may introduce uncertainty over the lifetime of the equipment. This requirement aligns with TNO’s strategic goals of operational continuity, cost-effectiveness, and regulatory compliance throughout the equipment’s lifecycle. | Click the drop-down box to select
Requirement 8.700.035 | The Contract will be executed through a purchase order, incorporating the Purchasing Conditions TNO for Goods, December 2025, and Incoterms 2020 Delivered Duty Paid (DDP), see paragraph 2.2.18. | Click the drop-down box to select

## validatie (hidden)
Click the drop-down box to select
Yes, compliant to requirement
No, not compliant to requirement
Click the drop-down box to select
Yes, Tenderer can comply with preference
No, Tenderer cannot comply with preference
note: dit blad 'verbergen' indien je dit gaat delen intern/extern (behoudens aan inkoop collega's)