## 1.Requirements Program for Auto
Programma van Eisen 1a. Automation & Integration
Failure to meet one or more (exclusion) knock-out (KO) criteria in this Requirements Document will result in your submission no longer being considered for the award of this contract. Meeting all knock-out criteria is regarded as a mandatory condition in this European Tender. Compliance with the Requirements Document does not earn any points.
The (question) statements in this Requirements Document are primarily intended to determine: 'Can the supplier meet the specified description?' Failure to meet one (or more) 'knock-out criteria' will result in rejection. Incorrect or incomplete filling in will lead to the immediate invalidation of the submission. The Requirements Document must be completed truthfully. PLEASE NOTE: The Excel sheet is not locked; the bidder should only fill in the fields that are highlighted in yellow and where it is indicated that something should be filled in.
|  |  | KO | Fill in with Yes or No; failure to fill in is not allowed. | Any explanation | Proof, documentation to be provided with the submission
Legal requirements for coagulation equipment and accessories
2 |  | The analysis equipment / manufacturer / supplier and consumables must demonstrably comply with the European directive (EU) 2017/746 when it comes into force. | KO | To be filled in by the bidder |  | To be provided with the submission
3 |  | The platform should have a biosafety level II enclosure providing full sample and user protection and fully tested to (BS EN 12469 2000 – Performance criteria for microbiological safety cabinets). | KO | To be filled in by the bidder |  | To be provided with the submission
General conditions of the platform.
4 |  | This platform is aimed at automation of high throughput expansion of newly derived pluripotent stem cells (PSCs). | KO | To be filled in by the bidder | Please add configura
5 |  | The platform should be able to run autonomously for at least 60 hours during expansion of the PSC lines, and can interact remotely with lab personel. | KO | To be filled in by the bidder
6 |  | The ability to automatically culture including refreshing medium, passaging and potentially freezing of PSCs. | KO | To be filled in by the bidder
7 |  | The platform should be able to refresh and expand a minimum of 84 PSC lines simultaneously within 24 hours (see Appendix workflow - 9 for the number of plates). | KO | To be filled in by the bidder | The vendor must submit documented proof, such as certifications, test reports, or case studies, to demonstrate compliance with the required specifications.
8 |  | The platform should be able to perform clump passaging of iPS cells. | KO | To be filled in by the bidder
9 |  | The platform should be able to perform single cell passaging. | KO | To be filled in by the bidder
10 |  | True Class II type biological safety enclosure providing full sample and user protection and is fully tested according to "BS EN 12469 2000-Performence criteria for microbiological safety cabinets". | KO | To be filled in by the bidder
11 |  | The platform should be a recirculating system that does not require external connection to the site HAVC. | KO | To be filled in by the bidder
12 |  | The airfiltration system should maintain a100% coverage of the entire work minimising risk of contamination entering the system. | KO | To be filled in by the bidder
13 |  | The platform should have internal environment monitoring with full real time data logging of temperature & humidity. | KO | To be filled in by the bidder
14 |  | The platform must have an illuminated visual alert system that indicates the system status and the stage of the process | KO | To be filled in by the bidder
15 |  | The platform should have anti-microbial worksurfaces to reduce potential sources of contamination. | KO | To be filled in by the bidder
16 |  | One portal should be opened for operator access without impacting on the internal environment. | KO | To be filled in by the bidder
17 |  | When the portal is opened the air flow over the open door should be increased to maintain the internal environment and prevent internal air from exiting the enclosure. | KO | To be filled in by the bidder
18 |  | The platform should have a large portal for a liquid handler to enable operator access and standalone use without impacting on the internal environment enclosure | KO | To be filled in by the bidder
19 |  | The platform enclosure should be configured for peroxide decontamination. | KO | To be filled in by the bidder
20 |  | The instrument slides for the imager and bulk dispenser must allow retraction from the workflow as needed to enable operator access, without compromising the internal environment. | KO | To be filled in by the bidder
21 |  | Areas where liquid is dispensed should be bunded to contain any spillages and include an operator alert in the event of a spill. | KO | To be filled in by the bidder
22 |  | The platform should include under-bench bins for waste tips and labware, equipped with an operator alert to indicate when they are full. | KO | To be filled in by the bidder
23 |  | The platform should have integrated power and service connections, including a main power panel for single-phase connection, a central mains switch, and accessible fuses. | KO | To be filled in by the bidder
24 |  | The platform should include a robot arm. | KO | To be filled in by the bidder | Please add configura
25 |  | The plaform should have one or more incubator(s) with the capacity of at least 300 plates. | KO | To be filled in by the bidder
26 |  | The platform must include drivers and control interfaces for selected instruments | KO | To be filled in by the bidder
27 |  | The platform should include an automated imager. | KO | To be filled in by the bidder
28 |  | The platform must have a Liquid Handler | KO | To be filled in by the bidder
29 |  | The enclosure should provide sufficient space to accommodate future upgrades and the addition of extra equipment. | KO | To be filled in by the bidder
30 |  | The instrument slides for the imager and bulk dispenser must allow retraction from the workflow as needed to enable operator access, without compromising the internal environment. | KO | To be filled in by the bidder
31 |  | The platform should include a media fridge & delivery system | KO | To be filled in by the bidder
32 |  | The plateform should be able to report Sendai virus clearence of the clones | KO | To be filled in by the bidder
33 |  | The platform should have an automated centrifuge compatebel with plate and tubes | KO | To be filled in by the bidder
34 |  | The platform is equipped with a PC and monitor to provide an interface for system control, data processing, and user interaction. | KO | To be filled in by the bidder
35 |  | All instruments are novel and or not used. | KO | To be filled in by the bidder | Please add configura
36 |  | All necessary equipment to use the device is included. | KO | To be filled in by the bidder
37 |  | Vendor must include installation and support of the system at Erasmus MC | KO | To be filled in by the bidder
38 |  | Vendor should provide a full design review including hardware and software specifications | KO | To be filled in by the bidder
39 |  | Vendor will provide system training to all employees of the iPS core team | KO | To be filled in by the bidder
40 |  | Vendor will provide user manuals for the complete system | KO | To be filled in by the bidder
41 |  | The vendor must perform a Factory Acceptance Test (FAT) and Site Acceptance Test (SAT) to validate that the system meets all performance specifications and safety requirements. | KO | To be filled in by the bidder
Automation and Integration of PSCs expansion
42 |  | A laboratory automation specialist is invited to tender for the system design to automate the following three workflows; high troughput iPS Expansion, Sendai clearence and cell line QC testing workflow driven by a central control software, every 4.5 to 5 weeks.(see Appendix workflow - 9) | KO | To be filled in by the bidder | Please add configura
43 |  | It should be possible manually place a minimal of 42 newly generated PSC lines, each line on a seprate 12-well plate and consist of 8 clones in 8 different wells in the platform. These 42 lines are called a batch. (see Appendix workflow - 9) | KO | To be filled in by the bidder
44 |  | The platform has to be able to report Sendai virus clearence, expand and perform QC testing of each batch (42 PSC lines) according to the iPS core facility workflow every 4.5 to 5 weeks. This 4.5 to 5 weeks workflow is called a cycle. (see Appendix workflow - 9) | KO | To be filled in by the bidder | The vendor must submit documented proof, such as certifications, test reports, or case studies, to demonstrate compliance with the required specifications.
45 |  | The platform should be able to maintain iPSCs, including daily imaging and reporting on phenotype quality by measuring the ratio of undifferentiated cells to spontaneously differentiated cells for the iPSC clones, confluency, and should be able to perform medium exchange, and passaging of each batch (42 newly generated iPSC lines) every 4 to 5 days. (see Appendix workflow - 9) | KO | To be filled in by the bidder | The vendor must submit documented proof, such as certifications, test reports, or case studies, to demonstrate compliance with the required specifications.
46 |  | The platform should have the capacity to work with an overlap of 3 weeks between batch 1 & 2 or batch 2 & 3 of PSC lines in parallel in a stable and efficient manner. (see Appendix workflow - 9) | KO | To be filled in by the bidder | The vendor must submit documented proof, such as certifications, test reports, or case studies, to demonstrate compliance with the required specifications.
47 |  | The system should be able to continuously expand iPS lines in a structured manner, ensuring a steady workflow during the expansion of 3 iPS batches in three cycles. (see Appendix workflow - 9) | KO | To be filled in by the bidder | The vendor must submit documented proof, such as certifications, test reports, or case studies, to demonstrate compliance with the required specifications.
48 |  | The platform should able to test Sendai clearence from each of the 42 lines at passage 7 by duplicating each line (8 clones per plate), harvesting cells for RNA isolation, cDNA prepration, and perform RT-PCR followed by feedback to the central software platform to indicate which clones need to expand. (see Appendix workflow - 9) | KO | To be filled in by the bidder | The vendor must submit documented proof, such as certifications, test reports, or case studies, to demonstrate compliance with the required specifications.
49 |  | The platform should be able to expand 4 selected clones from each of the 42 lines per batch, based on negative Sendai status and an undifferentiated phenotype.(see Appendix workflow - 9) | KO | To be filled in by the bidder | The vendor must submit documented proof, such as certifications, test reports, or case studies, to demonstrate compliance with the required specifications.
50 |  | The platform should be able to expand each selected clone from 1 well of a 12-well plate, and to all 12 wells of a 12-well plate per each line within 2 passages.(see Appendix workflow - 9) | KO | To be filled in by the bidder | The vendor must submit documented proof, such as certifications, test reports, or case studies, to demonstrate compliance with the required specifications.
51 |  | The platform should be able to perform downstream quality control for each batch of iPS lines (42 lines, each line 4 clones) including .(see Appendix workflow - 9): -Setting up trilineage differentiation assays (involving single cell seeding of the cells). -Conducting RT-PCR to assess the expression of stem cell and pluripotency markers in the newly generated lines (invloving harvesting of cells, RNA isolation, RNA measurement, cDNA synthesis and RT-PCR). -Collecting material for further analysis (harvesting of cells for DNA and RNA isolation, and DNA and RNA measurment). | KO | To be filled in by the bidder | The vendor must submit documented proof, such as certifications, test reports, or case studies, to demonstrate compliance with the required specifications.
52 |  | The platform should be able to collect medium for Mycoplasma and sterility testing and should document the result of these tests. (see Appendix workflow - 9) | KO | To be filled in by the bidder | The vendor must submit documented proof, such as certifications, test reports, or case studies, to demonstrate compliance with the required specifications.
53 |  | The platform should be able to perform Mycoplasma PCR test after expansion of each clone of each batch of iPS lines and should document and communicate the result of these tests to the central software. (see Appendix workflow - 9) | KO | To be filled in by the bidder | The vendor must submit documented proof, such as certifications, test reports, or case studies, to demonstrate compliance with the required specifications.
54 |  | The platform should able to harvest one well of each expanded clone for RNA isolation. (see Appendix workflow - 9) | KO | To be filled in by the bidder | The vendor must submit documented proof, such as certifications, test reports, or case studies, to demonstrate compliance with the required specifications.
55 |  | After the maintenance, expansion, and QC of three iPS batches (in three cycles) according to the workflow are completed, the platform will undergo maintenance. | KO | To be filled in by the bidder | The vendor must submit documented proof, such as certifications, test reports, or case studies, to demonstrate compliance with the required specifications.
56 |  | Once maintenance is complete, the platform begins processing the next three batches in a new term. | KO | To be filled in by the bidder | The vendor must submit documented proof, such as certifications, test reports, or case studies, to demonstrate compliance with the required specifications.
57 |  | The system should be designed for scalability to enhance throughput and support growing operational needs. | KO | To be filled in by the bidder | The vendor must submit documented proof, such as certifications, test reports, or case studies, to demonstrate compliance with the required specifications.
58 |  | The platform should automate manual tasks, enabling parallel operations that enhance both throughput and quality. It should also provide improved control and monitoring to ensure process consistency and repeatability. | KO | To be filled in by the bidder | The vendor must submit documented proof, such as certifications, test reports, or case studies, to demonstrate compliance with the required specifications.
59 |  | The system must be designed for long-term reliable and continuous operation over a timespan of at least 60 hours without operator support. | KO | To be filled in by the bidder | The vendor must submit documented proof, such as certifications, test reports, or case studies, to demonstrate compliance with the required specifications.
60 |  | The system must be designed for long-term reliable and continuous operation over a timespan of at least 10 weeks without interference due to system maintenance. | KO | To be filled in by the bidder | The vendor must submit documented proof, such as certifications, test reports, or case studies, to demonstrate compliance with the required specifications.
61 |  | The system should ensure that plates that are being moved are not dropped during a brief power interruption. | KO | To be filled in by the bidder | The vendor must submit documented proof, such as certifications, test reports, or case studies, to demonstrate compliance with the required specifications.
62 |  | The automation platform must be flexible and scalable to meet both current and future research needs. The ability to rapidly reconfigure workflows and system layouts will be an important advantage. | KO | To be filled in by the bidder | Please add configura
63 |  | Must have a configurable database to collect system data. | KO | To be filled in by the bidder
64 |  | There must be an audit trail of events in the system. | KO | To be filled in by the bidder
65 |  | Must record telemetric data from the automated platform, including temperature and humidity measurements. | KO | To be filled in by the bidder
66 |  | Must offer real time Gantt charts to show process progress. | KO | To be filled in by the bidder
67 |  | Should be able to change or disable resources during a run with automatic pooling of available resources. | KO | To be filled in by the bidder
68 |  | The user interface should be logically structured and require a training of maximal three days. | KO | To be filled in by the bidder
69 |  | Must have a configurable database to collect system data. | KO | To be filled in by the bidder
70 |  | Should be able to accommodate error handling on the fly without impacting on the entire process runs. | KO | To be filled in by the bidder
71 |  | Should provide both time resolved and event driven scheduling | KO | To be filled in by the bidder
72 |  | Should offer scripting capability,with high level commands to start the next step of a sample. | KO | To be filled in by the bidder
Scheduling software
73 |  | The software must provide a well-structured and visually clear interface for selected instruments accessible on both web browsers and mobile devices. | KO | To be filled in by the bidder
74 |  | Include tools for graphical planning of samples, devices, and resource utilization, alongside real-time simulation capabilities. | KO | To be filled in by the bidder
75 |  | Provide the ability to define and adjust protocols during active experiments, automatically managing changes across all subsequent steps. | KO | To be filled in by the bidder
76 |  | Ensure the system can seamlessly scale to accommodate growth in users and workload without compromising on speed or efficiency. | KO | To be filled in by the bidder | The vendor must submit documented proof, such as certifications, test reports, or case studies, to demonstrate compliance with the required specifications.
77 |  | Incorporate mechanisms to recover ongoing experiments and samples even in the event of critical failures like power outages. | KO | To be filled in by the bidder
78 |  | Should automatic re-start after a power cycle. | KO | To be filled in by the bidder
79 |  | Enable creation, assignment, and modification of tasks and events, with resource tracking, conflict resolution, and prioritization options. | KO | To be filled in by the bidder
80 |  | Deliver immediate updates on schedule changes with notifications to relevant users. | KO | To be filled in by the bidder
81 |  | Support generation of tailored reports and dashboards, alongside detailed performance metrics and time tracking. | KO | To be filled in by the bidder
82 |  | Allow protocols to manage diverse labware and plate processing without requiring protocol adjustments. | KO | To be filled in by the bidder
83 |  | Include functionality to pause new processes while allowing in-progress tasks to conclude efficiently. | KO | To be filled in by the bidder
84 |  | Implement automated barcode scanning to update physical locations and inform users of discrepancies in ordered items. | KO | To be filled in by the bidder
85 |  | Provide seamless import/export options for data, scripting Application Programming Interfaces (APIs) for system integrations, and open web service messaging. | KO | To be filled in by the bidder
86 |  | Ensure compatibility with future LIMS systems through configurable connectivity. | KO | To be filled in by the bidder
87 |  | Enable direct input of experiment and curated lists into the software for streamlined processing. | KO | To be filled in by the bidder
88 |  | Offer secure cloud-based deployment options with redundant data storage and regular automated backups for easy recovery. | KO | To be filled in by the bidder
89 |  | Include tools for tracking analytics and utilization metrics, ensuring optimal system performance. | KO | To be filled in by the bidder
90 |  | Updates and new versions of the software required for the operation of the device are provided free of charge until SAT. | KO | To be filled in by the bidder
Robot arm
91 |  | The platform should include a robot arm. | KO | To be filled in by the bidder
92 |  | The robot arm must be collaborative and safe to use | KO | To be filled in by the bidder
93 |  | The robot arm must be drag to teach and easy to configure new labware and instruments on the system | KO | To be filled in by the bidder
94 |  | The robot arm should have continuous 360°rotation at wrist and waist to optimise processing of the workflow | KO | To be filled in by the bidder
95 |  | The robot arm must have an integrated bar code scanner on the wrist | KO | To be filled in by the bidder
96 |  | The robot arm should be able to perform cross movements to evenly spread the iPS clumps after passaging in the well. | KO | To be filled in by the bidder | The vendor must submit documented proof, such as certifications, test reports, or case studies, to demonstrate compliance with the required specifications.
97 |  | The robot arm should have smooth movement profile to avoid attachment of iPS clumps during movement. | KO | To be filled in by the bidder
98 |  | The robot arm should be fully capable of accessing all critical positions on the liquid handler's deck, ensuring efficient handling of time-consuming tasks like the passaging of iPSCs. | KO | To be filled in by the bidder | The vendor must submit documented proof, such as certifications, test reports, or case studies, to demonstrate compliance with the required specifications.
99 |  | The arm must be able to pick and uncap the tip cap and place it in the right position and track the empty tip boxes and plates. | KO | To be filled in by the bidder
Incubators
100 |  | The plaform should have one or more incubator(s). | KO | To be filled in by the bidder
101 |  | The incubator(s) should have a total storage of minimum 300 low density plates. | KO | To be filled in by the bidder
102 |  | 50% of incubator capacity must be capable of run at hypoxic conditions | KO | To be filled in by the bidder
103 |  | The incubators must have programmable temprature approx. 37C, Co2 5%-20%. | KO | To be filled in by the bidder
104 |  | The incubators must have a decontamination protocol and proven to be contamination free. | KO | To be filled in by the bidder
105 |  | The incubators must be accesible to the robot arm to load and unload plates. | KO | To be filled in by the bidder
106 |  | The platform should include a plate sealer. | KO | To be filled in by the bidder
Media fridge & delivery system
107 |  | The supplier must ensure sufficient capacity for culture mediums, PBS, passaging reagents, coating material, (coated) plates and tips to support 60 hours of continuous cell culturing. | KO | To be filled in by the bidder
108 |  | The supplier must ensure sufficient capacity for DNA/RNA isolation reagents, cDNA synthesis, (q)PCR reagents during selected schedules. | KO | To be filled in by the bidder
109 |  | The supplier must ensure sufficient capacity to accommodate all the lab culture ware required for performing every expansion step outlined in the SOP. | KO | To be filled in by the bidder
110 |  | The supplier must ensure sufficient capacity to accommodate all the lab culture ware required for performing differentiation outlined in the SOP. | KO | To be filled in by the bidder
Additionally, Excel tab; Additional apparatus lists. includes automated equipment required for downstream analysis, which must also be integrated into the platform. However, the equipment in Additional apparatus lists is acquired independent of the tender, and suppliers are not required to bid on it. This Excel tab serves only to inform the integrator about the requirements and preferences for downstream analysis. Suppliers are welcome to submit a separate quote for these instruments directly to the Erasmus MC iPS core facility.
111 |  | All of these items must also be integrated, and Erasmus MC will supply this equipment separately, as outlined in the Excel tab 'Additional Apparatus Lists'. | KO | To be filled in by the bidder
ICT
IT: Regulatory and compliancy
112 | Compliancy | The system must be compliant with (international) laws and regulations in the field of privacy protection and data protection | KO | To be filled in by the bidder
115 | Privacy | The system complies to General Data Protection Regulation(GDPR). | KO | To be filled in by the bidder
116 | Archiving | The system has means and procedures for archiving of data. These means and procedures are in accordance with legislative and regulatory requirements. | KO | To be filled in by the bidder
117 | Changes in laws and legislation | All future changes in applicable laws and legislation will be supported by the system as soon as new laws and legislation will become effective. | KO | To be filled in by the bidder
118 | FDA Title 21 CFR Part 11 | Proven compliancy with FDA CFR 21 Part 11. | KO | To be filled in by the bidder
119 | Good Laboratory Practice | Systems conforms to Good Laboratory Practice (GLP). | KO | To be filled in by the bidder
120 | Property | All data will be property of the Erasmus MC. The supplier is prohibited to use the data for other purposes without concent of Erasmus MC. The data should be transferred to the Erasmus MC on demand (for example when license period expires). | KO | To be filled in by the bidder
121 | Availability of system after expiration of license | When the licence expires the supplier ensures his cooperation in a smooth migration to the new supplier and system. | KO | To be filled in by the bidder
IT: Roles and authorisation
122 | Authorization management | The system has a flexible and centrally managed authorization functionality based on groups and roles. Based on these groups and roles the researcher has access or is denied access to certain functionalities and / or data. | KO | To be filled in by the bidder
123 | Multiple roles / groups per user | A user can be authorized for multiple groups and roles. It is also possible to provide an individual researcher custom authorizations. | KO | To be filled in by the bidder
124 | Password demands | Strong password authentication supported by the system (minimal length 8 characters, combination of digits, characters and at least one special character). | KO | To be filled in by the bidder
125 | System breach prevention | System prohibits unlimited login attempts (configurable amount of login attempts). | KO | To be filled in by the bidder
IT: Security and Privacy
126 | Documentation in writing | The documentation of the data models, software data communication protocols and parts of the processing of the system is fully stated in writing. | KO | To be filled in by the bidder
128 | GDPR compliant audit trail | Audit trail is compliant with GDPR. | KO | To be filled in by the bidder
129 | Audit trail | Logging in the audit trail contains 1: type of action (e.g. delete), 2: time of action, 3: who performed the action, 4: the action is performed on what? (e.g. username) | KO | To be filled in by the bidder
130 | Audit trail on user logins | Audit trail on user logins | KO | To be filled in by the bidder
131 | Data transportation | The system is designed and implemented in a way that the data can be transported to another environment / platform (‘portability’). | KO | To be filled in by the bidder
132 | Data archiving | Erasmus MC is without assistance of the vendor able to archive data from the server. | KO | To be filled in by the bidder
133 | Safe removal of hardware | Removal of hardware containing data is done in a safe and certified manner. | KO | To be filled in by the bidder
134 | App privacy | Privacy of users and data is safeguarded when the system is used as an APP on mobile devices | KO | To be filled in by the bidder
135 | Automatic time-out if idle | System has automatic time-out if session is idle. | KO | To be filled in by the bidder
136 | Data in transit encryption | Data in transit is secured with the use of encryption (minimal 4.096 bit and use of a strong algorithm). Internet access is https with certificate of A or A+ level. | KO | To be filled in by the bidder
137 | Remote access for administrators | Personnel responsible for system management of the system only has access to the system in a safe way: - use of VPN connection (EMC validated) - 2 factor authentication - session time out of connection. | KO | To be filled in by the bidder
138 | System Hardening | The hard- and software on which the system is hosted, is demonstrable hardened against hacker attacks by removing default configurations, system passwords, etc. | KO | To be filled in by the bidder
139 | Administrator access | Erasmus MC has the right to get insight in how access of system administrators of the supplier is organized | KO | To be filled in by the bidder
140 | Couplings | Couplings between Erasmus MC and supplier are reviewed on security aspects by an Erasmus MC architect or Technical Security Officer | KO | To be filled in by the bidder
141 | Solving security incidents | Guidelines and procedures for the registration and solving of security incidents are stated in writing. | KO | To be filled in by the bidder
IT: General
142 | Web application | If the system is a web application, which is supported by generally used web browsers, in any case Microsoft Edge, Google Chrome, Apple Safari, Mozilla Firefox. | KO | To be filled in by the bidder
143 | Hybrid SaaS | The system is offered as Software as a Service (SaaS) with the ability to store data on Erasmus MC servers. | KO | To be filled in by the bidder
144 | Backup / restore | The system has backup and restore functionalities for all data. | KO | To be filled in by the bidder
145 | Maximum Data Loss | All data is backuped daily or more frequently. | KO | To be filled in by the bidder
146 | Access limitation and retainment policies | Access limitation and retainment policies exist to limit chances of data loss from backups. | KO | To be filled in by the bidder
147 | Session management | The system applies session management which ensures that no data from different user sessions is mixed. | KO | To be filled in by the bidder
148 | SSL security | The systems web based user interface runs exclusively by a secured https-connection (TCP port 443) and can be equipped with an official SSL certificate. | KO | To be filled in by the bidder
149 | Database access | The database of the system is directly and fully accessible by ODBC or comparable connection. | KO | To be filled in by the bidder
150 | Data model | All data in the system is structured and stored in a relational database. The data model is provided by the system supplier. | KO | To be filled in by the bidder
IT: Maintenance and support
151 | Updates and maintenance | Updates and upgrades may never lead to degradation of the functionality of the system. The supplier guarantees that all demands written in this list of demands are upheld after updates and upgrades are installed. | KO | To be filled in by the bidder
152 | Updates and maintenance | Software follows lifecycle of the operating system. For example compatbile with Windows 11 and future updates. | KO | To be filled in by the bidder
153 | Skip update | The Erasmus MC can choose to skip at least one upgrade. This doesn't result in the diminishment of guarantees by the supplier. | KO | To be filled in by the bidder
154 | Version management | System has a clear tracking system of different versions of the system, with version numbers, date published and a user defined description. | KO | To be filled in by the bidder
155 | Service Desk | The supplier has an English speaking service desk. | KO | To be filled in by the bidder
IT: Training & Help
156 | Manuals | Manuals and training are available to users and administrators in English | KO | To be filled in by the bidder
157 | Contextual help | Contextual help is available in the system | KO | To be filled in by the bidder
158 | Frequently asked questions | A list of frequently asked questions and their answers is available for reference | KO | To be filled in by the bidder
159 | E-learning | E-learning is available for the main functions of the system, it will also be kept up-to-date for new functions. | KO | To be filled in by the bidder
Maintenance Service, Response time and Repair
160 |  | Service includes; System design and documentation, Design Agreement System build, Software configuration, Method development Integration testing, Travel & accommodation. | KO | To be filled in by the bidder
161 |  | The supplier must provide a minimum 1-year warranty on all equipment delivered by the supplier. | KO | To be filled in by the bidder
162 |  | In case of equipment failure, the supplier must provide on-site repair services within 24 hours for critical components, with a maximum downtime of 48 hours for non-critical issues. | KO | To be filled in by the bidder
163 |  | Require the vendor to provide on-site repair services for issues that cannot be resolved remotely. | KO | To be filled in by the bidder
164 |  | Request the ability to troubleshoot issues remotely via secure connections, reducing downtime and minimizing the need for on-site visits. | KO | To be filled in by the bidder
165 |  | The vendor must provide detailed repair and service reports after each repair event, including the root cause, actions taken, and any parts replaced. | KO | To be filled in by the bidder
166 |  | Service includes Operator Training of at least 5 lab personel.(Vendor will provide system training to the current employees of iPS core team) | KO | To be filled in by the bidder
167 |  | The vendor must provide training for in-house maintenance by staff to handle repairs and maintenance. | KO | To be filled in by the bidder
168 |  | The vendor is responsibile for damage to the automated platform during the repair or maintenance process. | KO | To be filled in by the bidder
Documentation and Manuals
169 |  | A complete technical manual will be provided in English. | KO | To be filled in by the bidder
Technical
170 |  | The supplier must provide a full validation package for Installation Qualification (IQ), Operational Qualification (OQ), and Performance Qualification (PQ) to be performed upon installation at the site. | KO | To be filled in by the bidder
171 |  | The system must include automated QC processes to detect and alert operators to any out-of-spec conditions in real-time. | KO | To be filled in by the bidder
172 |  | The system must demonstrate that the setup is able to maintain 84 cell lines in culture per day. | KO | To be filled in by the bidder
Delivery Time, Milstones, Sanctions & Logistics
173 |  | The setup of the platform and the on-site training of at least 5 iPS core facility employees must be completed within 1 month after installation, including troubleshooting and emergency procedures. | KO | To be filled in by the bidder
174 |  | All end-users must be able to operate the equipment independently after the training by the bidder, without further assistance. | KO | To be filled in by the bidder
175 |  | If the project is delayed by more than 30 calendar days beyond the agreed delivery date, the issue will be escalated to senior management, and a formal meeting will be scheduled to discuss potential penalties, revised timelines, or further sanctions | KO | To be filled in by the bidder
176 |  | If the delivery is delayed by more than 90 calendar days from the final agreed-upon date without agreed upon reasons, the client reserves the right to terminate the contract without any further financial obligations, except for payment for work completed up to that point. | KO | To be filled in by the bidder
177 |  | Upon premature termination of the contract all acquired instruments paid by Erasmus MC should be shipped and delivered to the Erasmus MC. | KO | To be filled in by the bidder
178 |  | The client reserves the right to withhold 10% of the total contract value until final acceptance and successful system testing are completed, following the final Site Acceptance Test (SAT). | KO | To be filled in by the bidder
179 |  | The vendor will be liable for any costs incurred due to delays in delivery, including extra operational costs or expedited shipping of required components, up to a maximum of 15% of the total contract value. | KO | To be filled in by the bidder
180 |  | The shipment size of instruments packages should be specified. | KO | To be filled in by the bidder
181 |  | The supplier proactively informs the client about any changes in delivery times for equipment and required consumables. | KO | To be filled in by the bidder
182 |  | The supplier is required to provide a timely and accurate estimation of the delivery time. Failure to meet this requirement may result in sanctions. All communication regarding the delivery schedule must be transparent and free from exaggeration or misleading information. | KO | To be filled in by the bidder
Inkoop
183 |  | There are no costs associated with submitting a tender, and the bidder derives no rights (now or in the future) from being permitted to submit a tender at this time. | KO | To be filled in by the bidder
184 |  | The bidder agrees to the entire procedure described and all conditions as outlined in the accompanying Offer Guidelines, the other tender documents, and the attachments. | KO | To be filled in by the bidder
185 |  | The bidder agrees to the agreements and provisions contained therein, which are included as attachments to the Offer Guidelines. | KO | To be filled in by the bidder
186 |  | The provisions of the agreements included in the attachments to the Offer Guidelines shall prevail. For situations not covered by the provisions of these (framework) agreements, the NFU General Purchasing Conditions, version January 2024 (Appendix 8), shall apply. These general purchasing conditions are included as Appendix. | KO | To be filled in by the bidder
187 |  | The bidder accepts the applicability of the NFU purchasing conditions, to the explicit exclusion of your own general terms and conditions. | KO | To be filled in by the bidder
188 |  | The final agreements resulting from this Offer Guidelines will be based on all provided information and documentation. The bidder guarantees that their organization possesses the necessary capacities, skills, and resources to (continue to) meet the Program of Requirements and Preferences as set out in this Offer Guidelines. | KO | To be filled in by the bidder
189 |  | Any substantial change in the bidder's situation that may affect their ability to meet these requirements must be promptly communicated to Erasmus MC. | KO | To be filled in by the bidder
190 |  | The bidder hereby declares that they have reviewed the (draft) purchase agreement and the associated conditions applicable to this assignment and, by submitting this PvE, confirms their agreement with the contract conditions outlined therein. | KO | To be filled in by the bidder
191 |  | The bidder hereby declares to be responsible for the accuracy and completeness of all provided information. The bidder is aware that, if it is later found that incorrect and/or incomplete information has been provided, they may be excluded from further participation, or any agreements already made may be canceled without legal consequences for Erasmus MC. | KO | To be filled in by the bidder
192 |  | The bidder has fully completed the price sheet as included in Appendix 07 and Appendix it to the submission. | KO | To be filled in by the bidder
193 |  | The prices offered in the price sheet are binding. | KO | To be filled in by the bidder
194 | Information sharing | Information concerning ErasmusMC is not presented to third parties without previous approval of the ErasmusMC. | KO | To be filled in by the bidder
Sustainability
195 |  | The supplier shall complete the following tabs: 1. Instruction, 2. Company, and 3. Sustainability. | KO | To be filled in by the bidder

## 1. Wishes Program for Auto
Bijlage 1b. Programma van Wensen Automation & Integration
The sub-award criterion 'Program of Requirements' is worth a maximum of 484 points
For each wish, please indicate whether it applies by selecting "yes" or "no". For some wishes, additional information is requested. Based on your response, a score will be assigned. The maximum number of points that can be awarded per wish is shown under the ‘Score’ column.
REF | Categorie | Omschrijving | score (0-44) | Fill in with Yes or No; failure to fill in is not allowed. | Any explanation | Score toegekend door Erasmus MC
| General conditions of the platform. |  | 65
1 |  | The platform enclosure and instruments are compatible with UV decontamination. | 10 | To be filled in by the bidder
2 |  | It would be desirable for the platform to allow the placement of a portable UV lamp. | 10 | To be filled in by the bidder
3 |  | It would be desirable for the platform to be upgradeable with a -80°C cryo-storage in the future. | 5 | To be filled in by the bidder
4 |  | It would be desirable for the platform to be upgrade ready for inclusion of a pipeline that facilitates freezing of each batch of iPS lines (42 lines) per cyclus.((see attached workflow - Appendix 9) | 10 | To be filled in by the bidder | Please elaborate and add proof of concept
5 |  | The enclosure provides sufficient space to accommodate future upgrades. | 10 | To be filled in by the bidder
6 |  | It would be desirable to have real time visual feedback of both process and the safety status. | 10 | To be filled in by the bidder | Explanation/Example
7 |  | It would be desirable to have a solution for avoiding the attachment of iPS clumps immediately after passaging. | 10 | To be filled in by the bidder | Please elaborate and add proof of concept
| Automation and Integration |  | 28
8 |  | Incorporating of artificial intelligence or machine vision technologies would be highly beneficial for enhancing system capabilities. | 5 | To be filled in by the bidder | Explanation/Example; Higest score for system that has the possibility
9 |  | The ability to upgrade the automation for freezing of PSCs. | 10 | To be filled in by the bidder
10 |  | System cabelling should be managed, labelled a separated according to LVD | 3 | To be filled in by the bidder
11 |  | The vendor should provide a Declaration of Conformity (DoC) for the complete system. | 10 | To be filled in by the bidder
| Robot Arm |  | 60
12 |  | We wish for the robot arm to include integrated sensors or feedback mechanisms that can detect and recover from errors autonomously, minimizing downtime. | 10 | To be filled in by the bidder | Please elaborate and add proof of concept
13 |  | We wish for the platform to support adaptive scheduling, allowing the robot arm to dynamically adjust tasks and priorities based on real-time workflow requirements. | 10 | To be filled in by the bidder | Please elaborate and add proof of concept
14 |  | We wish for seamless integration with laboratory information management systems (LIMS) for real-time data exchange and process tracking. | 10 | To be filled in by the bidder
15 |  | We wish for a compact, modular robot arm design to allow for easy upgrades, scalability, and adaptation to different laboratory configurations. | 10 | To be filled in by the bidder | Please elaborate and add proof of concept
16 |  | We wish for the robot arm to be compatible with a wide range of laboratory instruments, enabling diverse applications without additional customization. | 10 | To be filled in by the bidder | Overview table and a example
17 |  | We wish for the robot arm to include remote monitoring and control capabilities, allowing users to supervise and adjust operations from outside the lab. | 10 | To be filled in by the bidder | Please elaborate in which form: Wired Remote Control, Wireless Remote Control, Software-Based Control, App-Based & Mobile Control
| Incubator |  | 15
18 |  | The movement in the incubator does not cause the iPS cell clumps to attach to each other. | 10 | To be filled in by the bidder | Please elaborate and add proof of concept
19 |  | The incubator is seamlessly welded, corrosion-resistant, and equipped with a stainless steel chamber with rounded edges for easy and thorough cleaning. | 5 | To be filled in by the bidder | Explanation/Example
| Media fridge & delivery system |  | 39
20 |  | It would be desirable for the platform to be capable of cold storage of medium for at least 60 hours. | 10 | To be filled in by the bidder | Please elaborate where, in which form)
21 |  | It would be desirable for the platform to be capable of storing half-liter aliquots of medium. | 4 | To be filled in by the bidder | Explanation/Example
22 |  | It would be desirable for the platform to be able to warm up half-liter aliquots of medium to room temperature within 30 to 60 minutes before medium refreshment or passaging cells. | 10 | To be filled in by the bidder
23 |  | It would be desirable for the platform to be able to coat needed plates for minimal of 60 hours. | 10 | To be filled in by the bidder
24 |  | The supplier must ensure the sufficient capacity to accommodate all the lab culture ware required for performing cryopreservation step outlined in the SOP. | 5 | To be filled in by the bidder
| Process scheduling software |  | 52
25 |  | It would be desirable to add AI-based optimization to adjust schedules in real-time for efficiency, resource usage, or cost reduction. | 5 | To be filled in by the bidder | Explanation/Example
26 |  | Data import/export capabilities, supporting multiple file formats such as CSV, Excel, XML, etc. | 5 | To be filled in by the bidder | Overview table and a example
27 |  | It would be desirable to have support for devices for real-time tracking and availability of equipment. | 5 | To be filled in by the bidder | Overview table and a example
28 |  | Ability to set up different scheduling views (e.g., daily, weekly, resource-specific). | 5 | To be filled in by the bidder | Please elaborate and add proof of concept
29 |  | Intuitive, user-friendly interface with drag-and-drop scheduling, task prioritization, and easy access to features. | 3 | To be filled in by the bidder | Explanation/Example
30 |  | Features for tracking and reporting on costs associated with scheduling (e.g., overtime, resources used, equipment). | 3 | To be filled in by the bidder | Explanation/Example
31 |  | Cost optimization algorithms for reducing operational expenses through efficient scheduling. | 5 | To be filled in by the bidder | Explanation/Example
32 |  | Budgeting tools that integrate with financial systems to track project or operational budget adherence. | 8 | To be filled in by the bidder | Explanation/Example
33 |  | User support portal with FAQs, troubleshooting guides, and helpdesk integration. | 5 | To be filled in by the bidder
34 |  | Regular software updates and patches to ensure the system is up to date. | 3 | To be filled in by the bidder
35 |  | Dedicated customer support team available through multiple channels (e.g., email, phone, live chat). | 5 | To be filled in by the bidder
| Operation, display, and user interface |  | 30
36 |  | Real-time monitoring of ongoing processes with dashboards displaying current statuses. | 10 | To be filled in by the bidder | Explanation/Example
37 |  | Alert system that notifies users of critical events such as delays, failures, or maintenance requirements. | 10 | To be filled in by the bidder | Explanation/Example
38 |  | Customizable alert settings based on time, severity, or specific criteria. | 10 | To be filled in by the bidder | Explanation/Example
65 |  |  | 1 | To be filled in by the bidder | n.v.t.
ICT
| IT: Regulatory and compliancy |  | 15
39 |  | The platform should be compatible with existing laboratory information management systems (LIMS) and support communication via OPC, REST API, or other industry-standard protocols??? | 3 | To be filled in by the bidder
40 | Central Authentication Service | Authentication of users is done via central authentication services of Erasmus MC (Active Directory). Erasmus MC uses AD (Active Directory). For Windows Based systems authentication via LDAP must be possible. | 3 | To be filled in by the bidder
41 | Continuity | The supplier guarantees to deliver service and maintenance for at least ten years, calculated from the start of agreement. | 3 | To be filled in by the bidder
42 | Updates and maintenance | The supplier delivers needed updates and upgrades for the system. Updates and upgrades are included within development fee. | 3 | To be filled in by the bidder
43 | Hosting | Hosting of the system takes place in a country that complies to the EU Directive 95/46/EC concerning privacy. | 3 | To be filled in by the bidder
| IT: Roles and authorisation |  | 30
44 | Central Authentication Service | Authentication of users is done via central authentication services of Erasmus MC (Active Directory). Erasmus MC uses AD (Active Directory). For Windows Based systems authentication via LDAP must be possible. | 3 | To be filled in by the bidder
45 | Guest account | Users without an EramsusMC Account (guest account) can be set-up to have access to the system with authentication through username and password verification | 3 | To be filled in by the bidder
46 | Account expiration | For guest accounts, passwords can be configured to automatically expire after a given period (This is not applicable with single sign-on) | 3 | To be filled in by the bidder
47 | Password expiration after period | For guest accounts, passwords handed out upon creation of an account (by administrator or by system) expire after a configurable period | 3 | To be filled in by the bidder
48 | Password expiration warning | For guest accounts, automatic warning of password expiration is available | 3 | To be filled in by the bidder
49 | Password reset | For guest accounts, administrators can send out global (massive) password reset requests (e.g. by e-mail) | 3 | To be filled in by the bidder
50 | Guest account reset | For guest accounts an inactivity time can be configured(e.g. 12 months) after which accounts are automatically disabled. Reason for disabling account should be registered, reactivation of accounts is possible | 3 | To be filled in by the bidder
51 | Forced password change | For guest accounts the system can block accounts or force change of password | 3 | To be filled in by the bidder
52 | Password reset by mail/sms | For guest accounts a new password can be requested to the system with temporary password/security code in mail/SMS | 3 | To be filled in by the bidder
53 | system-to-system accounts | Provide accounts for automated processing of data (system-to-system accounts), using a policy that takes into account that these accounts are not used by humans but by machines, while safeguarding the safety of the system (e.g. no password expiration but IP restriction) | 3 | To be filled in by the bidder
| IT: Security and Privacy |  | 21
New |  | The provider is ISO 27001 certified, or can demonstrably operate according to the requirements of the ISO/IEC 27001 standard for information security. | 3 | To be filled in by the bidder
54 | Layered architecture | The system has a layered architecture divided by: - a part for authorized users - within the authorized segment a part for information exchange ('front-end') and a part for data storage and data processing ('back-end'). | 3 | To be filled in by the bidder
55 | OTAP environment | Environments for development, testing and production of the virtual evironment are separated. Developments and updates can be pushed from one evironemnt to another. | 3 | To be filled in by the bidder
56 | Penetration testing (combi hosting) | If application has a webservice. A penetration test on the system is performed prior to the production date and annually afterwards. Erasmus MC is informed regarding the results of the test. | 3 | To be filled in by the bidder
57 | OWASP standard | Server/internet security is according to a commonly used standard (e.g. OWASP), including forced SSL connections, error messaging not revealing technical details, immediate error reporting to super users, intrusion detection, IP or account blocking, etc. | 3 | To be filled in by the bidder
58 |  | The data is buffered for at least 12 hours in case of network interruption or loss of connectivity with the LIMS. | 2 | To be filled in by the bidder
59 | Geographical location | Backups are stored on another geographical location | 2 | To be filled in by the bidder
60 |  | After a network interruption, the analyzers automatically send all unsent results to the middleware. | 2 | To be filled in by the bidder
| Maintenance Service, Respons time and Repair (We request that prospective suppliers submit a plan detailing how they will arrange maintenance and service (including repairs and support). The following topics will be evaluated. ) |  | 85
61 |  | The instrument can be remotely controlled. | 10 | To be filled in by the bidder | Please elaborate in which form: Wired Remote Control, Wireless Remote Control, Software-Based Control, App-Based & Mobile Control
62 |  | It is desirable the warranty begins from the moment the equipments are successfully integrated and operational at Eramus MC iPS core facility. | 10 | To be filled in by the bidder
63 |  | In the event of a system failure, it is desirable the vendor has a respond time within 2 hours and resolve critical failures within 24 hours. | 10 | To be filled in by the bidder | Please elaborate
64 |  | We wish for an access to 24/7 technical support to ensure uninterrupted system functionality. | 10 | To be filled in by the bidder
65 |  | The operator receives remote warnings. | 10 | To be filled in by the bidder | Explanation/Example
66 |  | Evidence should be provided for availability of spare parts and required installation time. | 8 | To be filled in by the bidder | overview/table
67 |  | Spare parts should be available for a minimum of 10 years after installation. | 7 | To be filled in by the bidder | Please elaborate
68 |  | The vendor is expected to provide detailed suggestions for a comprehensive maintenance strategy (executed by Erasmus MC) , including the frequency and scope of routine maintenance (e.g., monthly, quarterly, or annually) to ensure optimal platform performance, a preventative maintenance plan designed to minimize downtime and extend equipment lifespan, and a system for maintaining accurate logs of all maintenance activities, including repairs and inspections. | 4 | To be filled in by the bidder | Please elaborate
69 |  | The vendor is required to provide an escalation model. This model must clearly outline the steps to be taken when escalating issues or incidents, including the parties involved, timelines, and responsibilities. | 4 | To be filled in by the bidder | Please elaborate
70 |  | It would be desirable that the vendor provides support for equipment and software even after the end-of-life period, with clear terms for continued service or parts availability (equipment supplied by the automation specialist). | 4 | To be filled in by the bidder | Please elaborate
71 |  | It would be desirable if the vendor can provide any upgrade or replacement conditions if equipment becomes obsolete or incompatible with future technological developments (equipment supplied by the automation specialist). | 2 | To be filled in by the bidder | Please elaborate
72 |  | The tenderer shall perform a release run/system check after each repair, relocation or change to the software, after which the supplier shall officially release the system for use. | 3 | To be filled in by the bidder
73 |  | A user forum/community to exchange experience, tips and tricks is present | 3 | To be filled in by the bidder | Please elaborate
| Sustainability |  | 44
74 |  | The supplier shall complete the following tabs: 1. Instruction, 2. Company, and 3. Sustainability. | 44 | To be filled in by the bidder
484
Rechtsgeldige ondertekening:

## 2. imager Requirments
Programma van Eisen 2a. Imager
Failure to meet one or more (exclusion) knock-out (KO) criteria in this Requirements Document will result in your submission no longer being considered for the award of this contract. Meeting all knock-out criteria is regarded as a mandatory condition in this European Tender. Compliance with the Requirements Document does not earn any points.
The (question) statements in this Requirements Document are primarily intended to determine: 'Can the supplier meet the specified description?' Failure to meet one (or more) 'knock-out criteria' will result in rejection. Incorrect or incomplete filling in will lead to the immediate invalidation of the submission. The Requirements Document must be completed truthfully. PLEASE NOTE: The Excel sheet is not locked; the bidder should only fill in the fields that are highlighted in yellow and where it is indicated that something should be filled in.
|  |  | KO | Fill in with Yes or No; failure to fill in is not allowed. | Any explanation | Proof, documentation to be provided with the submission
Automated imager
1 |  | The imager should comply with all relevant requirments for an automated platform as listed in Requirments for Automation. | KO | To be filled in by the bidder
2 |  | Brightfield, 2x and 4x resolution with option for upgrates to higher resolution | KO | To be filled in by the bidder
3 |  | The imaging system must include detailed specifications regarding the exact time required to scan and process samples in both a 6-well plate and a 12-well plate. The vendor must provide verifiable data demonstrating these processing times under standard operating conditions. | KO | To be filled in by the bidder | The vendor must submit supporting evidence when responding 'yes,' demonstrating the time required time for scanning.
4 |  | Microwell plates (SBS formats 6, 12, 24, 48, 96 and 384 well) | KO | To be filled in by the bidder
5 |  | The imager should have a fast measurement time for 12 well plate and 6-well plates. Being able to load, image and eject plates to enabe the platform to refresh and expand a minimal of 84 cell lines with in 24 hours. | KO | To be filled in by the bidder | The vendor must submit supporting evidence when responding 'yes,' demonstrating the required time frame.
6 |  | The imager should have smooth movement profiles to avoid attachment of the iPS clumps during movement. (loading and ejection) | KO | To be filled in by the bidder
7 |  | The imager should be suitable for microwell plates. | KO | To be filled in by the bidder
8 |  | The imager should have imaging enhancement software, including capture of well edges. | KO | To be filled in by the bidder | The vendor must submit supporting evidence when responding 'yes'.
9 |  | The imager should be capable of ultrafast multiplex whole well imaging. | KO | To be filled in by the bidder | The vendor must submit supporting evidence when responding 'yes'.
10 |  | The imager should be capable of improved harmonic motion for imaging without agitation during plate scan. | KO | To be filled in by the bidder
11 |  | The imager should be able to perform image analysis during measurement. | KO | To be filled in by the bidder
12 |  | The imager should perform iPSC confluency measurements per well and provide information about quality of the PSCs, including random differentiation status of the colonies and morphology. | KO | To be filled in by the bidder | The vendor must submit supporting evidence when responding 'yes'.
13 |  | The imager should calculate/export/report measurements, for interpretation and decision making in follow up steps | KO | To be filled in by the bidder | The vendor must submit supporting evidence when responding 'yes'.
14 |  | All end-users must be able to operate the equipment independently after the training by the bidder, without further assistance. | KO | To be filled in by the bidder
ICT
IT: Regulatory and compliancy
15 | Good Laboratory Practice | Systems conforms to Good Laboratory Practice (GLP). | KO | To be filled in by the bidder
16 | Property | All data will be property of the Erasmus MC. The supplier is prohibited to use the data for other purposes without concent of Erasmus MC. The data should be transferred to the Erasmus MC on demand (for example when license period expires). | KO | To be filled in by the bidder
17 | Availability of system after expiration of license | When the licence expires the supplier ensures his cooperation in a smooth migration to the new supplier and system. | KO | To be filled in by the bidder
| IT: Roles and authorisation |  | KO | To be filled in by the bidder
18 | Authorization management | The system has a flexible and centrally managed authorization functionality based on groups and roles. Based on these groups and roles the researcher has access or is denied access to certain functionalities and / or data. | KO | To be filled in by the bidder
19 | Multiple roles / groups per user | A user can be authorized for multiple groups and roles. It is also possible to provide an individual researcher custom authorizations. | KO | To be filled in by the bidder
| IT: Security and Privacy |  | KO | To be filled in by the bidder
20 | Documentation in writing | The documentation of the data models, software data communication protocols and parts of the processing of the system is fully stated in writing. | KO | To be filled in by the bidder
21 | ISO 27001 compliance | The tenderer is demonstrably in compliance with the general accepted standard for IT-security, ISO 27001. | KO | To be filled in by the bidder
22 | GDPR compliant audit trail | Audit trail is compliant with GDPR. | KO | To be filled in by the bidder
23 | Safe removal of hardware | Removal of hardware containing data is done in a safe and certified manner. | KO | To be filled in by the bidder
24 | Administrator access | Erasmus MC has the right to get insight in how access of system administrators of the supplier is organized | KO | To be filled in by the bidder
| IT: General |  | KO | To be filled in by the bidder
25 | Hybrid SaaS | The system is offered as Software as a Service (SaaS) with the ability to store data on Erasmus MC servers. | KO | To be filled in by the bidder
26 | Access limitation and retainment policies | Access limitation and retainment policies exist to limit chances of data loss from backups. | KO | To be filled in by the bidder
27 | Session management | The system applies session management which ensures that no data from different user sessions is mixed. | KO | To be filled in by the bidder
| IT: Maintenance and support |  | KO | To be filled in by the bidder
28 | Updates and maintenance | Updates and upgrades may never lead to degradation of the functionality of the system. The supplier guarantees that all demands written in this list of demands are upheld after updates and upgrades are installed. | KO | To be filled in by the bidder
29 | Updates and maintenance | Software follows lifecycle of the operating system. For example compatbile with Windows 11 and future updates. | KO | To be filled in by the bidder
30 | Skip update | The Erasmus MC can choose to skip at least one upgrade. This doesn't result in the diminishment of guarantees by the supplier. | KO | To be filled in by the bidder
31 | Version management | System has a clear tracking system of different versions of the system, with version numbers, date published and a user defined description. | KO | To be filled in by the bidder
32 | Service Desk | The supplier has an English speaking service desk. | KO | To be filled in by the bidder
| IT: Training & Help |  | KO | To be filled in by the bidder
33 | Manuals | Manuals and training are available to users and administrators in English | KO | To be filled in by the bidder
34 | Contextual help | Contextual help is available in the system | KO | To be filled in by the bidder
35 | Frequently asked questions | A list of frequently asked questions and their answers is available for reference | KO | To be filled in by the bidder
Maintenance, Response time and Repair
36 |  | Service includes; System design and documentation, Design Agreement System build, Software configuration, Method development Integration testing, Travel & accommodation. | KO | To be filled in by the bidder
37 |  | The supplier must provide a minimum 1-year warranty on all equipment delivered by the supplier. | KO | To be filled in by the bidder
38 |  | In case of equipment failure, the supplier must provide on-site repair services within 24 hours for critical components, with a maximum downtime of 48 hours for non-critical issues. | KO | To be filled in by the bidder
39 |  | Require the vendor to provide on-site repair services for issues that cannot be resolved remotely. | KO | To be filled in by the bidder
40 |  | Request the ability to troubleshoot issues remotely via secure connections, reducing downtime and minimizing the need for on-site visits. | KO | To be filled in by the bidder
41 |  | The vendor must provide detailed repair and service reports after each repair event, including the root cause, actions taken, and any parts replaced. | KO | To be filled in by the bidder
42 |  | Service includes Operator Training of at least 5 lab personel.(Vendor will provide system training to the current employees of iPS core team) | KO | To be filled in by the bidder
44 |  | The vendor is responsibile for damage to the automated platform during the repair or maintenance process. | KO | To be filled in by the bidder
Documentation and Manuals
45 |  | A complete technical manual will be provided in English. | KO | To be filled in by the bidder
Delivery Time, Milstones, Sanctions & Logistics
46 |  | The setup of the platform and the on-site training of at least 5 iPS core facility employees must be completed within 1 month after installation, including troubleshooting and emergency procedures. | KO | To be filled in by the bidder
47 |  | All end-users must be able to operate the equipment independently after the training by the bidder, without further assistance. | KO | To be filled in by the bidder
48 |  | If the project is delayed by more than 30 calendar days beyond the agreed delivery date, the issue will be escalated to senior management, and a formal meeting will be scheduled to discuss potential penalties, revised timelines, or further sanctions | KO | To be filled in by the bidder
49 |  | If the delivery is delayed by more than 90 calendar days from the final agreed-upon date without agreed upon reasons, the client reserves the right to terminate the contract without any further financial obligations, except for payment for work completed up to that point. | KO | To be filled in by the bidder
50 |  | Upon premature termination of the contract all acquired instruments paid by Erasmus MC should be shipped and delivered to the Erasmus MC. | KO | To be filled in by the bidder
51 |  | The client reserves the right to withhold 10% of the total contract value until final acceptance and successful system testing are completed, following the final Site Acceptance Test (SAT). |  | To be filled in by the bidder
52 |  | The vendor will be liable for any costs incurred due to delays in delivery, including extra operational costs or expedited shipping of required components, up to a maximum of 15% of the total contract value. |  | To be filled in by the bidder
53 |  | The shipment size of instruments packages should be specified. | KO | To be filled in by the bidder
54 |  | The supplier proactively informs the client about any changes in delivery times for equipment and required consumables. | KO | To be filled in by the bidder
55 |  | The supplier is required to provide a timely and accurate estimation of the delivery time. Failure to meet this requirement may result in sanctions. All communication regarding the delivery schedule must be transparent and free from exaggeration or misleading information. | KO | To be filled in by the bidder
Technical
56 |  | The supplier must provide a full validation package for Installation Qualification (IQ), Operational Qualification (OQ), and Performance Qualification (PQ) to be performed upon installation at the site. | KO | To be filled in by the bidder
57 |  | The system must include automated QC processes to detect and alert operators to any out-of-spec conditions in real-time. | KO | To be filled in by the bidder
58 |  | The system must demonstrate that the setup is able to maintain 84 cell lines in culture per day. | KO | To be filled in by the bidder
Inkoop
59 |  | There are no costs associated with submitting a tender, and the bidder derives no rights (now or in the future) from being permitted to submit a tender at this time. | KO | To be filled in by the bidder
60 |  | The bidder agrees to the entire procedure described and all conditions as outlined in the accompanying Offer Guidelines, the other tender documents, and the attachments. | KO | To be filled in by the bidder
61 |  | The bidder agrees to the agreements and provisions contained therein, which are included as attachments to the Offer Guidelines. | KO | To be filled in by the bidder
62 |  | The provisions of the agreements included in the attachments to the Offer Guidelines shall prevail. For situations not covered by the provisions of these (framework) agreements, the NFU General Purchasing Conditions, version January 2024 (Appendix 8), shall apply. These general purchasing conditions are included as Appendix. | KO | To be filled in by the bidder
63 |  | The bidder accepts the applicability of the NFU purchasing conditions, to the explicit exclusion of your own general terms and conditions. | KO | To be filled in by the bidder
64 |  | The final agreements resulting from this Offer Guidelines will be based on all provided information and documentation. The bidder guarantees that their organization possesses the necessary capacities, skills, and resources to (continue to) meet the Program of Requirements and Preferences as set out in this Offer Guidelines. | KO | To be filled in by the bidder
65 |  | Any substantial change in the bidder's situation that may affect their ability to meet these requirements must be promptly communicated to Erasmus MC. | KO | To be filled in by the bidder
66 |  | The bidder hereby declares that they have reviewed the (draft) purchase agreement and the associated conditions applicable to this assignment and, by submitting this PvE, confirms their agreement with the contract conditions outlined therein. | KO | To be filled in by the bidder
67 |  | The bidder hereby declares to be responsible for the accuracy and completeness of all provided information. The bidder is aware that, if it is later found that incorrect and/or incomplete information has been provided, they may be excluded from further participation, or any agreements already made may be canceled without legal consequences for Erasmus MC. | KO | To be filled in by the bidder
68 |  | The bidder has fully completed the price sheet as included in Appendix 10 and attached it to the submission. | KO | To be filled in by the bidder
69 |  | The prices offered in the price sheet are binding. | KO | To be filled in by the bidder
70 | Information sharing | Information concerning ErasmusMC is not presented to third parties without previous approval of the ErasmusMC. | KO | To be filled in by the bidder
Sustainability
71 |  | The supplier shall complete the following tabs: 1. Instruction, 2. Company, and 3. Sustainability. | KO | To be filled in by the bidder

## 2. imager wishes
Bijlage 2b. Programma van Wensen Imager
The sub-award criterion 'Program of Requirements' is worth a maximum of 273 points
For each wish, please indicate whether it applies by selecting "yes" or "no". For some wishes, additional information is requested. Based on your response, a score will be assigned. The maximum number of points that can be awarded per wish is shown under the ‘Score’ column.
REF | Categorie | Omschrijving | score (0-10) | Fill in with Yes or No; failure to fill in is not allowed. | Any explanation | Score toegekend door Erasmus MC
| Automated Imager |  | 133
1 |  | The imager has a user friendly interface. | 3 | To be filled in by the bidder | Overview of the Graphical User Interface.
2 |  | The imager is capable of digital image recognition measurements. | 10 | To be filled in by the bidder | Overview table and a example Proof of concept, with an example of several iPS wells from a 12- or 6-well plate (not a different cell type), showing the percentage of confluency. Define how the percentage is calculated.
3 |  | The imager is able to accurately calculate the true confluency of the wells. | 10 | To be filled in by the bidder | Please elaborate and add proof of concept (e.g., an example of several iPS wells defining the percentage of differentiation).
4 |  | The imager can recognize and annotate differentiated colonies by detecting iPS colony and spontaneous differentiation and provide an estimate of the differentiated portion. | 10 | To be filled in by the bidder | Please elaborate and add proof of concept
5 |  | The imager should measure and report the turbidity of the well due to cell death. | 10 | To be filled in by the bidder | Please elaborate and add proof of concept
6 |  | The imager can be used in differentiation assays/workflows and analysis of differentiated cells. | 10 | To be filled in by the bidder | Explanation/Example
7 |  | Upgrade options for enhanced lens quality for beter resolution and sensitivity is desirable. | 10 | To be filled in by the bidder | Please elaborate
9 |  | It is desirable that the imager is compatible with microscope slides and Cell culture dishes | 8 | To be filled in by the bidder | Please elaborate and add proof of concept
10 |  | It would be desirable for the imaging system to be capable of imaging both unlabeled and fluorescent cells at the highest magnifications, with automated magnification changes. | 8 | To be filled in by the bidder | Please elaborate and add proof of concept
11 |  | It would be desirable the imaging system supports U-shaped and V-shaped bottom plates. | 10 | To be filled in by the bidder | Please elaborate and add proof of concept
12 |  | It would be desirable to have a free software platform which does not require buying extra licenses and free software updates is preffered. | 10 | To be filled in by the bidder
13 |  | It would be desirable to have anti-Blur settings to enhance the image clarity in case of debris in supernatant | 9 | To be filled in by the bidder | Please elaborate and add proof of concept
14 |  | It would be desirable to export the raw data images for future analysis. | 5 | To be filled in by the bidder | Overview table and a example
15 |  | It would be desirable that the imaging software has customization options to increase throughput | 10 | To be filled in by the bidder | Overview table and a example
16 |  | Software solution for high resolution at low magnification | 10 | To be filled in by the bidder | Explanation/Example
| IT: Regulatory and compliancy |  | 9
17 |  | The platform should be compatible with existing laboratory information management systems (LIMS) and support communication via OPC, REST API, or other industry-standard protocols??? | 3 | To be filled in by the bidder
18 | Central Authentication Service | Authentication of users is done via central authentication services of Erasmus MC (Active Directory). Erasmus MC uses AD (Active Directory). For Windows Based systems authentication via LDAP must be possible. | 3 | To be filled in by the bidder
19 | Updates and maintenance | The supplier delivers needed updates and upgrades for the system. Updates and upgrades are included within development fee. | 3 | To be filled in by the bidder
| IT: Training & Help |  | 5
20 | E-learning | E-learning is available for the main functions of the system, it will also be kept up-to-date for new functions. | 5 | To be filled in by the bidder
| Maintenance Service and Repair |  | 101
21 |  | It is desirablet the warranty begins from the moment the imager is successfully integrated and operational at Eramus MC iPS core facility (SAT), rather than purchase date of imager. | 10 | To be filled in by the bidder
|  | We request that prospective suppliers submit a plan detailing how they will arrange repairs and support. The following topics will be evaluated. |  | To be filled in by the bidder | Please elaborate
22 |  | We wish for an extended warranty and access to 24/7 technical support to ensure uninterrupted system functionality. | 5 | To be filled in by the bidder | Please elaborate
23 |  | It would be desirable that the supplier is available 75% of the 24 hours for faults and issues related to the platform. | 6 | To be filled in by the bidder | Please elaborate
24 |  | In the event of a system failure, it is desirable that the vendor can respond within 2 hours and resolve critical failures within 24 hours. | 6 | To be filled in by the bidder | Please elaborate
25 |  | Remote instrument control possible. | 7 | To be filled in by the bidder | Please elaborate in which form: Wired Remote Control, Wireless Remote Control, Software-Based Control, App-Based & Mobile Control
26 |  | The operator receives remote warnings. | 7 | To be filled in by the bidder | Explanation/Example
27 |  | Are spare parts readily available (how quickly?), and are the parts directly available during a visit? | 8 | To be filled in by the bidder | Please elaborate
28 |  | How is the availability of suppliers guaranteed outside of office hours (including weekends)? | 7 | To be filled in by the bidder | Please elaborate
29 |  | How quickly can a technician be physically present? | 7 | To be filled in by the bidder | Please elaborate
30 |  | It would be desirable that the supplier offers a comprehensive, preventive maintenance, corrective maintenance, and software updates. | 8 | To be filled in by the bidder | Please elaborate
31 |  | Spare parts should be available for a minimum of 10 years after installation. | 10 | To be filled in by the bidder | Please elaborate
32 |  | For critical equipment, request the vendor to ensure that key components are available for immediate replacement. | 8 | To be filled in by the bidder | Please elaborate
33 |  | It would be desirable that the vendor provides support for equipment and software even after the end-of-life period, with clear terms for continued service or parts availability. | 10 | To be filled in by the bidder | Please elaborate
34 |  | It would be desirable if vendor can provide any upgrade or replacement conditions if equipment becomes obsolete or incompatible with future technological developments. | 2 | To be filled in by the bidder | Please elaborate
| Sustainability |  | 25
35 |  | The supplier shall complete the following tabs: 1. Instruction, 2. Company, and 3. Sustainability. | 25 | To be filled in on tab: 1. Instruction, 2. Company, and 3. Sustainability.
273

## 3. Liquid handerler Requirments
Programma van Eisen 3a. Liquid Handler
Failure to meet one or more (exclusion) knock-out (KO) criteria in this Requirements Document will result in your submission no longer being considered for the award of this contract. Meeting all knock-out criteria is regarded as a mandatory condition in this European Tender. Compliance with the Requirements Document does not earn any points.
The (question) statements in this Requirements Document are primarily intended to determine: 'Can the supplier meet the specified description?' Failure to meet one (or more) 'knock-out criteria' will result in rejection. Incorrect or incomplete filling in will lead to the immediate invalidation of the submission. The Requirements Document must be completed truthfully. PLEASE NOTE: The Excel sheet is not locked; the bidder should only fill in the fields that are highlighted in yellow and where it is indicated that something should be filled in.
|  |  | KO | Fill in with Yes or No; failure to fill in is not allowed. | Any explanation | Proof, documentation to be provided with the submission
1 |  | The liquid handerler gripper arm must have the capability to transfer different tube types (from 1ml 96 format cryotubes to 50ml falcon-like tubes) for present or future wordkflow | KO | To be filled in by the bidder
2 |  | The liquid handerler gripper arm must be able to automatically switch (within one run, without manual user intervention) between different types of gripper-fingers (for plates, tubes etc). | KO | To be filled in by the bidder
3 |  | The pipetting arm of the liquid handler should support a pipetting range from 0.5 µl to 1000 µl without requiring manual exchange of parts. | KO | To be filled in by the bidder
4 |  | The pipetting arm of the liquid handler should be capable of dispensing 0.5 µl in free-dispense mode with high consistency and reliability, ensuring precise and reproducible liquid handling for demanding applications. | KO | To be filled in by the bidder
5 |  | The pipetting arm of the liquid handerler must have a physical protection measure, to ensure no liquid can contaminate/damage and/or enter the pipetting channel. | KO | To be filled in by the bidder
6 |  | The liquidhandler should have minimally 6 MTP/SBS positions in depth. | KO | To be filled in by the bidder
7 |  | The liquidhandler should have the (upgrade) option to install a MultiChannel Pipetting arm capable of pipetting in 384 well plates, with 384 tips at the same time. And also have the capability the switch automatically (without manual intervention) to 96 format pipetting in 96 well plates with 96 tips at the same time. | KO | To be filled in by the bidder
8 |  | Liquidhandler must have the ability to integrate modules like plate-readers/-washers below the pipetting deck and able to reach with the integrated plate gripper arm. | KO | To be filled in by the bidder
9 |  | The liquid handerler should ensure high accuracy and reliability in liquid handling & highly accurate low and/or dead volume detection | KO | To be filled in by the bidder
10 |  | The platform has multiple tilt stations to support enhanced medium refreshement and passaging of plates. | KO | To be filled in by the bidder
11 |  | the platform should have one or more warm blocks, 37 degrees for incubation of plates |  | To be filled in by the bidder
12 |  | The supplier must ensure that the liquid handler has sufficient capacity to accommodate all the lab culture ware required for performing every expansion step outlined in the SOP. | KO | To be filled in by the bidder
13 |  | The supplier must ensure that the system has sufficient capacity to support all the cell expansion steps outlined in the SOP (Annex 12) for the minimal number of 84 iPS lines within 24 hours simultaneously. | KO | To be filled in by the bidder
ICT
IT: Regulatory and compliancy
14 | Good Laboratory Practice | Systems conforms to Good Laboratory Practice (GLP). | KO | To be filled in by the bidder
15 | Property | All data will be property of the Erasmus MC. The supplier is prohibited to use the data for other purposes without concent of Erasmus MC. The data should be transferred to the Erasmus MC on demand (for example when license period expires). | KO | To be filled in by the bidder
16 | Availability of system after expiration of license | When the licence expires the supplier ensures his cooperation in a smooth migration to the new supplier and system. | KO | To be filled in by the bidder
IT: Roles and authorisation
17 | Authorization management | The system has a flexible and centrally managed authorization functionality based on groups and roles. Based on these groups and roles the researcher has access or is denied access to certain functionalities and / or data. | KO | To be filled in by the bidder
18 | Multiple roles / groups per user | A user can be authorized for multiple groups and roles. It is also possible to provide an individual researcher custom authorizations. | KO | To be filled in by the bidder
IT: Security and Privacy
19 | Documentation in writing | The documentation of the data models, software data communication protocols and parts of the processing of the system is fully stated in writing. | KO | To be filled in by the bidder
21 | GDPR compliant audit trail | Audit trail is compliant with GDPR. | KO | To be filled in by the bidder
22 | Safe removal of hardware | Removal of hardware containing data is done in a safe and certified manner. | KO | To be filled in by the bidder
23 | Administrator access | Erasmus MC has the right to get insight in how access of system administrators of the supplier is organized | KO | To be filled in by the bidder
IT: General
24 | Hybrid SaaS | The system is offered as Software as a Service (SaaS) with the ability to store data on Erasmus MC servers. | KO | To be filled in by the bidder
25 | Access limitation and retainment policies | Access limitation and retainment policies exist to limit chances of data loss from backups. | KO | To be filled in by the bidder
26 | Session management | The system applies session management which ensures that no data from different user sessions is mixed. | KO | To be filled in by the bidder
IT: Maintenance and support
27 | Updates and maintenance | Updates and upgrades may never lead to degradation of the functionality of the system. The supplier guarantees that all demands written in this list of demands are upheld after updates and upgrades are installed. | KO | To be filled in by the bidder
28 | Updates and maintenance | Software follows lifecycle of the operating system. For example compatbile with Windows 11 and future updates. | KO | To be filled in by the bidder
29 | Skip update | The Erasmus MC can choose to skip at least one upgrade. This doesn't result in the diminishment of guarantees by the supplier. | KO | To be filled in by the bidder
30 | Version management | System has a clear tracking system of different versions of the system, with version numbers, date published and a user defined description. | KO | To be filled in by the bidder
31 | Service Desk | The supplier has an English speaking service desk. | KO | To be filled in by the bidder
IT: Training & Help
32 | Manuals | Manuals and training are available to users and administrators in English | KO | To be filled in by the bidder
33 | Contextual help | Contextual help is available in the system | KO | To be filled in by the bidder
34 | Frequently asked questions | A list of frequently asked questions and their answers is available for reference | KO | To be filled in by the bidder
Maintenance, Response time and Repair
35 |  | Service includes; System design and documentation, Design Agreement System build, Software configuration, Method development Integration testing, Travel & accommodation. | KO | To be filled in by the bidder
36 |  | The supplier must provide a minimum 1-year warranty on all equipment delivered by the supplier. | KO | To be filled in by the bidder
37 |  | In case of equipment failure, the supplier must provide on-site repair services within 24 hours for critical components, with a maximum downtime of 48 hours for non-critical issues. | KO | To be filled in by the bidder
38 |  | Require the vendor to provide on-site repair services for issues that cannot be resolved remotely. | KO | To be filled in by the bidder
39 |  | Request the ability to troubleshoot issues remotely via secure connections, reducing downtime and minimizing the need for on-site visits. | KO | To be filled in by the bidder
40 |  | The vendor must provide detailed repair and service reports after each repair event, including the root cause, actions taken, and any parts replaced. | KO | To be filled in by the bidder
41 |  | Service includes Operator Training of at least 5 lab personel.(Vendor will provide system training to the current employees of iPS core team) | KO | To be filled in by the bidder
42 |  | The vendor must provide training for in-house maintenance by staff to handle repairs and maintenance. | KO | To be filled in by the bidder
43 |  | The vendor is responsibile for damage to the automated platform during the repair or maintenance process. | KO | To be filled in by the bidder
Documentation and Manuals
44 |  | A complete technical manual will be provided in English. | KO | To be filled in by the bidder
Delivery Time, Milstones, Sanctions & Logistics
45 |  | The setup of the platform and the on-site training of at least 5 iPS core facility employees must be completed within 1 month after installation, including troubleshooting and emergency procedures. | KO | To be filled in by the bidder
46 |  | All end-users must be able to operate the equipment independently after the training by the bidder, without further assistance. | KO | To be filled in by the bidder
47 |  | If the project is delayed by more than 30 calendar days beyond the agreed delivery date, the issue will be escalated to senior management, and a formal meeting will be scheduled to discuss potential penalties, revised timelines, or further sanctions | KO | To be filled in by the bidder
48 |  | If the delivery is delayed by more than 90 calendar days from the final agreed-upon date without agreed upon reasons, the client reserves the right to terminate the contract without any further financial obligations, except for payment for work completed up to that point. | KO | To be filled in by the bidder
49 |  | Upon premature termination of the contract all acquired instruments paid by Erasmus MC should be shipped and delivered to the Erasmus MC. | KO | To be filled in by the bidder
50 |  | The client reserves the right to withhold 10% of the total contract value until final acceptance and successful system testing are completed, following the final Site Acceptance Test (SAT). | KO | To be filled in by the bidder
51 |  | The vendor will be liable for any costs incurred due to delays in delivery, including extra operational costs or expedited shipping of required components, up to a maximum of 15% of the total contract value. | KO | To be filled in by the bidder
52 |  | The shipment size of instruments packages should be specified. | KO | To be filled in by the bidder
53 |  | The supplier proactively informs the client about any changes in delivery times for equipment and required consumables. | KO | To be filled in by the bidder
54 |  | The supplier is required to provide a timely and accurate estimation of the delivery time. Failure to meet this requirement may result in sanctions. All communication regarding the delivery schedule must be transparent and free from exaggeration or misleading information. | KO | To be filled in by the bidder
Technical
55 |  | The supplier must provide a full validation package for Installation Qualification (IQ), Operational Qualification (OQ), and Performance Qualification (PQ) to be performed upon installation at the site. | KO | To be filled in by the bidder
56 |  | The system must include automated QC processes to detect and alert operators to any out-of-spec conditions in real-time. | KO | To be filled in by the bidder
57 |  | The system must demonstrate that the setup is able to maintain 84 cell lines in culture per day. | KO | To be filled in by the bidder
Inkoop
58 |  | There are no costs associated with submitting a tender, and the bidder derives no rights (now or in the future) from being permitted to submit a tender at this time. | KO | To be filled in by the bidder
59 |  | The bidder agrees to the entire procedure described and all conditions as outlined in the accompanying Offer Guidelines, the other tender documents, and the attachments. | KO | To be filled in by the bidder
60 |  | The bidder agrees to the agreements and provisions contained therein, which are included as attachments to the Offer Guidelines. | KO | To be filled in by the bidder
61 |  | The provisions of the agreements included in the attachments to the Offer Guidelines shall prevail. For situations not covered by the provisions of these (framework) agreements, the NFU General Purchasing Conditions, version January 2024 (Appendix 8), shall apply. These general purchasing conditions are included as Appendix. | KO | To be filled in by the bidder
62 |  | The bidder accepts the applicability of the NFU purchasing conditions, to the explicit exclusion of your own general terms and conditions. | KO | To be filled in by the bidder
63 |  | The final agreements resulting from this Offer Guidelines will be based on all provided information and documentation. The bidder guarantees that their organization possesses the necessary capacities, skills, and resources to (continue to) meet the Program of Requirements and Preferences as set out in this Offer Guidelines. | KO | To be filled in by the bidder
64 |  | Any substantial change in the bidder's situation that may affect their ability to meet these requirements must be promptly communicated to Erasmus MC. | KO | To be filled in by the bidder
65 |  | The bidder hereby declares that they have reviewed the (draft) purchase agreement and the associated conditions applicable to this assignment and, by submitting this PvE, confirms their agreement with the contract conditions outlined therein. | KO | To be filled in by the bidder
66 |  | The bidder hereby declares to be responsible for the accuracy and completeness of all provided information. The bidder is aware that, if it is later found that incorrect and/or incomplete information has been provided, they may be excluded from further participation, or any agreements already made may be canceled without legal consequences for Erasmus MC. | KO | To be filled in by the bidder
67 |  | The bidder has fully completed the price sheet as included in Appendix 11 and attached it to the submission. | KO | To be filled in by the bidder
68 |  | The prices offered in the price sheet are binding. | KO | To be filled in by the bidder
69 | Information sharing | Information concerning ErasmusMC is not presented to third parties without previous approval of the ErasmusMC. | KO | To be filled in by the bidder
Sustainability
70 |  | The supplier shall complete the following tabs: 1. Instruction, 2. Company, and 3. Sustainability. | KO | To be filled in by the bidder

## 3. Liquid Handerler Wishes
Bijlage 3b. Programma van Wensen Liquid Handler
The sub-award criterion 'Program of Requirements' is worth a maximum of 183 points
For each wish, please indicate whether it applies by selecting "yes" or "no". For some wishes, additional information is requested. Based on your response, a score will be assigned. The maximum number of points that can be awarded per wish is shown under the ‘Score’ column.
REF | Categorie | Omschrijving | score (0-10) | Fill in with Yes or No; failure to fill in is not allowed. | Any explanation | Score toegekend door Erasmus MC
| Liquid Handler |  | 80
1 |  | It would be desirable that the Liquid handler has the (upgrade) option for the following two liquidhandling techniques: Air Displacement technology & Liquid Displacement technology | 10 | To be filled in by the bidder | Please elaborate
2 |  | It would be desirable to have remote monitoring and control capabilities. | 10 | To be filled in by the bidder | Please elaborate in which form: Wired Remote Control, Wireless Remote Control, Software-Based Control, App-Based & Mobile Control
3 |  | The liquidhandler should have the (upgrade) option to accommodate 3 separate, independent pipetting and gripper arms, working in parallel. | 10 | To be filled in by the bidder
4 |  | Preferly has an integrated robotized thermal cycler, refrigerated microplate positions, Thermal Cycler and a microplate reader | 10 | To be filled in by the bidder
5 |  | The liquid hander should have fast pipetting and movements of arms | 10 | To be filled in by the bidder | Validation report or proof of concept
6 |  | The liquid handerler software should be user firendly launch any protocol clearly, quickly, easily, and intuitively | 10 | To be filled in by the bidder | Overview of the GUI
7 |  | The liquid handlers should allow easy configuration and reconfiguration of the workbench | 10 | To be filled in by the bidder | Explanation/Example higest score for the instrument fit to platform
8 |  | The liquid handler deck can be optimally reached by the robot arm. | 10 | To be filled in by the bidder | Please elaborate and add proof of concept
| IT: Regulatory and compliancy |  | 12
9 |  | The platform should be compatible with existing laboratory information management systems (LIMS) and support communication via OPC, REST API, or other industry-standard protocols??? | 3 | To be filled in by the bidder
10 | Central Authentication Service | Authentication of users is done via central authentication services of Erasmus MC (Active Directory). Erasmus MC uses AD (Active Directory). For Windows Based systems authentication via LDAP must be possible. | 3 | To be filled in by the bidder
11 | Continuity | The supplier guarantees to deliver service and maintenance for at least ten years, calculated from the start of agreement. | 3 | To be filled in by the bidder
12 | Updates and maintenance | The supplier delivers needed updates and upgrades for the system. Updates and upgrades are included within development fee. | 3 | To be filled in by the bidder
| IT: Training & Help |  | 5
13 | E-learning | E-learning is available for the main functions of the system, it will also be kept up-to-date for new functions. | 5 | To be filled in by the bidder
| Maintenance Service and Repair |  | 70
14 |  | It is desirablet the warranty begins from the moment the imager is successfully integrated and operational at Eramus MC iPS core facility (SAT), rather than purchase date of imager. | 10 | To be filled in by the bidder
|  | We request that prospective suppliers submit a plan detailing how they will arrange repairs and support. The following topics will be evaluated. |  | To be filled in by the bidder | Please elaborate
15 |  | We wish for an extended warranty and access to 24/7 technical support to ensure uninterrupted system functionality. | 3 | To be filled in by the bidder | Please elaborate
16 |  | It would be desirable that the supplier is available 75% of the 24 hours for faults and issues related to the platform. | 4 | To be filled in by the bidder | Please elaborate
17 |  | In the event of a system failure, it is desirable that the vendor can respond within 2 hours and resolve critical failures within 24 hours. | 4 | To be filled in by the bidder | Please elaborate
18 |  | Remote instrument control possible. | 5 | To be filled in by the bidder | Please elaborate
19 |  | The operator receives remote warnings. | 5 | To be filled in by the bidder | Please elaborate in which form: Wired Remote Control, Wireless Remote Control, Software-Based Control, App-Based & Mobile Control
20 |  | Are spare parts readily available (how quickly?), and are the parts directly available during a visit? | 5 | To be filled in by the bidder | Explanation/Example
21 |  | How is the availability of suppliers guaranteed outside of office hours (including weekends)? | 5 | To be filled in by the bidder | Please elaborate
22 |  | How quickly can a technician be physically present? | 5 | To be filled in by the bidder | Please elaborate
23 |  | It would be desirable that the supplier offers a comprehensive, preventive maintenance, corrective maintenance, and software updates. | 5 | To be filled in by the bidder | Please elaborate
24 |  | Spare parts should be available for a minimum of 10 years after installation. | 6 | To be filled in by the bidder | Please elaborate
25 |  | For critical equipment, request the vendor to ensure that key components are available for immediate replacement. | 5 | To be filled in by the bidder | Please elaborate
26 |  | It would be desirable that the vendor provides support for equipment and software even after the end-of-life period, with clear terms for continued service or parts availability. | 6 | To be filled in by the bidder | Please elaborate
27 |  | It would be desirable if vendor can provide any upgrade or replacement conditions if equipment becomes obsolete or incompatible with future technological developments. | 2 | To be filled in by the bidder | Please elaborate
| Sustainability |  | 16
28 |  | The supplier shall complete the following tabs: 1. Instruction, 2. Company, and 3. Sustainability. | 16 | To be filled in on tab: 1. Instruction, 2. Company, and 3. Sustainability.
183

## Additional apparatus lists
Apparatuur: Additionally, Additional apparatus lists includes automated equipment required for downstream analysis, which must also be integrated into the platform. However, the equipment in Additional apparatus lists is acquired independent of the tender, and suppliers are not required to bid on it. This lot serves only to inform the integrator about the requirements and preferences for downstream analysis. Suppliers are welcome to submit a separate quote for these instruments directly to the Erasmus MC iPS core facility.
a | The platform should be able to perform RT-PCR.
b | The platform should include a plate sealer.
c | The platform should include a spectrophotometer.
d | The platform should include a bulk reagent dispenser.
e | The platform should be able to perform RNA and DNA quantification.
f | The plateform should be able to perform RNA & DNA isolation.
g | The platform should have an automated centrifuge compatebel with SBS plates and tubes
h | The platform should have the a bility to centrifuge across different ranges and conditions
i | The centrifuge maintain controlled tempreture of 4 C to 37 C
j | Compateble rack system for centrifuge and -80 cryo rack

## 1. Instruction
1. Instruction
Erasmus MC - (sub-)category Offer/Tender-Ned - reference number Erasmus MC (tenderee) wants to reward (potential) contractors (tenderer) who demonstrably take steps in their business operations and chain to improve "impact" by means of this IMPACT Sustainability List of Wishes. If the Tenderer is not the technical and/or legal manufacturer, but a distributor/importer for example, then the Tenderer must provide the answers together with at least the "legal manufacturer".
General instruction
1 | Please complete this Excel file in its entirety and submit your offer as required in the "Quotation / Tender-guideline" ("Offerteleidraad").
2. | Fill in all yellow cells/questions. These are mandatory.
3. | PLEASE NOTE tab 3. Sustainability List of Wishes. If an explanation/answer is required for a wish, the following requirement apply: - the answer must apply to the solution that is sold/delivered: the burden of proof lies with the Tenderer. For example, if tab. 3 asks for climate reduction plans, it must be demonstrated that these are implemented for the requested solution (in the specifically named plants in tab. 2, for example) - answers must be short, logical/coherent and sufficiently detailed
Completion of this document
| Sheet name: | Input required:
| 1. Instruction | Not required. For information only. Please read carefully.
| 2. Company & production info | Answer all question in this sheet.
| 3. Sustainability list of wishes | Answer all question in this sheet.
Maximaal te

## 2. Company & production info
2. Company & production information
Instruction
1 | Do not change the structure, layout and/or formulas in this worksheet.
2. | Answer all questions / fill in all yellow cells. It is allowed to add extra rows and/or columns, but not to delete them.
3. | This concerns the (final) solution that will be used by Erasmus MC. This can be the individual products that are purchased or a total "system" that is deployed (system = composition of related equipment and/or components, etc.).
2.1 Company information
| 2.1.1 | Company (legal) name:
2.2 Production information
| 2.2.1 | Locations: | Where does production of the above mentioned (end) products take place?
|  |  | Production locations: | Plant 1 | Plant 2 | Plant 3 | Plant 4 | Plant 5 | Plant 6 | Plant 7
Name (owner production / technical manufacturing)
Plant name:
Street + number
Postal code
City
Country
| 2.2.2 | FTE: | Number of FTE (all locations combined)
| 2.2.3 | Operational: | Number of years operational:
| 2.2.4 | Goods/products/systems: | Which products are produced/assembled where?
|  |  | Products/production lines: | Plant 1 | Plant 2 | Plant 3 | Plant 4 | Plant 5 | Plant 6 | Plant 7
Name:
Name:
Name:
Name:
Name:
Name:
Name:
Name:
Name:
Maximaal te

## 3.Sustainability list of wishes
BIJLAGE xxx - IMPACT Sustainability Programme of wishes yyyy
Assessment (as part of the award model/criteria): |  | The sub-award criterion "IMPACT Program of Wishes" yields a maximum of the number of points according to the Quotation/Tender-Guideline. The maximum score for the wishes below is 1525 points.
Explanation: |  | For all the wishes below, the Tenderer is requested to indicate in column F whether they are met at the time of Quotation/Tender. Not filling them in is not permitted. An explanation must be provided as requested in column G. The answer must comply with the instructions on tab/sheet "1. Instruction" PLEASE NOTE: the Excel sheet is not locked: only the fields marked in yellow must be filled in.
Nr. | Element | Description | Score max. | Does Yes/No Meet? Fill in with only Yes or No | Explanation: Complete and submit as required.
1 | UN Sustainable Development Goals The sustainable development goals of the United Nations (UN), known worldwide as the Sustainable Development Goals (SDGs). See: https://sdgs.un.org/goals https://www.rijksoverheid.nl/onderwerpen/ontwikkelingshulp/internationale-afspraken-ontwikkelingshulp | Are the UN SDGs integrated into the business strategy? Are the sustainable development goals an integral part of the business strategy? At a minimum (qualitatively) required is an answer to the question. | No points |  | If yes, answer is required. If no, explain why not.
2 | ESG reporting Providing transparency (Disclosure) about the company's own climate/environment (Environment), social conditions (Social) and governance (Governance). See: https://finance.ec.europa.eu/capital-markets-union-and-financial-markets/company-reporting-and-auditing/company-reporting/corporate-sustainability-reporting_en#legislation Zie: https://www.weforum.org/impact/stakeholder-capitalism-reporting-metrics-davos2023/ https://www.pwc.com/sk/en/environmental-social-and-corporate-governance-esg/esg-reporting.html | Is there reporting on the ESG (Environmental, Social and Governance) impact? If so, is an ESG reporting framework/standard used (mandatory or not), and if so, which standard is used for what? At a minimum (qualitatively) required is an answer to the question. | No points |  | If yes, answer is required. If no, explain why not.
3 | Supplier(s) Code of conduct Application of code(s) of conduct towards suppliers to comply (at least) with rules, laws and guidelines in the field of climate, environment, employment conditions, fair and respectful treatment of employees and ethical practices. This as part of company-wide principles in addition to specific codes around privacy, ethics, anti-bribery for example. See: https://ecovadis.com/nl/glossary/supplier-code-conduct https://ecovadis.com/glossary/supplier-code-conduct/ https://nl.wikipedia.org/wiki/Gedragscode | Is there a company-wide Supplier Code of Conduct? Score award based on: - Score 100 points (maximum) if Supplier Code of Conduct is shared from Tenderer, "legal manufacturer(s)" and "technical manufacturer(s)". - Score 50 points if Supplier Code of Conduct is shared from Tenderer, "legal manufacturer" but not from "technical manufacturer(s)". - Score 25 points if Supplier Code of Conduct is shared from Tenderer. - Score 0 points if no codes of conduct are used and/or shared NB: Tenderer is the company who submits a Tender and becomes (potential) contractor of Erasmus MC. "Legal manufacturer", or manufacturer, is a natural or legal person who manufactures or completely revises a device or has a device designed, manufactured or completely revised and markets it under his name or brand. The "technical manufacturer" here refers solely to the natural or legal person who, where applicable, is contracted by the "legal manufacturer" for production/manufacturing (and does not market) or a so-called OEM. Score of 50 points if there is a "technical manufacturer(s)", to be indicated by the Tenderer, but these Supplier Code(s) of Conduct of those "technical manufacturer(s)" cannot be shared. If there is no "technical manufacturer(s)" then only 100 points (Tendering Party and "legal manufacturer" SCoC), 25 points (Tendering Party) or 0 points. Condition: score only awarded if the Code(s) of Conduct are up to date (demonstrably updated to changing laws/regulations), are used for the solution that is sold and it is clear what the scope of the code of conduct is (social, environmental, governance, financial, etc.). | 100 |  | Answer is mandatory. Code(s) of conduct must be shared (digitally) upon Quotation/Tender.
4 | Sustainability annual report A sustainability report (also known as a social report) is a report by a company on its performance in the area of ​​sustainability or corporate social responsibility: both performance in the area of ​​social ethics, environment and economics (People-Planet-Profit). See: https://nl.wikipedia.org/wiki/Duurzaamheidsverslag | Can the sustainability annual report be shared? Directly upon response (digital) or if Erasmus MC intends to award/contract. Note: required is at least not older than 2023 and, if the size of the company requires it, in accordance with/demonstrable according to Global Reporting Initiative (GRI) or SASB/IR Framework. | No points | Not applicable | Answer is mandatory. Annual report must at least be able to be shared if Erasmus MC intends to award/contract. At a first request to that effect from Erasmus MC, as soon as possible but no later than within 10 working days.
5 | Climate: This wish concerns the way in which the climate impact in the supply chain (sourcing, production, transport, and packaging) is dealt with: CO2 impact or equivalents (NOx etc.): Scope 1 ("internal" / direct emissions from activities owned/controlled) Scope 2 ("energy" / indirect emissions from purchasing and use of energy) emissions Scope 3 ("value chain/supply chain" / indirect emissions; so-called "up" and "downstream" of organization, chain) Specifically, this concerns (in accordance with the CSRD Directive) the 'Environmental (topical) standard ESRS E1': - climate change mitigation - adaptation - energy See: https://www.ser.nl/nl/thema/duurzaamheid/eu-duurzaamheidsrapportage https://www.ey.com/en_ie/insights/sustainability/csrd-how-to-manage-your-double-materiality-assessment https://www.mckinsey.com/featured-insights/mckinsey-explainers/what-are-scope-1-2-and-3-emissions https://www.epa.gov/climateleadership/scope-1-and-scope-2-inventory-guidance#:~:text=Scope%201%20emissions%20are%20direct,boilers%2C%20furnaces%2C%20vehicles). https://www.pwc.nl/en/topics/sustainability/esg/corporate-sustainability-reporting-directive.html | How do you deal with climate impact as a company? The following must be demonstrated: 1. Research has been conducted or is ongoing into the climate impact, in which it has been/is being investigated a. which climate impact is "material" (significant/very relevant) and b. whether the "double materiality" concept has been used 2. Company-wide plan in progress: research results are used to reduce the specific material climate impact for Scope 1, Scope 2 and/or Scope 3: at least a reduction plan with targets/KPIs formulated; scope of plan: entire company / all components over which there is operational control? 3. Reporting available according to Greenhouse Gas Protocol (GHG Protocol) https://ghgprotocol.org/ 4. Audit report can be shared (digitally) or upon intention to award, and is audited by external auditor, if required by EU legislation Score award based on: - Score 200 points if points 1., 2. and 3. are demonstrated, and, if required for Tenderer, point 4. - Score 75 points if only points 1. and 2. are demonstrated: improvement in implementation - Score 25 points if only point 1. is demonstrated: only in research phase - Score 0 points if points 1., 2. and 3. are not demonstrated | 200 |  | Answer is mandatory. Report/report mandatory to be shared with Quotation/Tender if available. Audit report must be able to be shared at least if Erasmus MC intends to award/contract. At a first request to that effect from Erasmus MC, as soon as possible but no later than within 10 working days.
6 | Climate - Science based target Science Based Targets (SBT) are science-based targets that provide a thorough analysis of your emissions and reduction plans. See: https://www.mvonederland.nl/wat-zijn-science-based-targets#:~:text=Science%20Based%20Targets%20(SBT)%20zijn,Klimaatakkoord%20van%20Parijs%20te%20halen. https://sciencebasedtargets.org/ | Are the climate/CO2 reduction plans approved by Science Based Targets (SBT Initiative)? Score award based on: - Score 100 points if: scope 1, 2 and 3 emissions approved and in line with the 1.5 °C "global warming" scenario - Score 50 points if: scope 1 and 2 emissions approved and in line with the 1.5 °C "global warming" scenario - Score 25 points if: scope 1, 2 and/or 3 approved, but not in line with the 1.5 °C "global warming" scenario (but in line with "well below 2°C" scenario or 2°C "global warming" scenario) - Score 0 points: no approval or not requested Only award if evidence of SbTi direct is attached. | 100 |  | Answer is mandatory.. Evidence must be included with Quotation/Tender.
7 | Climate - renewable energy Renewable energy for use in own production/all components over which there is operational control. In this wish this refers only to the electricity usage and sourcing thereof. | Renewable energy Is the own/purchased electricity 100% (demonstrably) based on "upstream" renewable energy (wind/solar/hydropower) or is there a plan/plans in progress to achieve 100% within a maximum of 5 (five) years? Score award based on: - Score 75 points if demonstrated that 100% has been achieved - Score 25 points if there are plans and in progress (demonstrable) - Score is 0 points if other than required. Scores only and only if demonstrable by means of evidence. | 75 |  | Answer is mandatory.. Evidence must be included with Quotation/Tender.
8 | Environment: This wish concerns the way in which environmental impact is dealt with: impact on land, water, biodiversity, and toxicity (ozone, aerosol, acidification, biosphere, eutrophication, fresh water and oxygen radicals). Specifically, this concerns the CSRD 'Environmental standards ESRS': - E2 (pollution): pollution of air, water and soil; pollution of living organisms and food sources; very hazardous substances (SVHC); microplastics - E3 (water & sea resources): use of water, withdrawal, discharge, discharge in sea, use and withdrawal sea resources) - E4 (biodiversity & ecosystems): direct causes for loss of biodiversity; impact on species; degree of impact on ecosystems and their condition; impact on and dependencies of ecosystem services See: https://www.metabolic.nl/software/link/blog/business-meets-nature-biodiversity-decoding-the-csrd/ | How do you as a company deal with environmental impact? The following must be demonstrated: 1. Research has been conducted or is ongoing into the climate impact, whereby research has been/is being conducted into a. which environmental impact is "material" (significant/very relevant) and b. or "double materiality" concept is used 2. Company-wide plan in execution: research results are used to reduce the specific material environmental impact, at least reduction plan with targets/KPIs formulated, demonstrated that there is board-level commitment and what the scope of the plan is: entire company / all parts where there is operational control 3. Report / research / report available available for Erasmus MC 4. Audit report can be shared (digitally) or upon intention to award, and is audited by external auditor, if EU legislation requires Score award based on: - Score 200 points if points 1., 2. and 3. are demonstrated, and, if required for Tenderer point 4. - Score 75 points if only points 1. and 2. are demonstrated: improvement in execution - Score 25 points if only point 1. is demonstrated: only in research phase - Score 0 points if points 1., 2. and 3. are not demonstrated | 200 |  | Answer is mandatory. Report/report mandatory to be shared with Quotation/Tender, if available. Audit report must be able to be shared at least if Erasmus MC intends to award/contract. At a first request to that effect from Erasmus MC, as soon as possible but no later than within 10 working days.
9 | Value retention (material efficiency): Retention of value of (technical) raw materials, for example through reuse, refurbishment or remanufacturing. More specifically, the impact according to CSRD 'environmental topic standard ESRS': E5 (resource use and circular economy): inflow/inflow of resources; outflow/outflow related to products/services; waste. See: https://www.ellenmacarthurfoundation.org/completing-the-picture | Dealing with value retention: 1. Research has been carried out or is ongoing into material efficiency / value retention, in which it has been/is being investigated a. which material impact is "material" (significant/very relevant) and b. whether the "double materiality" concept has been used 2. Company-wide plan in progress: research results are used to reduce the specific impact, at least a company-wide initiative to include material efficiency / value retention in the development of new solutions, for example the application of eco-design and/or life cycle analysis (results) in R&D for the solutions that Erasmus MC intends to purchase Score allocation based on: - Score 200 points if points 1., 2. have been demonstrated - Score 75 points if only point 1. has been demonstrated: only in research phase - Score 0 points if points 1., and 2. have not been demonstrated | 200 |  | Answer is mandatory. Evidence must be included with Quotation/Tender.
10 | Impact product level measurement: Environmental-LCA (Life Cycle Analysis) and/or Social-LCA (UNEP) | LCA available of offered solution A Lifecycle Assessment (LCA) has been performed and independently validated, of the entire "Environmental" impact E-LCA and/or social impact Social-LCA is available. Score award based on: Score 75 points if: - Specifically for the products that are in-scope: LCA report is available and can be shared prior to any intention to award, LCA must demonstrably comply with ISO standards: ISO 14040: 2006 (principles/framework), ISO 14044: 2006 (requirements). - If only the climate impact at product level, this is also accepted (or required according to ISO 14067: 2018). - LCA has been validated by an independent institution (such as SGS, Bureau Veritas, Det Norske Veritas or Control Union World Group). Sharing data/data model is not required. | 75 |  | Answer is mandatory. Evidence may be included with the Quotation/Tender. Evidence must be able to be shared if Erasmus MC intends to award/contract. Upon a first request to that effect from Erasmus MC, as soon as possible but no later than within 10 working days.
11 | Socio-economic: This wish concerns the way in which socio-economic circumstances and relationships are dealt with in the entire supply chain. Specifically the impact according to 'CRSD - Social standards ESRS': S1 (own employees) S2 (employees in value chain) S3 (communities) S4 (consumers and/or end users) G1 (business conduct) See: https://www.ohchr.org/en/what-are-human-rights/international-bill-human-rights https://www.ilo.org/declaration//lang--en/index.htm https://unglobalcompact.org/what-is-gc/mission/principles | How do you as a company deal with the social impact? The following must be demonstrated: 1. Research has been conducted or is ongoing into the social impact, in which it has been/is being investigated a. which social impact is "material" (significant/very relevant) and b. whether the "double materiality" concept has been used. 2. Systematics/guidelines and specific plan in progress to address the material impact? Ensure that production methods, labor standards and anti-corruption practices comply with UN guidelines, more specifically OECD guidelines (Guidelines for Responsible Business Conduct), the ILO conventions / recommendations (The International Labour Organization Conventions and Recommendations (ILO)), the so-called 'UN Guiding Principles (UNGP) for Business and Human Rights, and the "United Nation Conventions on the Rights of the Child (UNCRC)' 3. Answer the following specific questions: is there a "Labor standard" statement, "Modern Slavery" statement, "Anti-corruption" statement? If yes, mandatory to share evidence. If no, mandatory answer and explain why not. 4. Report / research / report available for Erasmus MC 5. Audit report can be shared (digitally) or upon intention to award, and is audited by external auditor, if EU legislation requires Score award as follows: - Score 200 points as points 1. to 4. have been demonstrated, and, if required for the Tenderer by law, point 5. - Score 75 points if only points 1. and 2. have been demonstrated: improvement in implementation - Score 25 points if only point 1. has been demonstrated: only in research phase - Score 0 points if none of the points have been demonstrated | 200 |  | Answer is mandatory. Report/report mandatory to be shared with Quotation/Tender if available. Audit report must be able to be shared at least if Erasmus MC intends to award/contract. At a first request to that effect from Erasmus MC, as soon as possible but no later than within 10 working days.
12 | Supplier chain responsibility: As a company, dealing with sustainability towards its own chain/suppliers. | Suppliers' own 'governance' in terms of sustainability The following must be demonstrated: 1. Dealing with impact (on climate, environment, value retention (materials), socio-economic level) is integrated or will be integrated in the short term (<1 year) in at least: selection of suppliers, development of suppliers and agreements. 2. Specific answer to the following questions: a. are - in addition to the supply chain risks - the sustainability risks mapped out in the supplier network (factories, classification/key suppliers, requirements)?; b. are sustainability risks that go beyond own suppliers (Tier 1) discovered, and c. how/what is done to resolve abuses and what has been used to resolve abuses. 3. Audit systematics: which audits are performed, how often? If possible, share the most recent audit report. 4. 1 concrete example to demonstrate that non-compliance with sustainability guidelines leads to tightening/improvement of agreements and/or can lead to termination of supplier relationships. Score allocation as follows: - Score 250 points if points 1. to 4. are demonstrated - Score 150 points if only points 1. to 3 are demonstrated - Score 75 points if only points 1. and 2. are demonstrated - Score 25 points if only point 1. is demonstrated - Score 0 points if none of the points are demonstrated | 250 |  | Answer is mandatory. Report/report mandatory to be shared with Quotation/Tender, if available. Audit report must be able to be shared at least if Erasmus MC intends to award/contract. At a first request to that effect from Erasmus MC, as soon as possible but no later than within 10 working days.
13 | Packaging - recyclability of: - product packaging material - transport packaging - descriptions or instructions for use (DFU/IFU) | Degree to which packaging is recyclable Demonstrate the following: Packaging made from recycled materials and packaging fully recyclable: documentation/evidence can be provided directly upon response or upon request of Erasmus MC: weight of materials, % percentage of recyclable content; whether packaging is recyclable. Score award as follows: - Score 25 points if demonstrated - Score 0 points if not demonstrated | 25 |  | Answer is mandatory. Evidence to be included with Quotation/Tender or at a first request to that effect from Erasmus MC, as soon as possible but no later than within 10 working days.
14 | Hazardous Substances List (RSL) Substances of Very High Concern (SVHC) | Management and Phasing Out of Hazardous Chemicals Is there a 'chemicals management' strategy that, in addition to complying with strict global legislation and regulations, focuses on the voluntary phasing out of the supply chain of specific hazardous substances that cause serious and/or irreversible damage to the environment and/or human health but are not yet regulated? Demonstrate the following: 1. Documentation / evidence of a list of hazardous substances (legal/industry specific/voluntary) that suppliers/suppliers are contractually required to phase out and/or not use, e.g. "Regulated Substances List" in relation to the products that are in-scope. 2. Documentation / evidence of application and concentration of the list of very hazardous substances, specifically the "Substances of Very High Concern (SVHC)" (according to art. 57 EC REACH Regulation) in relation to the products that are in-scope. Scoring is as follows: - Score 50 points if points 1. and 2. are proven - Score 25 points if only 1 of the 2 points is proven - Score 0 points if both cannot be proven. | 50 |  | Answer is mandatory. Evidence must be included with Quotation/Tender.
15 | Socio-economic/conflict minerals: Conflict minerals: insight, transparency and taking responsibility (together) | Dealing with conflict minerals Demonstrate the following: 1. Policy on responsible sourcing of "conflict minerals" 2. Report / research / report available including evidence that "due diligence" meets (or exceeds) the OECD guidance on minerals (https://www.oecd.org/daf/inv/mne/mining.htm), available for Erasmus MC 3. Audit report can be shared (digitally) or upon intention to award/contract, and is audited by external auditor, if EU legislation requires Score award as follows: - Score 50 points if points 1. to 3. are demonstrated - Score 25 points if point 1. is demonstrated - Score 0 points if both cannot be demonstrated. | 50 |  | Answer is mandatory. Evidence must be included with Quotation/Tender.
1525
Company name:
Signature/date:
Maximaal te