## Requirements and Preferences
Annex A04, Compliance with Requirements - Answers to Preferences Applicable to the European open tender for a HPHT Triaxial Apparatus, Ref. WS3011436760, dated 10 March 2026
The Tenderer demonstrates the ability to meet the requirements by answering Yes/No to the questions, which must be structured in accordance with the requirements given in Chapter 8. Tenders that are incomplete and/or fail to meet all the formal requirements will be declared invalid by TNO and discarded. Missing answers, documents, or data will result in exclusion if there is no rectifiable omission. TNO assesses the Tenders against the Programme of Requirements and Preferences. The Programme of Requirements and Preferences includes Minimum Requirements, Performance Requirements and Preferences. The Minimum Requirements must be met on penalty of invalidity at the time of submission of the Tender, the Performance Requirements only at the time of performance of the Contract and the Preferences must be worked out by Tenderers as part of the award of the Contract.
Requirement | Description | Confirmation of fullfillment of (minimum) requirement Yes/No | Reference in Tenderer's documentation where Tenderer states his answer to the Preferences | Optional: Remarks
8.1 | Requirements
Requirement 8.1.1 | The maximum confining pressure capability shall reach at least 140 MPa with a pressure measurement resolution ≤ 1kPa. | Fill in here: Yes/No
Requirement 8.1.2 | The maximum temperature capability during testing shall reach at least 200 °C. | Yes/No
Requirement 8.1.3 | The maximum axial load shall reach at least 1000 kN. | Yes/No
Requirement 8.1.4 | Confining pressure system shall operate at max P–T without refilling: - confining oil tank with a large enough volume to accommodate the maximum variations of pressure and temperature or, - Dual‑pump system ensuring continuous flow. | etc
Requirement 8.1.5 | The HPHT Triaxial Apparatus must accommodate 1’’ or 25.4mm diameter samples and have to option to upgrade to 1.5’’ or 38.1mm diameter samples.
Requirement 8.1.6 | The HPHT Triaxial Apparatus axial deformation displacement rate shall be comprise between 0.1 µm/s and 0.1 mm/s. The HPHT Triaxial Apparatus shall have a range of axial deformation displacement rate that fits or encompasses this range.
Requirement 8.1.7 | The HPHT Triaxial Apparatus shall include Internal axial and radial displacement sensors fitting for samples of 25.4 mm diameter and 50.8 mm length. The displacement sensors shall be able to measure at least 10% axial and 5% radial deformation for the aforementioned sample dimension. Axial and radial displacement measurement resolution shall be ≤ 1 μm.
Requirement 8.1.8 | The HPHT Triaxial Apparatus shall include at least 1 thermocouple able to measure temperatures at the sample location, and at maximum operating pressure and temperature and a temperature measurement resolution ≤ 0.5 °C.
Requirement 8.1.9 | The electronic signal feedthrough system capacity shall allow simultaneous connections of: - 2 thermocouples. - Axial and Radial displacement sensors. - Ultrasonic measurement system. - Up to 6 Acoustic emission sensors (individual or multi-channel accepted). - Up to 4 strain gages.
Requirement 8.1.10 | The HPHT Triaxial Apparatus shall have fittings access for the confining pressure oil and fittings to accommodate a sample pore pressure system with inlet and outlet draining.
Requirement 8.1.11 | The HPHT Triaxial Apparatus shall have a control and data acquisition software capable of synchronized data acquisition across channels and data recording and compatible with TNO computers.
Requirement 8.1.12 | The Tenderer shall cover the delivery, transport and installation of the HPHT Triaxial Apparatus in the iM4RockLab located on TNO location Princetonlaan Utrecht, The Netherlands.
Requirement 8.1.13 | The Tenderer shall provide operator training for at least 2 persons on location following the HPHT Triaxial Tester installation.
Requirement 8.1.14 | The Tenderer shall provide complete user manuals, maintenance instructions and calibration guidelines.
Requirement 8.1.15 | The Tenderer shall provide documentation regarding: - Frame stiffness. - Calibrations for displacement and pressure sensors, and load cells. - Apparatus footprint, dimension and weight. - Electrical needs.
Requirement 8.1.16 | The tenderer shall provide a warranty of minimum twenty-four (24) months on all delivered Goods, covering both parts and labour, excluding only components subject to normal wear. This warranty shall ensure that the Goods are free from defects and fully suitable for their intended use for the duration of the warranty period.
Requirement 8.1.17 | The Tenderer shall ensure that the total cost of the proposed HPHT Triaxial Apparatus, including all necessary components, installation activities and training, does not exceed a total price of €540,000.
Requirement 8.1.18 | The offered HPHT Triaxial Apparatus and all relevant parts shall be CE-certified.
Requirement 8.1.19 | The offered HPHT Triaxial Apparatus must be brand new (recently produced) and free of signs of use and not previously used (not even as a demonstration or rental model) and provided with a first use certificate.
Requirement 8.1.20 | A full Site Acceptance Test (SAT) on-site at TNO location Princetonlaan, Utrecht, shall be performed by the tenderer in cooperation with TNO. As part of the tender bid, we require the Tenderer to submit a proposal for the SAT plan. TNO states that the minimum requirements for checkpoints that must be incorporated in the SAT plan are: - The HPHT Triaxial Apparatus will be checked for damages; - Any open issues from the Factory Acceptance Test (FAT) will be checked; - The proper functioning check of all functions of the HPHT Triaxial Apparatus; - The presence of the required documentation will be checked. The details of the FAT shall be defined and agreed between TNO and the final Tenderer.
8.6 | Preferences
Preference 8.2.1 a to g | To prevent complications arising from the use of third‑party equipment—such as issues with data acquisition synchronization—it is preferable for TNO that the HPHT Triaxial Apparatus be supplied with an integrated high‑pressure pump that can serve as the pore fluid pressure system and can be fully controlled from, and have its data recorded by, the same software platform referenced in Requirement 8.1.10. No pore fluid pressure system: 0 points; Pore fluid pressure system specifications as mentioned in the preferences below, points as follows:
Preference 8.2.1 a | Preference 8.2.1 a: Maximum pressure: ≥ 100 MPa: 50 points 70 to 99.9 MPa: 25 points < 70 MPa: 0 points
Preference 8.2.1 b | Preference 8.2.1 b: Gas tight: yes: 30 points no: 0 points
Preference 8.2.1 c | Preference 8.2.1 c: Corrosion resistant: yes: 30 points no: 0 points
Preference 8.2.1 d | Preference 8.2.1 d: Minimum Flow-rate: < 1 µm/min: 60 points ≥ 1 µm/min: 0 points
Preference 8.2.1 e | Preference 8.2.1 e: Maximum Flow rate ≥ 100 mL/min: 60 points 10-99.9 mL/min: 30 points < 10 mL/min: 0 points
Preference 8.2.1 f | Preference 8.2.1 f: Cylinder volume ≥ 200 mL: 60 points 100-199.9 mL: 40 points 50-99.9 mL: 20 points < 50 mL: 0 points
Preference 8.2.1 g | Preference 8.2.1 g: Pulse test circuit for permeability measurements: Yes: 60 points No: 0 points
Preference 8.2.2 | For flexibility regarding the type of samples to be tested, TNO would prefer if the HPHT Triaxial Apparatus has the capability to test samples of dimension : 1.5’’ or 38.1 mm diameter and a length to diameter ratio of at least 2:1: Yes: 100 points No: 0 points
Preference 8.2.3 | The HTHP Triaxial apparatus would gain in characterization capability from an in-situ ultrasonic velocity measurement system (Vp and Vs) with its data acquisition software: Yes: 50 points No: 0 points
Preference 8.2.4 | The HTHP Triaxial apparatus would gain in characterization capability from an in-situ acoustic emission detection system with its data acquisition software: 6 piezometer sensors: 50 points 4 piezometer sensors: 25 points 2 piezometer sensors: 10 points No acoustic emission detection system: 0 points
Preference 8.2.5 | For sake of efficiency and in order to reduce experiment time duration, TNO would prefer a pressure cell designed with an active cooling system: Active cooling system: 50 points. No active cooling system: 0 points
Preference 8.2.6 | TNO attaches great value to short delivery time: Delivery between 1 up to and incl. 34 weeks after confirmation of contract: 50 points Delivery between 35 up to and incl. 43 weeks after confirmation of contract: 30 points Delivery between 44 up to and incl. 52 weeks after confirmation of contract: 15 points Delivery beyond 52 weeks after confirmation of contract: 0 points.
Preference 8.2.7 | TNO would prefer if the warranty conditions detailed in requirement 8.1.15 is extended to: > 60 months: 50 points > 36 to 60 months: 30 points > 24 to 36 months: 15 points

## Preferences (hidden)
Annex A04, Prescribed format for answering questions/fullfilling preferences Guidelines for the European Public procurement tender for: one cell DLIT and one module DLIT, Ref. WS2616908858, dated 00 September 2025
The Tenderer demonstrates the ability to meet the preferences by means of his answers to the various questions/preferences, which must be structured in accordance with the instructions given in Chapter 8.
Preference | Description | Confirmation of fullfillment of preference Yes/No | Reference in Tenderer's documentation where Tenderer indicates in which way he can fulfil the preference | Remarks
8.2.4
8.4.6
8.4.7
8.5.2
8.5.3

## Blad2 (hidden)
Yes
No