## Vragen_en_antwoorden_576297_all
Ref. Nr. | Onderwerp | Vraag | Antwoord
1 | Tender documetnts language | Is it possible to get all the tender documents in English? Only the Dutch version is available on TenderNed. | All documents that are provided in Dutch will not be provided in English. However, it is allowed to ask questions and/or submit the tender in English.
2 | QUESTIONS #1 | 1. With reference to Bijlage 5 – Prijzenblad, could you please clarify whether the WGS singleton and trio analyses are expected to be performed as full WGS analyses (i.e. VCFs with more than 300,000 annotations), or whether a virtual gene panel will be applied to limit the number of annotations? In our software, pricing per sample is based on three annotation tiers: fewer than 50,000 annotations, 50,000–300,000 annotations, and more than 300,000 annotations. The gene panels you mentioned would likely fall into the lowest tier. Could you advise how we should reflect these annotation levels in our pricing proposal? 2. Please confirm whether quotations may be submitted in USD? By request of the LUMC, your account has been setup in USD, quoting in EUR would convert the account to EUR invoicing for all purchase orders in the future, including those unrelated to this tender. The Account can only be changed to a different currency once." | 1. Our WGS VCF files contain 5.000.000 annotations. No pre filtering. In case of captured NGS the VCF files contains ~200.000 annotations. 2. To ensure a level playing field, all tender prices must be in EUR.
3 | Question regarding Article 7 – Warranty | With reference to Article 7 (Warranty and Quality) of the draft agreement, could you please clarify in more detail what the warranty is expected to cover in practice? In particular, could you confirm whether the warranty is intended to cover only software defects and non-conformities affecting the agreed use, or whether it is also expected to cover other aspects such as performance levels, regulatory compliance over time, or changes in external requirements? | The warranty, as described in Article 7 of the draft agreement, relates to software defects and non-conformities that affect the agreed use. The warranty also covers the performance of the software, compliance with statutory requirements throughout the term of the agreement, and adaptation to changes in external requirements, as specified in the procurement documents.
4 | Price file | 1. In our offer, we are not permitted to include unrealistic pricing, such as listing an item at €0. If an item included in the price file is not available to us as a standalone item to quote, how should this be handled? Is it acceptable to leave the line blank, or should it be listed with a value of €0? | Including a price of €0.00 for an item is not permitted. If an item in the price list cannot be offered as a separate product, it should not be included in the tender submission. Leave the line blank; the contracting authority will take this into account.
5 | Bijlage 3 – Kwaliteitsblad Programma van Eisen | With reference to Bijlage 3 – Kwaliteitsblad Programma van Eisen, section 2.7 specifies a requirement of more than 75 concurrent users, while section 8.1 specifies a requirement of at least 100 users. As both are stated as mandatory requirements, could you please clarify the distinction between these two specifications? | The mininum requirement is 75 concurrent users, but scalable up to 100 users.
6 | Upload performance dependency (PvE 2.4) | Would it be acceptable for requirement PvE 2.4 to be limited to the processing of VCF and BAM files once uploaded, given that upload performance itself is highly dependent on the end user’s internet connection? | Processing time of >75GB should be <60min. Gedeelte upload weglaten. Requirement is adhusted..xxxxxxxxx
7 | Multi-user access vs. edit rights (PvE 6.7 / PvE 6.8 / PvW 6.8) | Could you please clarify the intent of requirements PvE 6.7 / 6.8 and PvW 6.8? Specifically, how should multi-user access to an analysis record (PvE 6.7) be interpreted in relation to the requirement that only a single user can edit an analysis record at any given time (PvE 6.8)? | To prevent adjustments are lost due to simultaneous editing of a single record.
8 | ONT and PacBio data format (PvE 1.1) | Can you confirm whether data generated from Oxford Nanopore Technologies (ONT) and PacBio platforms will be provided in VCF format? | Yes, that is correct.
9 | Repeat expansion disorders (PvE 1.2) | Could you please confirm whether LUMC has a predefined list of genes and/or specific repeat loci that should be included for the interpretation of repeat expansion disorders? In addition, can you clarify whether repeat expansion analysis is expected to be performed on Illumina short-read data, long-read sequencing data (e.g., ONT or PacBio), or both? | The list is in the DRAGEN VCF output and will be performed on all types of data
10 | Questions regarding Bijlage 3 | 1. With reference to document Bijlage 3 Kwaliteitsblad-Programma van Eisen, point 1.17 - The platform must support methylation analysis: Which types of variants and visualization would be expected to be assessed for methylation related data for clinical interpretation? 2. With reference to document Bijlage 3 Kwaliteitsblad-Programma van Eisen, point 1.31 - The platform must support datasharing with National and International cooperative efforts (VKGL, SolveRD, etc.): What is the expected file format for sharing information with and from SolveRD? 3. With reference to document Bijlage 3 Kwaliteitsblad-Programma van Eisen, point 6.8 - The platform must ensure that only one user can edit an analysis record at the same time.: Is the intention of this requirement to prevent conflicting information being added by multiple users working on a case simultaneously or to limit case access to a single user at a time? | 1. e.g. allelic specific, UPD's (imprinting dissorders), Episign 2. CSV and or VCF 3. Yes, that is correct.
11 | Question regarding 1) software agreement 2) DPA 3) ARBIT 2022 | Would the Contracting Authority be amenable to negotiating certain provisions within the proposed Terms and Conditions? Following a comprehensive review of the tender documentation by our Legal and Privacy Teams, we kindly request confirmation on whether the Contracting Authority would entertain some proposed amendments to the Software Agreement, applicable clauses of ARBIT 2022, and the Data Processing Agreement (DPA). The anticipated modifications primarily concern, but are not limited to, provisions regarding the limitation of liability and specific timeframes for data breach notifications. | The contractual documents included in the tender, including the Software Agreement, ARBIT 2022 provisions, and the Data Processing Agreement, form an integral part of the Tender and are generally not subject to negotiation. Requests for clarification/modification may be submitted and, where applicable, will be addressed in a formal notice applicable to all tenderers.
12 | Bijlage 3: Kwaliteitsblad-Programma van Eisen | (PvE), Req 5.2 — "Transfer MVL from current software" Given the transition from the current platform to a new system, please clarify expected data migration support: a) In what format and volume are existing classified variants (MVL) available? ACMG classifications only, or also internal annotations, curation notes, and audit trails? b) Should historical analysis results (previous cases) be preserved and accessible in the new platform, or is archival in the current system sufficient? c) What is the expected transition timeline, and does LUMC foresee a parallel running period? If so, what is the expected duration? | a) De bestaande geclassificeerde varianten (MVL) zijn beschikbaar in een CSV-export. Deze export omvat ACMG-classificaties, interne annotaties, curatie-aantekeningen en audit trails. Het volume betreft ongeveer 15.000 varianten. b) Historische analyse resultaten (voorgaande cases) moeten bewaard en toegankelijk zijn in het nieuwe platform. Archivering in het huidige systeem is niet voldoende. c) De verwachte transitieperiode is 6 maanden. Het LUMC voorziet een parallelle draaiperiode van 3 maanden.
13 | Bijlage 3: Kwaliteitsblad-Programma van Eisen | (PvE), Req 5.2 — "Transfer MVL from current software" For MVL migration from the current Franklin platform: a) What format are classified variants stored in (VCF, CSV, proprietary)? b) Approximate number of classified variants in current MVL? c) Beyond classification, are there free-text fields, internal annotations, or metadata to transfer? d) Will the data transfer be limited to reviewed, classified variants (i.e., a curated subset with interpretations), or is it expected to include all called variants regardless of classification status, and does this extend to raw or uninterpreted historical patient data? | a) Most likely CSV format b) 80.0000 variants (SNV), 450.000 variants (SV) c) Yes, free-text fields, internal annotations and metadata. d) Reviewed and classified variants
14 | CE-IVDR | LUMC Clinical Genetics operates within an ISO 15189-accredited diagnostic environment. For use in accredited diagnostic workflows: a) Does LUMC require the platform to carry a CE-IVD or IVDR marking for use in ISO 15189-accredited clinical diagnostic reporting? b) Is a Research Use Only (RUO) designation acceptable for use in accredited diagnostic workflows, or is CE-IVD status the minimum? | a) No, that is not required, but preferred. b) RUO is also acceptable, we perform a validation of the software platform in our ISO15189 accredited lab.
15 | GDPR requirements | Regarding GDPR requirements for processing special category data (genetic data, health data — Art. 9 GDPR): a) Must all patient data be processed and stored exclusively within the EEA? b) Does LUMC require that the platform vendor's corporate headquarters and primary development operations be located within the EEA? Or is EEA-based data hosting sufficient regardless of vendor domicile? c) Is processing outside the EEA (e.g., Israel, United States, Switzerland) permitted with safeguards such as SCCs? | a) Yes, that is correct. b) Yes, the platform vendor's corporate headquarters and primary development operations must be located within the EEA c) No, this is not permitted.
16 | Bijlage 3: Kwaliteitsblad-Programma van Eisen | Req 1.1 Regarding QC files a) What formats of QC files are expected as input (e.g., MultiQC, FastQC, Picard metrics)? b) What level of QC support is required — visualization of metrics within the platform, automated pass/fail alerting based on configurable thresholds, or both? | a) DRAGEN QC output. Metrics JSON-files b) visualization of metrics within the platform without automated pass/fail alerting
17 | Bijlage 3: Kwaliteitsblad-Programma van Eisen | Tab Programma van Wensen, the dropdown criteria to fill column G “uw antwoord” are not matching with the associated question most of the time. How should we handle this? | The dropdown criteria in column G of the Program of Requirements have been corrected. You can now select the correct answers corresponding to the relevant questions.
18 | Bijlage 3: Kwaliteitsblad-Programma van Eisen | Session timeout: configurability and re-authentication behaviour (Req 6.9) Req 6.9 requires customisable session timeout limits. Please clarify what level of configurability is needed for LUMC's use case: a) Is a configurable session timeout at the organisation level sufficient, or does LUMC require control at a finer granularity such as per user group? b) What minimum and maximum timeout durations does LUMC's IT security policy require? c) Should timeout be triggered by user inactivity only, or also by a maximum session duration regardless of activity? d) After session timeout, does LUMC's access policy require mandatory re-authentication (username/password or MFA), or is background session renewal via a secure token acceptable? | The main goal of this requirement is to prevent users from having to sign in too frequently after inactivity (e.g. every 30 minutes). a) Yes, authentication at the organisational level is sufficient. b) There are no specific software requirements. c) Re-authentication should only be triggered by user inactivity. d) No, a single authentication per day or per 8-hour session is sufficient.
19 | Bijlage 3: Kwaliteitsblad-Programma van Eisen | Concurrent editing prevention: locking mechanism and administration (Req 6.8) Req 6.8 requires prevention of concurrent editing of analysis records. Please clarify the expected behaviour: a) What locking behaviour does LUMC expect — should the analysis be locked for editing as soon as one user opens it (preventing simultaneous edits), or is a conflict warning upon attempted concurrent save sufficient? b) If an editing user's session expires or they close their browser, should the edit lock be automatically released after a defined timeout? c) Should a platform administrator be able to release a lock if the editing user is unreachable (e.g., absent from office)? d) Should users with read-only access (e.g., a supervising clinician or RAQA reviewer) be able to view the analysis in real time while another user holds the edit lock? | a) Yes, analysis should be locked for editing as soon as one user opens it. b) Yes, the edit lock should be automatically released after a defined timeout c) Yes, a platform admin should be able to release a lock if the user is unreachable. d) No, this is not necessary.
20 | Bijlage 3: Kwaliteitsblad-Programma van Eisen | Adding family members to an existing analysis post-hoc (Req 1.55) Req 1.55 requires the ability to add family members to an existing analysis after the initial run. Please clarify the expected behaviour: a) When a family member is added to an existing analysis, should the analysis be extended incrementally (existing results and annotations preserved, new family context added), or is a full re-run from scratch acceptable? b) Should all prior variant interpretations, ACMG classifications, and analyst annotations be preserved when the family structure is updated? c) Should there be an audit trail recording that the analysis was extended — including which family member was added, at what date, and by which user? d) Which family relationship types does LUMC expect to add to an existing analysis at a later stage — for example, a biological parent of the proband, an affected sibling, or other first- or second-degree relatives? | a) a full re-run from scratch is acceptable. The original run and findings should be preserved b) Yes, that is correct. c) Yes, that is correct. d) all mentioned relationships
21 | Bijlage 3: Kwaliteitsblad-Programma van Eisen | Filter tree visualisation: definition and expected behaviour (Req 1.32) Req 1.32 refers to a "filter tree visualisation." Please clarify the expected behaviour and scope: a) Does "filtertree" refer to a hierarchical visual representation of all currently active filters and their logical relationships (AND/OR nesting), or to a guided decision-tree-style filtering workflow? b) Should the complete filter state (all active filters and their current values) be visible at a glance on the variant table page, without requiring the user to open a separate settings panel? c) Should a configured filter tree be saveable as a named preset, shareable with colleagues within the same lab, and reusable across analyses? | a) Yes, we prefer both. b) Yes, that is correct. c) Yes, that is correct.
22 | Bijlage 3: Kwaliteitsblad-Programma van Eisen | Clinical report generation: variant selection, conclusion, and report structure (Req 13.1) Req 13.1 requires the platform to generate a report in which selected variants can be included and an overall clinical conclusion can be entered. Please clarify the expected report structure: Should the variant selection and the clinical conclusion field be part of a single unified clinical report, or is a two-document model acceptable for example, a separate technical variant report combined with a clinical interpretation document? | All in one lab report. It concerns a lab report and not a clinical report.
23 | Bijlage 3: Kwaliteitsblad-Programma van Eisen | Multiple gene panels in a single analysis (Req 1.9) Req 1.9 requires the ability to apply multiple gene panels simultaneously in one analysis. Please clarify: What is the maximum number of gene panels that LUMC expects to combine in a single analysis run? | A maximum of 5 panels.
24 | Bijlage 3: Kwaliteitsblad-Programma van Eisen | Family concordance check: scope and output (Req 1.7) Req 1.7 requires a family concordance check (trio-check). Please clarify the expected scope and output: a) What specific checks are expected — Mendelian error rate per variant, per sample, genome-wide summary statistics, or a combination? b) How should discordant variants be flagged — per-variant row annotation in the variant table, a dedicated summary panel, or both? c) Should the concordance check support non-standard family configurations, such as half-siblings or trios where one parental sample has limited coverage? d) Is the concordance check result expected to be included in the exported clinical or analysis report? | The following answers are in context of family concordance a) trio check (parental check genome wide) b) No, this is not what we want. c) No, not necessary d) If possible yes
25 | Bijlage 3: Kwaliteitsblad-Programma van Eisen | (PvE), Req 2.4 — "Upload and process 75 GB within 60 minutes" What files are meant by "75 GB" in Req 2.4? a) Full fastr/BAM/CRAM files (~50-75 GB) or VCF files (~1-5 GB) for tertiary analysis? b) If BAM: is the 60-minute requirement including network transfer or processing only? c) Measured per sample or per analysis (e.g., trio = 3 samples)? | a) VCF files (1-5GB) b) Processing only c) Per sample
26 | Bijlage 3: Kwaliteitsblad-Programma van Eisen | (PvE), Req 7.1 — "99.5% uptime during business hours" What methodology does LUMC use for measuring uptime? a) Synthetic monitoring, user reporting, or other? b) Definition of 'business hours' for SLA? c) Are planned maintenance windows excluded? | a) User reporting b) 8am-18pm c) No, maintenance preferably out side business hours
27 | Bijlage 5: Prijzenblad | Genepanel (captured NGS) (Singleton) is optional, could you clarify the following? - What is the timeline to include this in the solution? - Is there 1 gene panel or multiple expected? - What is the size of the gene panel(s) in number of genes? - What is/will be the secondary analysis solution for the genepanel(s)? | The timeline is unsure and depends on offer. Multiple panels are expected. 2700 genes are captured and the panels are smaller (<100 genes)
28 | Bijlage 3: Kwaliteitsblad-Programma van Eisen | (PvE), Req 1.14 — "UPD detection including segmental heterodisomy" For the mandatory UPD detection requirement (1.14), including segmental heterodisomy: a) Should the platform detect both isodisomy AND heterodisomy with parental origin determination? b) Is UPD detection expected in singleton analyses (ROH-based) as well as trio analyses (parental comparison)? c) Which input files are expected: VCF-based (WGS), or BAM with calling performed by the vendor ? both? d) Does LUMC use a specific UPD calling tool as reference? | a) Yes, that is correct. b) No, not ROH based c) If possible VCF, in the future also with methylation d) Yes we have a UPD calling tool.
29 | Bijlage 3: Kwaliteitsblad-Programma van Eisen | (PvE), Req 1.2 — "SNV, CNV/SV, repeat analysis from VCF" Can you clarify whether repeat analysis means refreshing in full the annotation of the case ? or doing a second review ? or doing a new report on a case that was already reported? | Performing a second review with new phenotype information or gene panel.
30 | Bijlage 4: Beoordelingsformulier NGS-analysesoftware | Pilot test uses Dragen pipeline data. Please clarify file formats: a) VCFs as gVCF, hard-filtered VCF, or both? b) BAM or CRAM files included? c) STR repeat VCFs (e.g., ExpansionHunter) for repeat analysis cases? d) CNV/SV VCFs separate or combined? | a) hard-filtered VCF b) Depending on the funcitionality c) Yes, that is correct. d) separate at this moment
31 | Bijlage 6: I-EU-25-03 Overeenkomst NGS Analyseplatform | Article 3.2b refers to "[number]" users. Can you specify the expected number of concurrent users the license should accommodate? | We expect 75 concurrent users.
32 |  | We hebben een vraag over het tabblad ‘PvW’ van 'Bijlage 3 Kwaliteitsblad-Programma van Eisen'. In de kolom 'Uw antwoord' zou overal de optie 'ja' of 'nee' te zien moeten zijn. Voor verschillende wensen zijn er echter andere keuzemogelijkheden, maar niet 'ja' of 'nee'. Een voorbeeld is wens 1.17 'The platform must support methylation analysis.' (rij 23 in de Excel file), waarbij de keuzemogelijkheden 'Meer dan 9 jaar', 'Meer dan 8 jaar', 'Meer dan 7 jaar', 'Meer dan 6 jaar', 'Meer dan 5 jaar' of '5 jaar of minder' zijn, maar dus niet 'ja' of 'nee'. Hierdoor kan deze kolom dus niet correct ingevuld worden voor alle wensen. | The dropdown criteria in column G of the Program of Requirements have been corrected. You can now select the correct answers corresponding to the relevant questions.
33 |  | We hebben nog een vraag i.v.m. de GVA (Gedragsverklaring Aanbesteden) en VvdB (Verklaring van de Belastingsdienst) documenten die enkel aan te leveren zijn na eventuele selectie. Als we het goed voor hebben, kunnen we als Belgisch bedrijf deze specifieke documenten niet bekomen en aanleveren, maar moeten we hiervoor terugvallen op de Belgische tegenhangers. Volgens ons zijn dit het 'uittreksel uit het strafregister' (tegenhanger GVA) en het 'attest voor openbare aanbesteding' (tegenhanger VvdB). Kunnen jullie bevestigen dat we hiermee zouden voldoen aan deze eisen? Kunnen jullie ook nog even meegeven hoeveel tijd we na eventuele selectie exact zouden hebben om alle documenten aan te leveren die niet meegestuurd moeten worden bij submissie, maar wel na eventuele selectie? Zo kunnen we beter inschatten of we al deze documenten al dan niet op voorhand moeten opvragen. | You may submit the documents 'uittreksel uit het strafregister' and 'attest voor openbare aanbesteding'. These are accepted as equivalent to the Dutch GVA and VvdB. After selection, you will have 5 calendar days to submit all requested documents that did not need to be included with the initial submission.
34 |  | we have find contradictory information regarding the language required to submit documents: In TN576297 - EF16 Aankondiging van een opdracht - algemene richtlijn, standaardregeling 20260326170743.pdf page 3, it says language is Dutch In 01. Offerteaanvraag NGS-analyseplatform.pdf chapter 14, it says Dutch or English. Could you please let us know in which language we can submit the documents? | All documents that are provided in Dutch will not be provided in English. However, it is allowed to ask questions and/or submit the tender in English.
35 |  | Welke sequencer is het meest geschikt voor uw dry lab-pipeline? Welk bestandstype is het meest geschikt voor het uploaden naar onze AI-software: FASTQ, VCF of BAM? Graag specificeren. Moeten uw database/variantflags worden geïmporteerd? Wilt u RNA-seq routinematig gebruiken in combinatie met DNA- en/of epigenetische modules? Kunt u de genenlijst voor uw custom capture NGS-panel met ons delen? | Sequencing is performed elsewhere. This concerns data analysis only. File type: DRAGEN VCF files. In the future, this may also include data from long-read sequencing (VCF). The database must be imported. With regard to RNA-seq in combination with DNA and/or epigenetic modules, this may be considered for the future.
36 |  | 1.15 Heranalyse. Kunt u de reikwijdte van de heranalyse-vereisten toelichten? 1.33 Het platform moet filtering ondersteunen op basis van (eerdere) interne annotaties. Kunt u details geven over de reikwijdte van de (eerdere) interne annotaties? Overige vragen: Droog lab of bundel (reagentia + dry lab)? Is validatie vereist? Is er behoefte aan MCP-validatie en een prestatierapport & eventueel non-regressierapporten? Zijn er vereisten met betrekking tot kwaliteitsbeoordeling? Zijn er vereisten met betrekking tot het delen van gegevens met samenwerkingspartners/secundaire accounts (bijv. peer-netwerk)? Zijn er vereisten met betrekking tot servicelevering (bijv. levering van NGS-gegevens en analyses aan klanten)? | 1.15: The platform must support reanalysis of existing, previously analyzed data without overwriting or deleting the results of the primary analysis. 1.33: We would like to see the previous classification of the variant. Sequencing is performed externally. Yes, validation is necessary. Yes, there are requirements related to collaborations.