## Requirements
Answers to Program of Requirements XRay Defraction tool If to any of the requirements, are not met by Tenderer, he will be excluded for any further assessment. By answering a question with “compliant” will mean that in the event of awarding the contract, non-conformity to the requirement shall count as failure to meet an obligation from the Agreement to be concluded.
Requirement | Type | Description | Validation method & moments | Compliant to requirement? | Reference in Tenderer's documentation where compliance with the requirement is demonstrated
General
R-1000-005 | Handling InP 6" Wafers | The equipment shall be capable to operate correctly on wafers with the following specifications: - Substrates: InP and Silicon (Si); - Double side polished; - Diameter: 150 ± 0.5 mm (6” wafers); - Thickness: 675 ± 200 μm; - For InP: Alignment on notch. Notch axis [010], SEMI M23-0623; - Allowable warp and bow: 50 μm. | R-TD, D – FAT, SAT | select
R-1000-010 | Cassette loading | The equipment shall provide for automated 6” wafer loading with cassette-to-cassette handling (open cassette configuration) | R-TD, D – FAT, SAT | select
R-1000-015 | Stress-free wafer holding | The equipment shall be capable of holding wafers with no significant additional stress to the wafers. | R-TD, D – FAT, SAT
Equipment performance
R-1100-005 | High accuracy and precision with fast positioning feedback time | The equipment shall have the following characteristics: - A motorized goniometer; - Motions include Chi, Phi, X, Y and Z; - Range Phi: 360 unlimited, - Range Chi -10 to 100 °, - Chi and Phi resolution and reproducibility < 0.01° - X and Y translation to cover full wafer surface and with resolution and reproducibility < 0.1 mm - Z range > 10 mm, resolution and reproducibility < 0.001 mm - Omega (or Theta) and 2Theta drives with resolution and - Reproducibility on Omega < 1'' (= 0.000278°) and resolution on 2 Tetha < 1'' (=0.000278°); - Fully automated alignment of the sample under X-ray beam; - beam knife for grazing incidence measurements; - Support of 002, 004 and 006 measurements. | R – TD | select
R-1100-010 | X-ray power | X-ray source power operating range shall be (2.0 kW..4 kW) or higher | R – TD | select
R-1100-015 | Detector capabilities | The equipment shall have the ability to do reciprocal space mapping. | R – TD | select
R-1100-020 | Detector types | The solid-state hybrid pixel detector shall be capable of 0-D, 1-D and 2-D diffraction compatible with all common X-ray emission lines. | R - TD | select
R-1100-025 | Analytical capabilities | The main purpose of the equipment is to determine the properties and quality of epitaxially grown InP, InGaAs, InAlGaAs and InGaAsP layers on InP substrates. The equipment hardware and analysis software, shall be able to measure: - Thickness of InGaAsP layers on InP layers with <0.5 nm resolution; - Determination of InGaAsP & InAlGaAs alloy composition based on lattice constants with accuracy of < ± 10 ppm; - Compressive and tensile strain of InGaAsP and InAlGaAs quantum well and barrier layers with a resolution of < 0.05%. | R – TD, D – SAT | select
Instrument control and software
R-1200-005 | Instrument control | The equipment shall be controlled by a computer that: - Controls the equipment and performs automated measurements; - Has the ability to run analysis software during measurements. | D – SAT | select
R-1200-010 | Operator interface and Automation | The equipment shall offer the following characteristics: - An operator interface allowing to pre-program, queue and execute several jobs (same/different scans on different wafers with autofitting and reporting); - The ability of batch processing. | D – SAT | select
R-1200-015 | Automated analysis | The equipment shall offer the following characteristics: - The presence of modelling software and Autofitting technology to extract critical parameters such as layer thickness, lattice constants, stress/strain or composition (see R-2000-015); - The ability to generate data reports in HTML, pdf, excel format and others with SPC reporting; - Software that shall plot wafer maps of the above critical parameters as well as reciprocal space maps; - This software shall be able to run on additional computers, not necessarily connected to the equipment, for the purpose of further data analysis and post processing Rationale: raw diffraction data alone is not enough. | R – TD, D – SAT | select
Productivity requirements
R-1300-005 | Throughput | The equipment shall have the capacity to process 60 wafers per 8 hours as a minimum, whereby each wafer shall be measured multiple times (at least 9 points on each wafer) and with enough accuracy to achieve the requirements in R-2000-015. | R – TD, D – SAT | select
Customer support
R-2000-005 | Spare parts and accessories availability | Spare Parts and Accessories kit shall include a reference sample that can be used to calibrate the system and perform Statistical Process Control (SPC). | R – TD, D – SAT
Reliability and availability
R-2100-005 | Equipment | Following the guidelines for definition and measurement of equipment reliability and availability (SEMI-E10), the equipment shall achieve these specifications: • The equipment-dependent up/down times are 4-weeks moving averages. For calculation of the uptime the waiting times for internal staff and the facilities related downtimes are not counted. • Uptime: > 85%. • Scheduled and unscheduled downtime: < 15%. | R – TD | select
R-2100-010 | Software | The following specifications shall be achieved: • Rebooting due to software failures shall not happen more than once a month; • Software failures will be logged and reported to the Supplier on a monthly basis. The Supplier will respond within one week after such a report with a corrective action plan. | R – TD | select
R-2100-015 | Lifetime | The equipment and all its parts, except those parts that are designated to be consumable or regular to be exchanged due to foreseen “wear and tear” by normal use, shall be designed for a useful life of 10 years. |  | select
R-2100-020 | Cleanroom compatibility | The equipment will be installed in a cleanroom with ISO5 air cleanliness conditions. The equipment shall not degrade this cleanroom to a condition below the parameters as formulated in ISO5. |  | select
Safety
R-2200-005 | Regulations | The design and implementation of the equipment shall adhere to the relevant and applicable legislation and regulations in conformity with the Semiconductor equipment safety standards primarily governed by SEMI, ISO/IEC, NFPA, Machinery Directive 2006/42/EC essential health and safety requirements including ergonomic principles, and OSHA regulations and with certification for European Conformity trademarks including product labelling (CE). The Tenderer shall provide comprehensive information to enable TNO to meet its obligations regarding the safe use of the equipment. This includes detailed written instructions on the procedures for installation and operation, as well as guidelines for the safe handling, storage, and disposal of any materials necessary for the equipment's operation, especially those that are hazardous. Additionally, the Tenderer shall supply sufficient information to ensure that TNO's use of the equipment complies with local environmental protection regulations. Rationale: TNO safety requirement | R – TD | select
R-2200-010 | CE Certification | The equipment shall be constructed in accordance with CE-directives and regulations and will be marked accordingly. Rationale: Safety and quality of the equipment. | R – TD | select
R-2200-015 | Emergency power off | The equipment shall fail safely in the event of an electrical power failure or operating system malfunction, in compliance with ISO13850. Rationale: Machine and operator safety | D – FAT, SAT | select
R-2200-020 | Marks & Labels | All connections and fittings for electrical in- and outputs, water in- and outlets, process gas inputs, vacuum and exhaust ports are provided with distinguishing marks, text and labels. All interconnecting cables and lines will be clearly marked at both ends. The identification of these cables and lines will be in correspondence with the identification used in the system documentation. Gas lines are marked with the applicable medium and flow direction. Rationale: labelled connections and fittings will increase safety and decrease installation and troubleshooting time. | D – FAT, SAT | select
R-2200-025 | Information Security | To safeguard reliability of the system’s internal communication and its communication with host computers and other systems, shall comply with the following: 1. All software delivered with the system is checked on ‘alien elements’ (e.g. viruses, malware). Remark: If required, the Tenderer will provide evidence that the software is checked and that there is no risk for viruses etc. 2. Different access levels can be defined for operator, engineer, specialist etc. to prevent unauthorized access to e.g. process recipe settings, system constants, maintenance routines, achieved data etc. Access to the different levels is password protected and controlled. 3. Tenderer will provide information and instructions on back-up and recovery procedures (e.g. after facilities are powered down). Emphasis shall be on saving the relevant data in such cases. 4. Computers are supplied with up to date Windows operating system. Rationale: compliance and security | D – SAT | select
R-2200-030 | X-ray protection | The equipment shall provide for: - An effective safety enclosure, including safety interlocks compliant with X-ray radiation safety standards; - A highly visible "X-ray on" and "Shutter open" sign/display/light.
Electrical equipment interfaces
R-2300-005 | Legislation | The system design and implementations shall implement relevant and applicable legislation and regulations such as EU Machinery Regulation EU2023/1230, REACH and RoHS and global standards (SEMI, IEC60204-1, ISO13849, IEC62061). Rationale: TNO safety requirement | R - TD | select
Communication interfaces
R-2400-005 | Basic connectivity | The equipment shall be capable of integration with a Network Attached Storage (NAS) device for the purpose of backing up log files and recipes. Rationale: TNO safety requirement. | R – TD, SAT | select
R-2400-010 | SECS GEM | The equipment shall support an up-to-date version of the SECS/GEM communication standard that allows monitoring of the equipment, collecting process data and remote control of the equipment. The Tenderer shall supply the documentation of supported SECS/GEM commands and features for review. | R - TD | select
Contractual aspects
R-3000-005 | (Critical) design review (CDR) | After acceptance of the order, the Tenderer will prepare a Critical Design Review (CDR) in the form of detailed design documentation and will organize a presentation and review session with TNO. The objective of the CDR is to confirm that the equipment and all deliverables satisfies the requirements whereby 75 percent to 90 percent of (manufacturing quality) product drawings, software design specification(s), and associated instructions should be complete, and that 100 percent of all safety-critical components drawings are complete. It establishes the compatibility of external and internal interfaces and the compliance of the design with contractual requirements (predicted and/or measured) based on the product baseline documents of the Tenderer. The successful completion of the CDR permits the Contractor to start the manufacturing of the equipment. The CDR shall contain the following details: • Physical dimensions of the system and system components. • External and internal interface location and dimensions of the system. • Showing the design meets the functional and performance requirements. • Certification and safety compliance. • List of required utilities for operating the equipment including (water, electricity, special gasses, exhaust, vacuum, Clean Dry Air (CDA)) and capacity. • The acceptance of the CDR shall by achieved using a written document, to be signed by both parties. If the CDR fails to meet the requirements, the Contractor or TNO shall set out such failure in a written statement, to be signed by both parties. The Contractor shall then provide TNO with a corrective action plan within fourteen (14) calendar days from the date of signing. The Contractor shall remedy these deficiencies at no additional cost within the timeframe set out in the corrective action plan, but not exceeding 1 month. | R-After contract is granted, before manufacturing starts | select
R-3000-010 | Factory Acceptance Test | Factory Acceptance Test (FAT) is a process that evaluates the equipment during and after the assembly process by verifying that it is built and operating in accordance with design specifications. The equipment will undergo testing at the Tenderer’s location, with one or two TNO representatives potentially present, prior to shipment approval to TNO. The FAT shall include, at a minimum, the testing of basic functions and the requirements outlined in the program of requirements that are indicated with FAT in the third column. The specifics of the FAT will be determined collaboratively between TNO and the selected Tenderer. TNO requires that the selected Tenderer first submit a proposal to TNO, based on TNO’s requirements and the Tenderer’s accepted preferences. Acceptance for shipment will only be granted when all data corresponds to the specifications, and all Acceptance Parts are manufactured in compliance. A report of the FAT shall be submitted to TNO | D, I, R, FAT | select
R-3000-015 | Site Acceptance Test | Site Acceptance Test (SAT) is a process in which a machine is tested and accepted at the customer's site. The purpose of the SAT is to ensure that the machine or system has been properly installed and configured and is ready for operation. The SAT contains the following checkpoints: • The equipment will be checked for damages; • The proper functioning of the equipment will be checked; o Mechanical functions, o Electrical safety- and switching functions, o Gas safety and functions (if applicable), o Vacuum functionality (if applicable), o User interface(s), including software. • All requirements that have a “SAT” indication in the program of requirements 4th column shall be checked; • The presence and quality of the documentation will be checked. A comprehensive user acceptance test on-site shall be conducted by the Tenderer in collaboration with TNO. TNO requires that the selected Tenderer first submit a proposal to TNO, based on TNO’s requirements and the Tenderer’s accepted preferences. Upon successful completion of the on-site user acceptance test, the equipment shall be at TNO's unrestricted disposal for at least one month, with application support from the Tenderer to set up and verify the most critical processes. If the equipment operates according to TNO’s functional and technical specifications, without any faults or malfunctions during this period, the test period will be deemed successfully concluded by TNO. Should the equipment fail to meet these specifications, any faults or malfunctions will be addressed, and another one-month test period shall be conducted and approved by TNO, up to a maximum of three one-month test periods. The specifics of the site acceptance test shall be mutually agreed upon by the user and the Contractor. The Tenderer shall submit a draft test protocol for the site acceptance test with its Tender. If the deliverables do not meet the user Acceptance Criteria for the SAT, the Tenderer shall document these failures in a written statement, to be signed by both parties. The Tenderer shall then provide TNO with a corrective action plan for these failures within fourteen (14) calendar days from the date of signing the statement. The Tenderer shall rectify these deficiencies at no additional cost within the timeframe outlined in the corrective action plan, not exceeding three months. A new acceptance test will then be conducted. | D, I , R, SAT | select
R-3000-020 | Terms and conditions of payment | Terms and conditions of payment TNO shall apply the following payment terms and conditions: • 70% upon delivery at location of TNO; • 30% after commissioning and Site Acceptance Test (SAT) including test periods, approved by TNO, Management Technology and Procurement. |  | select
R-3000-025 | Mill Test | When supplying the equipment, the Tenderer shall provide a Mill Test Certificate (MTC), or similar certificate such that it meets the requirements and obligations to establish the country of origin (further specified in paragraph 5.2.4 of the Procurement Guide). |  | select
R-3000-030 |  | The Tenderer shall have comprehensive maintenance facilities within Europe. These facilities should be equipped to handle all aspects of maintenance, including diagnostics, repairs, and parts replacement. The facilities shall be strategically located to ensure a quick and efficient delivery of services. The Tenderer shall provide details of the current service organization. |  | select
R-3000-035 |  | The Tenderer shall provide a user manual in English with detailed description of the hardware and software, including any custom-made options. Rationale: To provide clear understanding of the system's functionalities, facilitate effective utilization and help with the troubleshooting of the instrument. |  | select
R-3000-040 |  | The Contract will be executed through a purchase order, incorporating the Purchasing Conditions for Goods TNO 2022 and Incoterms®2020 Delivered Duty Paid (DDP), see paragraph 2.2.18 of the Procurement Guide. |  | select
Maintenance/Services contract
R-4000-005 | General | Tenderer shall submit a Maintenance and Services contract that complies with: • Section 1.5 of the Procurement Guide; • Requirements R-4000-010 and R-4000-015 | R-TD | select
R-4000-010 | Maintenance activities | The activities within the maintenance that Tenderer offers, are mentioned in broad lines below. 1) Preventive maintenance In addition to Annex C04: Tenderer shall conduct preventive maintenance with a frequency necessary with respect to the throughput and uptimes as needed. 2) Predictive maintenance Tenderer shall implement predictive maintenance in order to improve the effectiveness and efficiency of maintenance planning. 3) Corrective maintenance Tenderer shall execute corrective maintenance within the time frames as formulated in requirement R-4000-015 4) Condition based maintenance In addition to Annex C04: For now, condition based maintenance is not part of de maintenance and services contract (and it is also not applied during the warranty period). The reason for this lies in risk analyses to be performed concerning cyber security. 5) Software maintenance In addition to Annex C04: Tenderer shall provide software updates free of charge for the entire expected lifetime of the equipment. If a hardware modification is required to implement a software update, the Tenderer shall present this option to TNO for approval. Should TNO choose to proceed with the hardware modification, no additional license fees shall apply for the software update. The cost of the required hardware shall be itemized separately in accordance with the pricing schedule. Tenderer shall provide technical support for the implementation and operation of software updates throughout the minimum expected lifetime of the system. | R-TD | select
This support shall include assistance with installation, configuration, troubleshooting, and user guidance related to the updates. The support as mentioned shall be at no additional cost for TNO; Upgrades in itself are not included in the standard maintenance agreement and may be subject to additional fees. Tenderer shall inform TNO of available upgrades at least 30 days prior to release, including details of functionality, pricing, and potential impact on existing systems. TNO reserves the right to accept or decline any offered upgrade.
R-4000-015 | Other activities | Tenderer shall have a support organization in place for the equipment including software applications, that complies to the following: a) Response (acknowledgement) to service requests within 2 hours in all cases and: o Start of solving the problem within 24 hours in cases with severe impact on business operations; o Start of solving the problem within 5 working days in cases with significant impact on business operations; o Start of solving the problem within 10 working days in cases with no significant impact on business operations; The impact level is to the discretion of TNO. b) Tenderer shall maintain an inventory of critical and/or expensive spare parts, including but not limited to reactor liners, susceptor, gas injection units, and temperature control units. These parts shall be readily available to accommodate service requests with severe impact on business operations; c) 1. The Equipment shall be delivered with basic training on hardware, software, and safety, as well as FAB automation (semiconductor fabrication automation), MES integration, and SECS/GEM, with a minimum of 6 personnel trained 2. Additionally advanced training is provided for personnel on maintenance, calibration, troubleshooting, and process knowledge, with a minimum of 4 personnel trained. d) Detailed service reports shall be provided after each maintenance or repair activity. These reports should include the nature of the issue, actions taken, parts replaced, and recommendations for future maintenance. The latter is part of a yearly summary log of maintenance performed. | R-TD | select

## Informative Questions
Answers to Questions
X000-05 | Corporate Social Responsibility is becoming increasingly important in our own organization and in the future TNO will also ask more from our Contractors to play a part in this. What possibilities has your company already initiated and where does your company see opportunities for itself and for TNO in future?
X000-10 | As your company is aware off, ISO 20400 is an international standard that provides guidelines for sustainable procurement. The standard can be used by any organization that intends to improve their social, economic, and environmental sustainability. Considering the growing awareness with regard to sustainable business practices, nationally and internationally, whether required by laws and regulations, promoted by the UN Sustainable Development Goals, or even by the changing societal expectations – there are a myriad of reasons in today’s world that drive organizations to be more sustainable, including TNO. Does your company embrace the ISO 20400 guideline and, if so, is your company currently implementing it in your own organisation or will your company be implementing it in the coming period?
X000-15 | Laboratory instruments that TNO uses for its scientific research consume a lot of power and also consist of very high-quality raw materials. For the benefit of corporate social responsibility (CSR), is your company making or willing to make technological investments reducing power consumption when using the instrument, without negatively affecting the technical application of the instrument?
X000-20 | Biobased Goods are made of materials derived from fully renewable sources. TNO is considering in future that devices be made (in part) from biobased materials. Can your company give TNO some insight into whether this is feasible for the Goods in your portfolio, and if so, which biobased Goods does your company think could be used as alternative Goods?
X000-25 | Due to global supply chain disruptions - driven by geopolitical, health, and environmental crises - many organizations have faced significant operational challenges. What measures has your organization taken to ensure timely delivery under these conditions?

## validatie (hidden)
select
Yes
No
note: dit blad 'verbergen' indien je dit gaat delen intern/extern (behoudens aan inkoop collega's)