	
	
	



Market consultation Full Anechoic Room - The Dutch Authority for Digital Infrastructure
As part of this market consultation, the Dutch Authority for Digital Infrastructure is seeking market parties with demonstrated expertise in the supply, installation, and commissioning of a Full Anechoic Room. 

Through this market consultation, the Dutch Authority for Digital Infrastructure seeks to obtain insight into the solutions and capabilities available in the market for the delivery of a Full Anechoic Room at the intended location of the Dutch Authority for Digital Infrastructure. Furthermore, this market consultation is intended to gather relevant market information and insights to support the preparation of a potential future procurement procedure.

Participation in the Market Consultation
The Dutch Authority for Digital Infrastructure invites market parties to participate in this market consultation and greatly appreciates a response. 

Participation confers no rights, and participation does not lead to a preferential position or any other form of advantage in any potential follow-up process.

Background and Requirements
The Dutch Authority for Digital intends to rent additional space in which a Full Anechoic Room is to be constructed. This space is located at Piet Mondriaanlaan 54, 3812 GV Amersfoort, near the central station in Amersfoort. The space is situated on the ground floor above a parking garage and has a net available height of 3.8 meters. The width and length can still be determined as needed to meet the specification requirements.

The Dutch Authority for Digital Infrastructure requires a fully compliant Full Anechoic Room to be installed within these premises for the testing of electrical and radio equipment in accordance with the EMCD/RED. The Full Anechoic Room shall comply with the following requirements:

	Full Anechoic Room FAR 3m specifications 

	· Full compliant radiated emission CISPR 16-1-4, NSA 30MHz – 1000MHz, SVSWR 1 – 18GHz

	· Shielding effectiveness tested to EN 50147-1

	· ETSI TR 102-273-2 V1.2.1

	· Full compliant radiated immunity EN 61000-4-3, 80 – 1000MHz, 1 – 18 GHz

	· Antenna measuring distance 3m

	· Measuring volume 1.5m

	· Turntable

	· 3 phase AC mains 400VAC 3x 25A filtered

	· Interface/feedthrough panel: Ethernet, USB, 4 x N connector, 4 x BNC, waveguide 1x3 inch, waveguide1x 1 inch, both waveguides with endcap

	· 1x water supply, 1 x waterdrain

	· Antenna position at control room side

	· Anechoic chamber lighting

	· 2 x Honeycomb air filter for air ventilation

	· Interlock on entrance door

	There are no restrictions on length x width dimensions of the Full Anechoic Room. The height restriction of the Full Anechoic Room is 3.8m. (also placing/mounting of the Full Anechoic Room (must be realized within this height).



	Questions:
	Answers:

	Are you able to deliver and install a Full Anechoic Room that complies with the specified requirements?
	

	Are you able to deliver and install a compliant Full Anechoic Room (FAR) within the maximum available height of 3.8 m?
	

	If a compliant Full Anechoic Room (FAR) cannot be created and installed, what can you offer instead? Please explain why a compliant Full Anechoic Room cannot be achieved and identify the constraints or limitations.
	

	For the design process, please indicate the dimensions (length × width × height) of the Full Anechoic Room you are able to deliver.
	

	What is the expected delivery and installation timeframe for the Full Anechoic Room?
	

	What essential information should the Dutch Authority for Digital Infrastructure provide to enable you to prepare for a potential procurement procedure and the subsequent implementation of the project?
	

	What floor load capacity is required for the installation of the Full Anechoic Room?
	

	Do you foresee any implications or constraints associated with installing the Full Anechoic Room in premises located above a parking garage? If so, which implications or constraints?
	

	Do you foresee any consequences arising from the location near the railway station? (For example, in terms of vibrations and noise). If so, please describe these consequences
	

	What factors must the RDI take into account regarding the additional space required around the anechoic chamber when constructing/placing it?
	

	What opportunities do you see for incorporating sustainability into this project?
	



Responses to this market consultation may be submitted via TenderNed until Monday, 
Monday, August 31, 2026.

Responses received after this deadline will not be taken into consideration as part of this market consultation.

The Dutch Authority for Digital Infrastructure would like to thank you in advance for your cooperation with this market consultation.
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