## Lot 2 - data logger
Response form Lot 2 - Program of Requirements and awarding criteria
Awarding criteria - Lot 2: data logger
Type | # | Datalogger quality consisting of: | Answer | Answer tenderer
KO | D1 | The provided datalogger will have simultaneous sampling on all channels | Yes/No
KO | D2 | The provided datalogger will have a resolution of at least 24 bit @ 200 sps | Yes/No
KO | D3 | The provided datalogger will support channel input configurations within the 6–16 channel range, including 6, 12, and 16 channels. | Yes/No
KO | D4 | The provided datalogger is capable of continuous recording with MiniSEED support | Yes/No
KO | D5 | The provided datalogger will be able to continuously stream data, seismic and state-of-health, supporting the SeedLink protocol | Yes/No
KO | D6 | The provided datalogger provides automatic data backfill for when communication is interrupted | Yes/No
KO | D7 | The provided datalogger will have a timing accuracy < 5 µs via internal GNSS receiver and external antenna | Yes/No |  | more than 3 extra channel
KO | D8 | The provided datalogger will have maximum dimensions of 56 x 64 x 30 cm (lxwxd). If multiple components need to be assembled together, this is done in the most space-efficient way possible. | Yes/No |  | 2 or 3 extra channel
KO | D9 | The provided datalogger will have an internal data archive storage (separate from OS storage) of at least 32 GB | Yes/No |  | 1 extra channel
KO | D10 | The provided datalogger supports (incidental) file transfer through internet | Yes/No |  | no extra channel
KO | D11 | The provided datalogger has remote self diagnostics, configuration, tests and calibration capabilities through internet, in combination with the offered sensor | Yes/No
KO | D12 | The provided datalogger has an operating temperature range of -20°C to 60°C | Yes/No
KO | D13 | The provided datalogger has an IP67 equivalent or higher rating | Yes/No
KO | D14 | Full datasheet with specifications of the datalogger is submitted | Yes/No
KO | D15 | The nominal response of the data logger is available online (for example, in the IRIS Nominal Response Library (NRL) | Yes/No
KO | D16 | Including connection on the datalogger for (bare-wire) geophone channels, to be connected in field conditions (i.e. break-out box/terminal board) | Yes/No
KO | D17 | Poweradapter for connection to 220V with schuko male is included | Yes/No
KO | D18 | The provided datalogger has undergone standardized performance tests by an independent laboratory (e.g. Sandia National Laboratories). The evaluation report is available and will be provided. | Yes/No
KO | D19 | The provided datalogger will have one physical connection (RJ45) and one IP address | Yes/No
Type | # | Other criteria: | Answer | Answer tenderer
KO | O1 | Delivery time up to 120 days, longer is a knockout | Yes/No |  | ≤60 days
KO | O2 | Warranty Free of charge replacment of faulty components for (at least) 12 months. A component is deemd faulty if it causes the station to deviate from the knockout criteria | Yes/No |  | >60 days ≤90 days

## Blad1 (hidden)
Yes | ≤
No | ≥
S12
0-1.5 | 10 points
1.5-3 | 5 points
>3 | 0 points
S13 (Hz)
<0.01->50 | 20 points
>0.01-<50 | 10 points
0.01 Hz-40 H | 0 points
S14
>10 Hz | 25 points
>0.08<10 | 12 points
0.1-5 Hz | 0 points
S15
>150 | 10 points
140≤150 | 5 points
135 dB to 140 dB | 0 points
D14
Resolution > 24 bit @ 100 sps
≤24 | 0 points
>24 | 10 points
D15
>3 | 30 points
2-3 | 10 points
1 | 5 points
0 | 0 points
D16
Yes | 5 points
No | 0 points
D17
Yes | 5 points
No | 0 points
D18
Yes | 10 points
No | 0 points
D19
1.5-3 | 5 points
>3 | 0 points
W05
0≤30 | 20 points
31≤60 | 10 points
61≤90 | 0 points
W06
≥36 | 10 points
>12<36 | 5 points
≤12 | 0 points