## Lot 1 - Surface accelerometer
Response form Lot 1 - Program of Requirements and awarding criteria
Awarding criteria - Lot 1: Surface Accelerometer
Type | # | Sensor quality consisting of: | YES/NO / Max. points | Answer tenderer
KO | S1 | The provided instrument has an internal datalogger | Yes/No
KO | S2 | The provided instrument will have a leveling bubble (physical and/or digital) | Yes/No
KO | S3 | The provided instrument will have a horizontal alignment mark (North for example) | Yes/No
KO | S4 | The provided instrument will have maximum dimensions of 20 x 20 x 15 cm (lxwxh) | Yes/No
KO | S5 | The provided instrument will have an operating temperature range of -20°C to 60°C | Yes/No
KO | S6 | The provided sensor will be a triaxial force balance accelerometer | Yes/No
KO | S7 | The provided sensor will have a flat instrument response to ground acceleration from 0 Hz to 100 Hz | Yes/No
KO | S8 | The provided sensor will have a self-noise that is below the -130 dB (1 m/s^2 reference value) from 1 Hz to 30 Hz | Yes/No
KO | S9 | The provided sensor will have a dynamic range of at least 130 dB between 1 and 15 Hz @100 sps | Yes/No
KO | S10 | The provided sensor will have an adjustable Full Scale range, from +/-0.25g - +/-2g | Yes/No
KO | S11 | The internal datalogger will have simultaneous sampling on all channels | Yes/No
KO | S12 | The internal datalogger will have a resolution of at least 24 bit @ 200 sps | Yes/No
KO | S13 | The internal datalogger will have at least 3 channel input; more is better (see D14), less is a knockout | Yes/No
KO | S14 | The internal datalogger is capable of continuous recording with MiniSEED support | Yes/No
KO | S15 | The internal datalogger will be able to continuously stream data, seismic and state-of-health, supporting the SeedLink protocol | Yes/No
KO | S16 | The internal datalogger provides automatic data backfill for when communication is interrupted | Yes/No |  | more than 3 extra channel
KO | S17 | The internal datalogger will have a timing accuracy < 5 µs via internal GNSS receiver and external antenna | Yes/No |  | 2 or 3 extra channel
KO | S18 | The internal datalogger will have an internal data archive storage (separate from OS storage) of at least 8 GB | Yes/No |  | no extra channel
KO | S19 | The internal datalogger supports (incidental) file transfer through internet | Yes/No
KO | S20 | The internal datalogger has remote self diagnostics, configuration, tests and calibration capabilities through internet, in combination with the internal sensor | Yes/No
KO | S21 | The nominal responses of the sensor and data logger are available online (for example, in the IRIS Nominal Response Library (NRL)) | Yes/No
KO | S22 | The provided instrument and cables will have an IP67 equivalent or higher rating | Yes/No
KO | S23 | Poweradapter for connection to 220V with schuko male is included | Yes/No
KO | S24 | The provided instrument has traceable calibration and has undergone standardized calibration/performance tests by an independent laboratory (e.g. Sandia National Laboratories). The evaluation report is available and will be provided. | Yes/No
KO | S25 | The full datasheet with specifications, factory instrument response, and self noise level curves is submitted | Yes/No
Type | # | Other criteria: | Max. points | Answer tenderer | ≤60 days
KO | O1 | Delivery time up to 120 days, longer is a knockout | Yes/No |  | >60 days ≤90 days
KO | O2 | Warranty Free of charge replacement of faulty components for (at least) 12 months. A component is deemd faulty if it causes the station to deviate from the knockout criteria | Yes/No |  | >90 ≤ 120 days

## Blad1 (hidden)
Yes | ≤
No | ≥
S12
0-1.5 | 10 points
1.5-3 | 5 points
>3 | 0 points
S13 (Hz)
<0.01->50 | 20 points
>0.01-<50 | 10 points
0.01 Hz-40 H | 0 points
S14
>10 Hz | 25 points
>0.08<10 | 12 points
0.1-5 Hz | 0 points
S15
>150 | 10 points
140≤150 | 5 points
135 dB to 140 dB | 0 points
D14
Resolution > 24 bit @ 100 sps
≤24 | 0 points
>24 | 10 points
D15
>3 | 30 points
2-3 | 10 points
1 | 5 points
0 | 0 points
D16
Yes | 5 points
No | 0 points
D17
Yes | 5 points
No | 0 points
D18
Yes | 10 points
No | 0 points
D19
1.5-3 | 5 points
>3 | 0 points
W05
0≤30 | 20 points
31≤60 | 10 points
61≤90 | 0 points
W06
≥36 | 10 points
>12<36 | 5 points
≤12 | 0 points